## Scale (G1-1)
Scale Calibration Worksheet
| Equipment ID: | G1 - 1 |  |  |  |  | Date: | 2006-09-28
| Manufacturer: | Ohaus |  |  |  |  | Performed by: | R. Conley
| Model #: | EOD 120
| Serial #: | B129005780 |  |  |  |  | Next Calibration Due: | September 2007
| UofA ID: | 216761 |  |  |  |  | Last Calibration: | NA
| Location: | STO (218)
| Calibration Item: | Verify mass readings for at least 5 points in the range of the balance
Check sensitivity (0.1%) for at least 5 points in the range of the balance
| Calibration Procedure: | AASHTO M-231
| Calibration Equipment: | Class F, Class I, and 10-kg Calibration Weights
Tolerances:
| Class |  | Readability and Sensitivity |  |  |  | Accuracy
| G1 |  |  | 0.01 g |  | 0.02 g or 0.1 percent
| G2 |  |  | 0.1 g |  | 0.2 or 0.1 percent
| G5 |  |  | 1 g |  | 2 g or 0.1 percent
| G20 |  |  | 5 g |  | 5 g or 0.1 percent
| G100 |  |  | 20 g |  | 20 g or 0.1 percent
| Weight Applied (g) |  | Scale Reading (g) |  | Tolerance (g) |  |  | Pass/Fail
| 1 |  | 1 |  | 0.02
| 1.01 |  | 1.01 |  | 0.02
| 2 |  | 2 |  | 0.02
| 2.01 |  | 2.01 |  | 0.02
| 5 |  | 5 |  | 0.02
| 5.01 |  | 5.01 |  | 0.02
| 10 |  | 10 |  | 0.02
| 20 |  | 20 |  | 0.02
| 50 |  | 49.99 |  | 0.05
| 100 |  | 100 |  | 0.1
| 200 |  | 200.01 |  | 0.2
| 500 |  | 500 |  | 0.5
| 1000 |  | 999.9 |  | 1
| 4000 |  | 4000 |  | 4
| 4000.01 |  | 4000.3 |  | 4
[9 empty] | Pass/Fail: | Pass
| Comments: | Good
[9 empty] | Initial By: | sgw

## Scale (G1-2)
Scale Calibration Worksheet
| Equipment ID: | G1 - 2 |  |  |  |  | Date: | 2006-09-28
| Manufacturer: | Ohaus |  |  |  |  | Performed by: | R. Conley
| Model #: | EOD 120
| Serial #: | E2411119231431 |  |  |  |  | Next Calibration Due: | September 2007
| UofA ID: | [5 empty] | Last Calibration: | NA
| Location: | STO (218)
| Calibration Item: | Verify mass readings for at least 5 points in the range of the balance
Check sensitivity (0.1%) for at least 5 points in the range of the balance
| Calibration Procedure: | AASHTO M-231
| Calibration Equipment: | Class F, Class I, and 10-kg Calibration Weights
Tolerances:
| Class |  | Readability and Sensitivity |  |  |  | Accuracy
| G1 |  |  | 0.01 g |  | 0.02 g or 0.1 percent
| G2 |  |  | 0.1 g |  | 0.2 or 0.1 percent
| G5 |  |  | 1 g |  | 2 g or 0.1 percent
| G20 |  |  | 5 g |  | 5 g or 0.1 percent
| G100 |  |  | 20 g |  | 20 g or 0.1 percent
| Weight Applied (g) |  | Scale Reading (g) |  | Tolerance (g) |  |  | Pass/Fail
| 1 |  | 1 |  | 0.02
| 1.01 |  | 1.01 |  | 0.02
| 2 |  | 2 |  | 0.02
| 2.01 |  | 2.01 |  | 0.02
| 5 |  | 5 |  | 0.02
| 5.01 |  | 5.01 |  | 0.02
| 10 |  | 10 |  | 0.02
| 20 |  | 20 |  | 0.02
| 50 |  | 49.99 |  | 0.05
| 100 |  | 100 |  | 0.1
| 200 |  | 199.99 |  | 0.2
| 500 |  | 499.98 |  | 0.5
| 1000 |  | 999.94 |  | 1
| 4000 |  | 4000.03 |  | 4
| 4000.01 |  | 4000.03 |  | 4
[9 empty] | Pass/Fail: | Pass
| Comments: | Good
[9 empty] | Initial By: | sgw

## Scale (G2-A)
Scale Calibration Worksheet
| Equipment ID: | AG - A |  |  |  |  | Date: | 2006-09-28
| Manufacturer: | Ohaus |  |  |  |  | Performed by: | R. Conley
| Model #: | EP8101C
| Serial #: | H1761122222411 P |  |  |  |  | Next Calibration Due: | September 2007
| UofA ID: | NA |  |  |  |  | Last Calibration: | September 2005
| Location: | AG (233) Rear Table (West)
| Calibration Item: | Verify mass readings for at least 5 points in the range of the balance
Check sensitivity (0.1%) for at least 5 points in the range of the balance
| Calibration Procedure: | AASHTO M-231
| Calibration Equipment: | Class F, Class I, and 10-kg Calibration Weights
Tolerances:
| Class |  | Readability and Sensitivity |  |  |  | Accuracy
| G1 |  |  | 0.01 g |  | 0.02 g or 0.1 percent
| G2 |  |  | 0.1 g |  | 0.2 or 0.1 percent
| G5 |  |  | 1 g |  | 2 g or 0.1 percent
| G20 |  |  | 5 g |  | 5 g or 0.1 percent
| G100 |  |  | 20 g |  | 20 g or 0.1 percent
| Weight Applied (g) |  | Scale Reading (g) |  | Tolerance (g) |  |  | Pass/Fail
| 1 |  | 1 |  | 0.2
| 5 |  | 5 |  | 0.2
| 10 |  | 10 |  | 0.2
| 50 |  | 49.9 |  | 0.2
| 100 |  | 100 |  | 0.2
| 100.1 |  | 100.1 |  | 0.2
| 200 |  | 200 |  | 0.2
| 500 |  | 500 |  | 0.5
| 500.1 |  | 500.1 |  | 0.5
| 1000 |  | 1000 |  | 1
| 2000 |  | 2000.1 |  | 2
| 2000.1 |  | 2000.2 |  | 2
| 5000 |  | 5000.2 |  | 5
| 5000.1 |  | 5000.3 |  | 5
| 7500 |  | 7500.2 |  | 7.5
[9 empty] | Pass/Fail: | Pass
Comments:
[9 empty] | Initial By: | sgw

## Scale (G2-B)
Scale Calibration Worksheet
| Equipment ID: | AG - B |  |  |  |  | Date: | 2006-09-28
| Manufacturer: | Ohaus |  |  |  |  | Performed by: | S. Williams
| Model #: | EP12001
| Serial #: | 1125430302 |  |  |  |  | Next Calibration Due: | September 2007
| UofA ID: | [5 empty] | Last Calibration: | September 2005
| Location: | AG (233) Front Table (East)
| Calibration Item: | Verify mass readings for at least 5 points in the range of the balance
Check sensitivity (0.1%) for at least 5 points in the range of the balance
| Calibration Procedure: | AASHTO M-231
| Calibration Equipment: | Class F, Class I, and 10-kg Calibration Weights
Tolerances:
| Class |  | Readability and Sensitivity |  |  |  | Accuracy
| G1 |  |  | 0.01 g |  | 0.02 g or 0.1 percent
| G2 |  |  | 0.1 g |  | 0.2 or 0.1 percent
| G5 |  |  | 1 g |  | 2 g or 0.1 percent
| G20 |  |  | 5 g |  | 5 g or 0.1 percent
| G100 |  |  | 20 g |  | 20 g or 0.1 percent
| Weight Applied (g) |  | Scale Reading (g) |  | Tolerance (g) |  |  | Pass/Fail
| 1 |  | 1 |  | 0.2
| 5 |  | 5 |  | 0.2
| 10 |  | 10 |  | 0.2
| 50 |  | 50 |  | 0.2
| 100 |  | 100 |  | 0.2
| 100.1 |  | 100.1 |  | 0.2
| 200 |  | 200 |  | 0.2
| 500 |  | 500 |  | 0.5
| 500.1 |  | 500.1 |  | 0.5
| 1000 |  | 1000.1 |  | 1
| 2000 |  | 2000.2 |  | 2
| 2000.1 |  | 2000.3 |  | 2
| 5000 |  | 5000.3 |  | 5
| 5000.1 |  | 5000.4 |  | 5
| 7500 |  | 7500.5 |  | 7.5
[9 empty] | Pass/Fail: | Pass
Comments:
[9 empty] | Initial By: | sgw

## Scale (G2-E)
Scale Calibration Worksheet
| Equipment ID: | AG - E |  |  |  |  | Date: | 2006-09-28
| Manufacturer: | Ohaus |  |  |  |  | Performed by: | R. Conley
| Model #: | EP12001
| Serial #: | 1125330661 |  |  |  |  | Next Calibration Due: | September 2007
| UofA ID: | NA |  |  |  |  | Last Calibration: | NA (new)
| Location: | AG (233) Front SpG (East)
| Calibration Item: | Verify mass readings for at least 5 points in the range of the balance
Check sensitivity (0.1%) for at least 5 points in the range of the balance
| Calibration Procedure: | AASHTO M-231
| Calibration Equipment: | Class F, Class I, and 10-kg Calibration Weights
Tolerances:
| Class |  | Readability and Sensitivity |  |  |  | Accuracy
| G1 |  |  | 0.01 g |  | 0.02 g or 0.1 percent
| G2 |  |  | 0.1 g |  | 0.2 or 0.1 percent
| G5 |  |  | 1 g |  | 2 g or 0.1 percent
| G20 |  |  | 5 g |  | 5 g or 0.1 percent
| G100 |  |  | 20 g |  | 20 g or 0.1 percent
| Weight Applied (g) |  | Scale Reading (g) |  | Tolerance (g) |  |  | Pass/Fail
| 1 |  | 1 |  | 0.2
| 10 |  | 10 |  | 0.2
| 100 |  | 100 |  | 0.2
| 100.1 |  | 100 |  | 0.2
| 200 |  | 199.8 |  | 0.2
| 200.1 |  | 199.9 |  | 0.2
| 500 |  | 499.6 |  | 0.5
| 500.1 |  | 499.7 |  | 0.5
| 1000 |  | 999 |  | 1
| 2000 |  | 1998.1 |  | 2
| 5000 |  | 4995.3 |  | 5
| 5000.1 |  | 4995.4 |  | 5
| 10000 |  | 9991.2 |  | 10
[9 empty] | Pass/Fail: | Pass
Comments:
[9 empty] | Initial By: | sgw

## Scale (G2-F)
Scale Calibration Worksheet
| Equipment ID: | AG - F |  |  |  |  | Date: | 2006-09-28
| Manufacturer: | Ohaus |  |  |  |  | Performed by: | R. Conley
| Model #: | EP12001
| Serial #: | 1126051479 |  |  |  |  | Next Calibration Due: | September 2007
| UofA ID: | NA |  |  |  |  | Last Calibration: | NA (new)
| Location: | AG (233) Rear SpG (West)
| Calibration Item: | Verify mass readings for at least 5 points in the range of the balance
Check sensitivity (0.1%) for at least 5 points in the range of the balance
| Calibration Procedure: | AASHTO M-231
| Calibration Equipment: | Class F, Class I, and 10-kg Calibration Weights
Tolerances:
| Class |  | Readability and Sensitivity |  |  |  | Accuracy
| G1 |  |  | 0.01 g |  | 0.02 g or 0.1 percent
| G2 |  |  | 0.1 g |  | 0.2 or 0.1 percent
| G5 |  |  | 1 g |  | 2 g or 0.1 percent
| G20 |  |  | 5 g |  | 5 g or 0.1 percent
| G100 |  |  | 20 g |  | 20 g or 0.1 percent
| Weight Applied (g) |  | Scale Reading (g) |  | Tolerance (g) |  |  | Pass/Fail
| 1 |  | 1 |  | 0.2
| 10 |  | 10 |  | 0.2
| 100 |  | 100 |  | 0.2
| 100.1 |  | 100.1 |  | 0.2
| 200 |  | 199.9 |  | 0.2
| 200.1 |  | 200 |  | 0.2
| 500 |  | 499.6 |  | 0.5
| 500.1 |  | 499.7 |  | 0.5
| 1000 |  | 999.1 |  | 1
| 2000 |  | 1998.4 |  | 2
| 5000 |  | 4995.8 |  | 5
| 5000.1 |  | 4995.9 |  | 5
| 10000 |  | 9992.3 |  | 10
[9 empty] | Pass/Fail: | Pass
Comments:
[9 empty] | Initial By: | sgw

## Scale (G2-IO)
Scale Calibration Worksheet
| Equipment ID: | Ignition Oven |  |  |  |  | Date: | 2006-09-28
| Manufacturer: | Ohaus |  |  |  |  | Performed by: | R. Conley
| Model #: | GT 8000
| Serial #: | 945960260909 |  |  |  |  | Next Calibration Due: | September 2007
| UofA ID: | NA |  |  |  |  | Last Calibration: | September 2005
| Location: | HM (225) Ignition Oven
| Calibration Item: | Verify mass readings for at least 5 points in the range of the balance
Check sensitivity (0.1%) for at least 5 points in the range of the balance
| Calibration Procedure: | AASHTO M-231
| Calibration Equipment: | Class F, Class I, and 10-kg Calibration Weights
Tolerances:
| Class |  | Readability and Sensitivity |  |  |  | Accuracy
| G1 |  |  | 0.01 g |  | 0.02 g or 0.1 percent
| G2 |  |  | 0.1 g |  | 0.2 or 0.1 percent
| G5 |  |  | 1 g |  | 2 g or 0.1 percent
| G20 |  |  | 5 g |  | 5 g or 0.1 percent
| G100 |  |  | 20 g |  | 20 g or 0.1 percent
| Weight Applied (g) |  | Scale Reading (g) |  | Tolerance (g) |  |  | Pass/Fail
| 10 |  | 10 |  | 0.2
| 10.1 |  | 10.1 |  | 0.2
| 20 |  | 20 |  | 0.2
| 50 |  | 50 |  | 0.2
| 100 |  | 100 |  | 0.2
| 100.1 |  | 100.1 |  | 0.2
| 200 |  | 200 |  | 0.2
| 500 |  | 500.1 |  | 0.5
| 1000 |  | 1000.2 |  | 1
| 1000.1 |  | 1000.3 |  | 1
| 2000 |  | 2000.5 |  | 2
| 5000 |  | 5001.2 |  | 5
| 5000.1 |  | 5001.3 |  | 5
[9 empty] | Pass/Fail: | Pass
Comments:
[9 empty] | Initial By: | sgw

## Scale (G5-MX1)
Scale Calibration Worksheet
| Equipment ID: | MX - 1 |  |  |  |  | Date: | 2006-09-28
| Manufacturer: | Mettler |  |  |  |  | Performed by: | R. Conley
| Model #: | SG 32000
| Serial #: | 2115275579 |  |  |  |  | Next Calibration Due: | September 2007
| UofA ID: | NA |  |  |  |  | Last Calibration: | September 2005
| Location: | MX (234) (Right)
| Calibration Item: | Verify mass readings for at least 5 points in the range of the balance
Check sensitivity (0.1%) for at least 5 points in the range of the balance
| Calibration Procedure: | AASHTO M-231
| Calibration Equipment: | Class F, Class I, and 10-kg Calibration Weights
Tolerances:
| Class |  | Readability and Sensitivity |  |  |  | Accuracy
| G1 |  |  | 0.01 g |  | 0.02 g or 0.1 percent
| G2 |  |  | 0.1 g |  | 0.2 or 0.1 percent
| G5 |  |  | 1 g |  | 2 g or 0.1 percent
| G20 |  |  | 5 g |  | 5 g or 0.1 percent
| G100 |  |  | 20 g |  | 20 g or 0.1 percent
| Weight Applied (g) |  | Scale Reading (g) |  | Tolerance (g) |  |  | Pass/Fail
| 1 |  | 1 |  | 2
| 5 |  | 5 |  | 2
| 10 |  | 10 |  | 2
| 50 |  | 50 |  | 2
| 100 |  | 100 |  | 2
| 500 |  | 500 |  | 2
| 1000 |  | 1000 |  | 2
| 5000 |  | 5000 |  | 5
| 5001 |  | 5001 |  | 5
| 10000 |  | 10001 |  | 10
| 10001 |  | 10002 |  | 10
| 20000 |  | 20001 |  | 20
| 20001 |  | 20002 |  | 20
[9 empty] | Pass/Fail: | Pass
| Comments: | Very good
[9 empty] | Initial By: | sgw

## Scale (G5-MX2)
Scale Calibration Worksheet
| Equipment ID: | MX - 2 |  |  |  |  | Date: | 2006-09-28
| Manufacturer: | Mettler |  |  |  |  | Performed by: | R. Conley
| Model #: | SG 32000
| Serial #: | 1120303652 |  |  |  |  | Next Calibration Due: | September 2007
| UofA ID: | NA |  |  |  |  | Last Calibration: | NA
| Location: | MX (234) (Left)
| Calibration Item: | Verify mass readings for at least 5 points in the range of the balance
Check sensitivity (0.1%) for at least 5 points in the range of the balance
| Calibration Procedure: | AASHTO M-231
| Calibration Equipment: | Class F, Class I, and 10-kg Calibration Weights
Tolerances:
| Class |  | Readability and Sensitivity |  |  |  | Accuracy
| G1 |  |  | 0.01 g |  | 0.02 g or 0.1 percent
| G2 |  |  | 0.1 g |  | 0.2 or 0.1 percent
| G5 |  |  | 1 g |  | 2 g or 0.1 percent
| G20 |  |  | 5 g |  | 5 g or 0.1 percent
| G100 |  |  | 20 g |  | 20 g or 0.1 percent
| Weight Applied (g) |  | Scale Reading (g) |  | Tolerance (g) |  |  | Pass/Fail
| 1 |  | 1 |  | 2
| 5 |  | 5 |  | 2
| 10 |  | 10 |  | 2
| 50 |  | 50 |  | 2
| 100 |  | 100 |  | 2
| 500 |  | 500 |  | 2
| 1000 |  | 1000 |  | 2
| 5000 |  | 5000 |  | 5
| 5001 |  | 5001 |  | 5
| 10000 |  | 10000 |  | 10
| 10001 |  | 10001 |  | 10
| 20000 |  | 20001 |  | 20
| 20001 |  | 20002 |  | 20
[9 empty] | Pass/Fail: | Pass
| Comments: | Very good
[9 empty] | Initial By: | sgw