## Oven AG2
Oven Calibration Worksheet
Equipment ID: | AG2 |  | Serial #: | 187253
Manufacturer: | Despatch |  | UofA ID: | NA
Model #: | LBB1-69A-1 |  | Location: |  | Agg Lab (233)
Calibration Item: | Verify temperature readings |  |  | Typical Set Temperature: | 230 ° F
Calibration Procedure: | In-House Procedure #00-1 |  |  | Tolerance : | ± 9 °
Calibration Equipment: | Thermometer ASTM | 67 F
Temp. Readings
Date: | 2014-02-18 | Set Temp. | 1 | 2 | 3
Performed by: | M. Fleck | 230 | 224.7 | 223.8 | 225
Next Calibration Due: | Oct. 2014 | 235 | 225.9 | 230 | 229.5
Last Calibration: | New
Requirement: | ± 9 ° F
| Pass/Fail: | Pass |  |  | Initial by: | sgw
| Comments: | Digital Setpoint @ 235° F.
Constant temperature.
Temp. Readings
Date: | 2014-07-18 | Set Temp. | 1 | 2 | 3
Performed by: | M. Fleck | 240 | 227.4 | 230.5 | 229.4
Next Calibration Due: | Oct. 2015
Last Calibration: | 2/18/14
Requirement: | ± 9 ° F
| Pass/Fail: | Pass |  |  | Initial by: | sgw
| Comments: | Installed rear exhaust fan. Oven temperature set
to 240° F.
Temp. Readings
Date: | 2014-10-02 | Set Temp. | 1 | 2 | 3
Performed by: | Trish S. | 230 | 230.6 | 231.6 | 231.3
Next Calibration Due: | Oct. 2016
Last Calibration: | Oct. 2015
Requirement: | ± 9 ° F
| Pass/Fail: | Pass |  |  | Initial by:
| Comments: | Oven set point at 240°F.

## Oven HM1
Oven Calibration Worksheet
Equipment ID: | HM - 1 |  | Serial #: | 170436
Manufacturer: | Despatch |  | UofA ID: | 255801
Model #: | LBB2 - 18 - 1 |  | Location: |  | Hot Mix Lab (225)
Calibration Item: | Verify temperature readings |  |  | Typical Set Temperature: | 300 ° F
Calibration Procedure: | In-House Procedure #00-1 |  |  | Tolerance : | ± 5 ° F
Calibration Equipment: | Thermometer ASTM | 67F / 68F
Temp. Readings
Date: | 2012-10-31 | Set Temp. | 1 | 2 | 3
Performed by: | M. Fleck | 270 | 270.4 | 274.9 | 271.1
Next Calibration Due: | October 2013 | 300 | 301.9 | 302.4 | 302.8
Last Calibration: | October 2011 | 328 | 325.5 | 326 | 326.5
Requirement: | ± 5 ° F
| Pass/Fail: | Pass |  |  | Initial by: | sgw
| Comments: | Digital setpoints.
3° F SP Offset Note added to Oven
Temp. Readings
Date: | 2013-10-11 | Set Temp. | 1 | 2 | 3
Performed by: | M. Fleck | 230 | 233.5 | 233.7 | 234
Next Calibration Due: | October 2014 | 280 | 282.4 | 282.7 | 282
Last Calibration: | October 2012 | 315 | 313.1 | 314.1 | 313.8
Requirement: | ± 5 ° F
| Pass/Fail: | Pass |  |  | Initial by: | sgw
| Comments: | Digital setpoints.
3° F SP Offset Note added to Oven
Temp. Readings
Date: | 2014-10-01 | Set Temp. | 1 | 2 | 3
Performed by: | Cardia R. | 250 | 251.1 | 250.9 | 250.9
Next Calibration Due: | October 2015 | 290 | 288.9 | 288.9 | 289.1
Last Calibration: | October 2013 | 325 | 320.7 | 321.3 | 321.7
Requirement: | ± 5 ° F
| Pass/Fail: | Pass |  |  | Initial by:
| Comments: | Digital setpoints.
3° F SP Offset Note added to Oven

## Oven MX5
Oven Calibration Worksheet
Equipment ID: | MX - 5 |  | Serial #: | 9040305
Manufacturer: | Sheldon Mfg., Inc. |  | UofA ID: | 259372
Model #: | 1685 |  | Location: |  | Mixing Lab (234)
Calibration Item: | Verify temperature readings |  |  | Typical Set Temperature: | 330 ° F
Calibration Procedure: | In-House Procedure #00-1 |  |  | Tolerance : | ± 5 ° F
Calibration Equipment: | Thermometer ASTM | 68 F
Temp. Readings
Date: | 2011-10-21 | Set Temp. | 1 | 2 | 3
Performed by: | Haley M. | 250 | 254.6 | 253.4 | 248.8
Next Calibration Due: | October 2012 | 290 | 291 | 286.3 | 289.4
Last Calibration: | October 2010 | 330 | 328.5 | 330.4 | 332.1
Requirement: | ± 5 ° F
| Pass/Fail: | Pass |  |  | Initial by: | sgw
| Comments: | Digital Set @ 250, 290, 330 °F.
Temp. Readings
Date: | 2012-10-23 | Set Temp. | 1 | 2 | 3
Performed by: | M. Fleck | 250 | 247.7 | 246.9 | 247.3
Next Calibration Due: | October 2013 | 290 | 286.1 | 285.7 | 285.1
Last Calibration: | October 2011 | 330 | 322.7 | 323.2 | 323.8
Requirement: | ± 5 ° F
| Pass/Fail: | Pass |  |  | Initial by: | sgw
| Comments: | Digital Set @ 250, 290, 330 °F.
5° F Correction Tag Added to Oven
Temp. Readings
Date: | 2013-10-09 | Set Temp. | 1 | 2 | 3
Performed by: | M. Fleck | 254 | 252.3 | 252.3 | 252.5
Next Calibration Due: | October 2014 | 300 | 293.6 | 294.1 | 294.7
Last Calibration: | October 2012 | 350 | 340.5 | 340.8 | 340.7
Requirement: | ± 5 ° F
| Pass/Fail: | Pass |  |  | Initial by: | sgw
| Comments: | Digital Set @ 254, 300, 350 °F.
5° F Correction Tag Added to Oven

## Oven MX5 (2)
Oven Calibration Worksheet
Equipment ID: | MX - 5 |  | Serial #: | 9040305
Manufacturer: | Sheldon Mfg., Inc. |  | UofA ID: | 259372
Model #: | 1685 |  | Location: |  | Mixing Lab (234)
Calibration Item: | Verify temperature readings |  |  | Typical Set Temperature: | 330 ° F
Calibration Procedure: | In-House Procedure #00-1 |  |  | Tolerance : | ± 5 ° F
Calibration Equipment: | Thermometer ASTM | 68 F
Temp. Readings
Date: | 2014-10-03 | Set Temp. | 1 | 2 | 3
Performed by: | M. Fleck | 250 | 247.6 | 246.7 | 247.1
Next Calibration Due: | October 2015 | 300 | 294.9 | 294.7 | 294.7
Last Calibration: | October 2013 | 350 | 342.1 | 342.5 | 342.7
Requirement: | ± 5 ° F
| Pass/Fail: | Pass |  |  | Initial by:
| Comments: | Digital Set @ 250, 300, 350 °F.
5° F Correction Tag Added to Oven
Temp. Readings
Date: |  | Set Temp. | 1 | 2 | 3
Performed by: |  | 230
Next Calibration Due: | October 2016 | 280
Last Calibration: | October 2014 | 330
Requirement: | ± 5 ° F
| Pass/Fail: |  |  |  | Initial by:
| Comments: | Digital Set @ 230, 280, 330 °F.
5° F Correction Tag Added to Oven
Temp. Readings
Date: |  | Set Temp. | 1 | 2 | 3
Performed by: |  | 240
Next Calibration Due: | October 2017 | 290
Last Calibration: | October 2015 | 340
Requirement: | ± 5 ° F
| Pass/Fail: |  |  |  | Initial by:
| Comments: | Digital Set @ 240, 290, 340 °F.
5° F Correction Tag Added to Oven

## Oven MX1
Oven Calibration Worksheet
Equipment ID: | MX - 1 |  | Serial #: | 9602246
Manufacturer: | Quincy Lab |  | UofA ID: | 234300
Model #: | 21 - 350 |  | Location: |  | Mixing Lab (234)
Calibration Item: | Verify temperature readings |  |  | Typical Set Temperature: | 230 ° F
Calibration Procedure: | In-House Procedure #00-1 |  |  | Tolerance : | ± 9 ° F
Calibration Equipment: | Thermometer ASTM | 67 F
Temp. Readings
Date: | 2012-10-22 | Set Temp. | 1 | 2 | 3
Performed by: | M. Fleck
Next Calibration Due: | October 2013 | 230 | 224.5 | 228 | 233.4
Last Calibration: | October 2011
Requirement: | ± 9 ° F
| Pass/Fail: | Pass |  |  | Initial by: | sgw
| Comments: | Dial set @ 230° mark.
Constant temperature.
Temp. Readings
Date: | 2013-10-09 | Set Temp. | 1 | 2 | 3
Performed by: | M. Fleck
Next Calibration Due: | October 2014 | 230 | 232.1 | 235.9 | 225.6
Last Calibration: | October 2012
Requirement: | ± 9 ° F
| Pass/Fail: | Pass |  |  | Initial by: | sgw
| Comments: | Dial set @ 230° mark.
Constant temperature.
Temp. Readings
Date: | 2014-02-01 | Set Temp. | 1 | 2 | 3
Performed by: | M. Fleck
Next Calibration Due: | October 2015 |  | OOS
Last Calibration: | October 2013
Requirement: | ± 9 ° F
| Pass/Fail: | Fail |  |  | Initial by:
| Comments: | Oven fails to maintain required temperature.
Replaced oven with AG2.

## Oven MX4
Oven Calibration Worksheet
Equipment ID: | MX - 4 |  | Serial #: | 3064606
Manufacturer: | Sheldon Mfg., Inc. |  | UofA ID: | N / A
Model #: | 1370FZZ |  | Location: |  | Mixing Lab (234)
Calibration Item: | Verify temperature readings |  |  | Typical Set Temperature: | 330 ° F
Calibration Procedure: | In-House Procedure #00-1 |  |  | Tolerance : | ± 5 ° F
Calibration Equipment: | Thermometer ASTM | 68 F
Temp. Readings
Date: | 2012-10-22 | Set Temp. | 1 | 2 | 3
Performed by: | M. Fleck | 200 | 202.7 | 198.1 | 195.1
Next Calibration Due: | October 2013 | 280 | 277.7 | 274 | 271.3
Last Calibration: | October 2011 | 330 | 320.3 | 316 | 312.9
Requirement: | ± 5 ° F
| Pass/Fail: | Pass |  |  | Initial by: | sgw
| Comments: | Calibration sheet added to oven door.
Temp. Readings
Date: | 2013-10-10 | Set Temp. | 1 | 2 | 3
Performed by: | M. Fleck | 300 | 302.9 | 304 | 301.5
Next Calibration Due: | October 2014 | 257 | 254.4 | 256.5 | 258.9
Last Calibration: | October 2012 | 202 | 201.7 | 199.1 | 197.3
Requirement: | ± 5 ° F
| Pass/Fail: | Pass |  |  | Initial by: | sgw
| Comments: | Set temperature is according to calibration sheet
based on set point for oven.
Temp. Readings
Date: | 2014-02-01 | Set Temp. | 1 | 2 | 3
Performed by: | M. Fleck
Next Calibration Due: | October 2015 |  | OOS
Last Calibration: | October 2013
Requirement: | ± 5 ° F
| Pass/Fail: |  |  |  | Initial by:
| Comments: | Oven no longer needed.