## 2" Sieve
Sieve Calibration Worksheet
|  | Equipment ID: | 2" Sieve |  |  |  |  | Serial #: | 11147403
|  | Manufacturer: | Gilson |  |  |  |  | UofA ID: | NA
|  | Model #: | V8SF"200 |  |  |  |  | Location: | ST (218)
| Calibration Item: | Verify physical condition and critical dimensions for sieves (#4 and larger)
| Calibration Procedure: | In-House Procedure #00-5
| Calibration Equipment: | Mititoyo Digimatic Calipers
| Requirement (according to AASHTO M 92) = |  |  |  |  | 50 ± 1.50 mm
(48.5mm - 51.5mm)
|  | Date: | 9/18/14 |  |  |  |  | Date:
|  | Performed by: | Cardia R. |  |  |  |  | Performed by:
|  | Next Calib. Due: | Sept. 2015 |  |  |  |  | Next Calib. Due: | Sept. 2016
|  | Last Calibration: | Sept. 2013 |  |  |  |  | Last Calibration: | Sept. 2014
|  | 1 | 50.21 |  | 50.5 |  |  | 1
|  | 2 | 50.57 |  | 50.58 |  |  | 2
|  | 3 | 50.6 |  | 50.52 |  |  | 3
|  | 4 | 50.62 |  | 50.7 |  |  | 4
|  | 5 | 50.48 |  | 50.37 |  |  | 5
|  | 6 | [5 empty] | 6
|  | 7 | [5 empty] | 7
|  | 8 | [5 empty] | 8
|  | 9 | [5 empty] | 9
|  | 10 | [5 empty] | 10
|  | Frame: | Excellent |  |  |  |  | Frame:
|  | Joints: | Excellent |  |  |  |  | Joints:
|  | Cloth: | Excellent |  |  |  |  | Cloth:
|  | Tightness: | Excellent |  |  |  |  | Tightness:
|  | Other: | [5 empty] | Other:
| Comments: | [5 empty] | Comments:
|  |  | Pass/Fail: | Pass |  |  |  |  | Pass/Fail:
|  |  | Initial By: | [5 empty] | Initial By:

## 1.5" Sieve
Sieve Calibration Worksheet
|  | Equipment ID: | 1 1/2" Sieve |  |  |  |  | Serial #: | 11147404
|  | Manufacturer: | Gilson |  |  |  |  | UofA ID: | NA
|  | Model #: | V8SS"150 |  |  |  |  | Location: | ST (218)
| Calibration Item: | Verify physical condition and critical dimensions for sieves (#4 and larger)
| Calibration Procedure: | In-House Procedure #00-5
| Calibration Equipment: | Mititoyo Digimatic Calipers
| Requirement (according to AASHTO M 92) = |  |  |  |  | 37.5 ± 1.10 mm
(36.4mm - 38.6mm)
|  | Date: | 9/18/14 |  |  |  |  | Date:
|  | Performed by: | Cardia R. |  |  |  |  | Performed by:
|  | Next Calib. Due: | Sept. 2015 |  |  |  |  | Next Calib. Due: | Sept. 2016
|  | Last Calibration: | Sept. 2013 |  |  |  |  | Last Calibration: | Sept. 2014
|  | 1 | 37.51 |  | 37.53 |  |  | 1
|  | 2 | 37.45 |  | 37.5 |  |  | 2
|  | 3 | 37.54 |  | 37.52 |  |  | 3
|  | 4 | 37.47 |  | 37.49 |  |  | 4
|  | 5 | 37.61 |  | 37.52 |  |  | 5
|  | 6 | 37.64 |  | 37.72 |  |  | 6
|  | 7 | 37.74 |  | 37.71 |  |  | 7
|  | 8 | 37.6 |  | 37.59 |  |  | 8
|  | 9 | 37.6 |  | 37.76 |  |  | 9
|  | 10 | 37.66 |  | 37.68 |  |  | 10
|  | Frame: | Excellent |  |  |  |  | Frame:
|  | Joints: | Excellent |  |  |  |  | Joints:
|  | Cloth: | Excellent |  |  |  |  | Cloth:
|  | Tightness: | Excellent |  |  |  |  | Tightness:
|  | Other: | [5 empty] | Other:
| Comments: | [5 empty] | Comments:
|  |  | Pass/Fail: | Pass |  |  |  |  | Pass/Fail:
|  |  | Initial By: | [5 empty] | Initial By:

## 1" Sieve
Sieve Calibration Worksheet
|  | Equipment ID: | 1" Sieve |  |  |  |  | Serial #: | 11147405
|  | Manufacturer: | Gilson |  |  |  |  | UofA ID: | NA
|  | Model #: | V8SF"100 |  |  |  |  | Location: | ST(218)
| Calibration Item: | Verify physical condition and critical dimensions for sieves (#4 and larger)
| Calibration Procedure: | In-House Procedure #00-5
| Calibration Equipment: | Mititoyo Digimatic Calipers
| Requirement (according to AASHTO M 92) = |  |  |  |  | 25 ± 0.80 mm
(24.2mm - 25.8mm)
|  | Date: | 9/18/14 |  |  |  |  | Date:
|  | Performed by: | Cardia R. |  |  |  |  | Performed by:
|  | Next Calib. Due: | Sept. 2015 |  |  |  |  | Next Calib. Due: | Sept. 2016
|  | Last Calibration: | Sept. 2013 |  |  |  |  | Last Calibration: | Sept. 2014
|  | 1 | 24.89 |  | 24.81 |  |  | 1
|  | 2 | 25 |  | 24.99 |  |  | 2
|  | 3 | 25.01 |  | 25.04 |  |  | 3
|  | 4 | 24.88 |  | 25.13 |  |  | 4
|  | 5 | 24.86 |  | 24.96 |  |  | 5
|  | 6 | 25.1 |  | 24.9 |  |  | 6
|  | 7 | 24.9 |  | 24.98 |  |  | 7
|  | 8 | 24.95 |  | 25.08 |  |  | 8
|  | 9 | 25.03 |  | 25.1 |  |  | 9
|  | 10 | 24.9 |  | 25 |  |  | 10
|  | Frame: | Excellent |  |  |  |  | Frame:
|  | Joints: | Excellent |  |  |  |  | Joints:
|  | Cloth: | Excellent |  |  |  |  | Cloth:
|  | Tightness: | Excellent |  |  |  |  | Tightness:
|  | Other: | [5 empty] | Other:
| Comments: | [5 empty] | Comments:
|  |  | Pass/Fail: | Pass |  |  |  |  | Pass/Fail:
|  |  | Initial By: | [5 empty] | Initial By:

## .75" Sieve
Sieve Calibration Worksheet
|  | Equipment ID: | 3/4" Sieve |  |  |  |  | Serial #: | 11147406
|  | Manufacturer: | Gilson |  |  |  |  | UofA ID: | NA
|  | Model #: | V8SF"075 |  |  |  |  | Location: | ST(218)
| Calibration Item: | Verify physical condition and critical dimensions for sieves (#4 and larger)
| Calibration Procedure: | In-House Procedure #00-5
| Calibration Equipment: | Mititoyo Digimatic Calipers
| Requirement (according to AASHTO M 92) = |  |  |  |  | 19 ± 0.60 mm
(18.4mm - 19.6mm)
|  | Date: | 9/18/14 |  |  |  |  | Date:
|  | Performed by: | Cardia R. |  |  |  |  | Performed by:
|  | Next Calib. Due: | Sept. 2015 |  |  |  |  | Next Calib. Due: | Sept. 2016
|  | Last Calibration: | Sept. 2013 |  |  |  |  | Last Calibration: | Sept. 2014
|  | 1 | 19.33 |  | 19.25 |  |  | 1
|  | 2 | 19.32 |  | 19.3 |  |  | 2
|  | 3 | 19.35 |  | 19.26 |  |  | 3
|  | 4 | 19.22 |  | 19.28 |  |  | 4
|  | 5 | 19.31 |  | 19.33 |  |  | 5
|  | 6 | 19.31 |  | 19.23 |  |  | 6
|  | 7 | 19.3 |  | 19.39 |  |  | 7
|  | 8 | 19.37 |  | 19.33 |  |  | 8
|  | 9 | 19.38 |  | 19.36 |  |  | 9
|  | 10 | 19.27 |  | 19.33 |  |  | 10
|  | Frame: | Excellent |  |  |  |  | Frame:
|  | Joints: | Excellent |  |  |  |  | Joints:
|  | Cloth: | Excellent |  |  |  |  | Cloth:
|  | Tightness: | Excellent |  |  |  |  | Tightness:
|  | Other: | [5 empty] | Other:
| Comments: | [5 empty] | Comments:
|  |  | Pass/Fail: | Pass |  |  |  |  | Pass/Fail:
|  |  | Initial By: | [5 empty] | Initial By:

## .5" Sieve
Sieve Calibration Worksheet
|  | Equipment ID: | 1/2" Sieve |  |  |  |  | Serial #: | 11147407
|  | Manufacturer: | Gilson |  |  |  |  | UofA ID: | NA
|  | Model #: | V8SF"050 |  |  |  |  | Location: | ST(218)
| Calibration Item: | Verify physical condition and critical dimensions for sieves (#4 and larger)
| Calibration Procedure: | In-House Procedure #00-5
| Calibration Equipment: | Mititoyo Digimatic Calipers
| Requirement (according to AASHTO M 92) = |  |  |  |  | 12.5 ± 0.39 mm
(12.11mm-12.89mm)
|  | Date: | 9/18/14 |  |  |  |  | Date:
|  | Performed by: | Cardia R. |  |  |  |  | Performed by:
|  | Next Calib. Due: | Sept. 2015 |  |  |  |  | Next Calib. Due: | Sept. 2016
|  | Last Calibration: | Sept. 2013 |  |  |  |  | Last Calibration: | Sept. 2014
|  | 1 | 12.57 |  | 12.64 |  |  | 1
|  | 2 | 12.6 |  | 12.66 |  |  | 2
|  | 3 | 12.54 |  | 12.62 |  |  | 3
|  | 4 | 12.61 |  | 12.66 |  |  | 4
|  | 5 | 12.63 |  | 12.64 |  |  | 5
|  | 6 | 12.59 |  | 12.63 |  |  | 6
|  | 7 | 12.56 |  | 12.58 |  |  | 7
|  | 8 | 12.59 |  | 12.62 |  |  | 8
|  | 9 | 12.64 |  | 12.71 |  |  | 9
|  | 10 | 12.56 |  | 12.58 |  |  | 10
|  | Frame: | Excellent |  |  |  |  | Frame:
|  | Joints: | Excellent |  |  |  |  | Joints:
|  | Cloth: | Excellent |  |  |  |  | Cloth:
|  | Tightness: | Excellent |  |  |  |  | Tightness:
|  | Other: | [5 empty] | Other:
| Comments: | [5 empty] | Comments:
|  |  | Pass/Fail: | Pass |  |  |  |  | Pass/Fail:
|  |  | Initial By: | [5 empty] | Initial By:

## 3-8" Sieve
Sieve Calibration Worksheet
|  | Equipment ID: | 3/8" Sieve |  |  |  |  | Serial #: | 11147408
|  | Manufacturer: | Gilson |  |  |  |  | UofA ID: | NA
|  | Model #: | V8SF"038 |  |  |  |  | Location: | ST(218)
| Calibration Item: | Verify physical condition and critical dimensions for sieves (#4 and larger)
| Calibration Procedure: | In-House Procedure #00-5
| Calibration Equipment: | Mititoyo Digimatic Calipers
| Requirement (according to AASHTO M 92) = |  |  |  |  | 9.5 ± 0.30 mm
(9.20mm-9.80mm)
|  | Date: | 9/18/14 |  |  |  |  | Date:
|  | Performed by: | Cardia R. |  |  |  |  | Performed by:
|  | Next Calib. Due: | Sept. 2015 |  |  |  |  | Next Calib. Due: | Sept. 2016
|  | Last Calibration: | Sept. 2013 |  |  |  |  | Last Calibration: | Sept. 2014
|  | 1 | 9.73 |  | 9.63 |  |  | 1
|  | 2 | 9.53 |  | 9.64 |  |  | 2
|  | 3 | 9.61 |  | 9.56 |  |  | 3
|  | 4 | 9.75 |  | 9.71 |  |  | 4
|  | 5 | 9.55 |  | 9.6 |  |  | 5
|  | 6 | 9.53 |  | 9.45 |  |  | 6
|  | 7 | 9.51 |  | 9.58 |  |  | 7
|  | 8 | 9.68 |  | 9.6 |  |  | 8
|  | 9 | 9.52 |  | 9.57 |  |  | 9
|  | 10 | 9.66 |  | 9.63 |  |  | 10
|  | Frame: | Excellent |  |  |  |  | Frame:
|  | Joints: | Excellent |  |  |  |  | Joints:
|  | Cloth: | Excellent |  |  |  |  | Cloth:
|  | Tightness: | Excellent |  |  |  |  | Tightness:
|  | Other: | [5 empty] | Other:
| Comments: | [5 empty] | Comments:
|  |  | Pass/Fail: | Pass |  |  |  |  | Pass/Fail:
|  |  | Initial By: | [5 empty] | Initial By:

## #4 Sieve
Sieve Calibration Worksheet
|  | Equipment ID: | #4 Sieve |  |  |  |  | Serial #: | 11147409
|  | Manufacturer: | Gilson |  |  |  |  | UofA ID: | NA
|  | Model #: | V8SF#004 |  |  |  |  | Location: | ST(218)
| Calibration Item: | Verify physical condition and critical dimensions for sieves (#4 and larger)
| Calibration Procedure: | In-House Procedure #00-5
| Calibration Equipment: | Mititoyo Digimatic Calipers
| Requirement (according to AASHTO M 92) = |  |  |  |  | 4.75 ± 0.15 mm
(4.60mm - 4.90mm)
|  | Date: | 9/18/14 |  |  |  |  | Date:
|  | Performed by: | Cardia R. |  |  |  |  | Performed by:
|  | Next Calib. Due: | Sept. 2015 |  |  |  |  | Next Calib. Due: | Sept. 2016
|  | Last Calibration: | Sept. 2013 |  |  |  |  | Last Calibration: | Sept. 2014
|  | 1 | 4.65 |  | 4.78 |  |  | 1
|  | 2 | 4.81 |  | 4.75 |  |  | 2
|  | 3 | 4.78 |  | 4.83 |  |  | 3
|  | 4 | 4.81 |  | 4.82 |  |  | 4
|  | 5 | 4.67 |  | 4.77 |  |  | 5
|  | 6 | 4.74 |  | 4.77 |  |  | 6
|  | 7 | 4.73 |  | 4.67 |  |  | 7
|  | 8 | 4.75 |  | 4.76 |  |  | 8
|  | 9 | 4.75 |  | 4.8 |  |  | 9
|  | 10 | 4.72 |  | 4.74 |  |  | 10
|  | Frame: | Excellent |  |  |  |  | Frame:
|  | Joints: | Excellent |  |  |  |  | Joints:
|  | Cloth: | Excellent |  |  |  |  | Cloth:
|  | Tightness: | Excellent |  |  |  |  | Tightness:
|  | Other: | [5 empty] | Other:
| Comments: | [5 empty] | Comments:
|  |  | Pass/Fail: | Pass |  |  |  |  | Pass/Fail:
|  |  | Initial By: | [5 empty] | Initial By:

## # 8 Sieve
Sieve Calibration Worksheet
| Equipment ID: | # 8 Sieve |  |  | Serial #: | 11147410
| Manufacturer: | Gilson |  |  | UofA ID: | NA
| Model #: | V8SF#008 |  |  | Location: | ST (218)
[5 empty] | Calibration Interval: | 12 months
| Calibration Item: | Verify physical condition of sieve and critical dimensions for
sieves larger than #4 (4.75 mm)
| Calibration Procedure: | In-House Procedure #00-5
| Calibration Equipment: | Mititoyo Digimatic Calipers
Condition
Date | Frame | Joints | Cloth | Tightness | Pass / Fail | Checked By | Initial by
2011-04-14 | New | New | New | New | Pass | MF | sgw
2011-09-20 | Excellent | Excellent | Excellent | Excellent | Pass | Haley M. | sgw
2012-09-04 | Excellent | Excellent | Excellent | Excellent | Pass | Thomas M. | sgw
2013-09-10 | Excellent | Excellent | Excellent | Excellent | Pass | Abiy G. | sgw
2014-09-19 | Excellent | Excellent | Excellent | Excellent | Pass | MF

## # 10 Sieve
Sieve Calibration Worksheet
| Equipment ID: | # 10 Sieve |  |  | Serial #: | 11147411
| Manufacturer: | Gilson |  |  | UofA ID: | NA
| Model #: | V8SF#010 |  |  | Location: | ST (218)
[5 empty] | Calibration Interval: | 12 months
| Calibration Item: | Verify physical condition of sieve and critical dimensions for
sieves larger than #4 (4.75 mm)
| Calibration Procedure: | In-House Procedure #00-5
| Calibration Equipment: | Mititoyo Digimatic Calipers
Condition
Date | Frame | Joints | Cloth | Tightness | Pass / Fail | Checked By | Initial by
2011-04-14 | New | New | New | New | Pass | MF | sgw
2011-09-20 | Excellent | Excellent | Excellent | Excellent | Pass | Haley M. | sgw
2012-09-04 | Excellent | Excellent | Excellent | Excellent | Pass | Thomas M. | sgw
2013-09-10 | Excellent | Excellent | Excellent | Excellent | Pass | Abiy G. | sgw
2014-09-19 | Excellent | Excellent | Excellent | Excellent | Pass | MF

## # 16 Sieve
Sieve Calibration Worksheet
| Equipment ID: | # 16 Sieve |  |  | Serial #: | 11147412
| Manufacturer: | Gilson |  |  | UofA ID: | NA
| Model #: | V8SF# 016 |  |  | Location: | ST (218)
[5 empty] | Calibration Interval: | 12 months
| Calibration Item: | Verify physical condition of sieve and critical dimensions for
sieves larger than #4 (4.75 mm)
| Calibration Procedure: | In-House Procedure #00-5
| Calibration Equipment: | Mititoyo Digimatic Calipers
Condition
Date | Frame | Joints | Cloth | Tightness | Pass / Fail | Checked By | Initial by
2011-04-14 | New | New | New | New | Pass | MF | sgw
2011-09-20 | Excellent | Excellent | Excellent | Excellent | Pass | Haley M. | sgw
2012-09-04 | Excellent | Excellent | Excellent | Excellent | Pass | Thomas M. | sgw
2013-09-10 | Excellent | Excellent | Excellent | Excellent | Pass | Abiy G. | sgw
2014-09-19 | Excellent | Excellent | Excellent | Excellent | Pass | MF

## # 30 Sieve
Sieve Calibration Worksheet
| Equipment ID: | # 30 Sieve |  |  | Serial #: | 11147413
| Manufacturer: | Gilson |  |  | UofA ID: | NA
| Model #: | V8SF#030 |  |  | Location: | ST (218)
[5 empty] | Calibration Interval: | 12 months
| Calibration Item: | Verify physical condition of sieve and critical dimensions for
sieves larger than #4 (4.75 mm)
| Calibration Procedure: | In-House Procedure #00-5
| Calibration Equipment: | Mititoyo Digimatic Calipers
Condition
Date | Frame | Joints | Cloth | Tightness | Pass / Fail | Checked By | Initial by
2011-04-14 | New | New | New | New | Pass | MF | sgw
2011-09-20 | Excellent | Excellent | Excellent | Excellent | Pass | Haley M. | sgw
2012-09-04 | Excellent | Excellent | Excellent | Excellent | Pass | Thomas M. | sgw
2013-09-10 | Excellent | Excellent | Excellent | Excellent | Pass | Abiy G. | sgw
2014-09-19 | Excellent | Excellent | Excellent | Excellent | Pass | MF

## # 40 Sieve
Sieve Calibration Worksheet
| Equipment ID: | # 40 Sieve |  |  | Serial #: | 11147414
| Manufacturer: | Gilson |  |  | UofA ID: | NA
| Model #: | V8SF# 040 |  |  | Location: | ST (218)
[5 empty] | Calibration Interval: | 12 months
| Calibration Item: | Verify physical condition of sieve and critical dimensions for
sieves larger than #4 (4.75 mm)
| Calibration Procedure: | In-House Procedure #00-5
| Calibration Equipment: | Mititoyo Digimatic Calipers
Condition
Date | Frame | Joints | Cloth | Tightness | Pass / Fail | Checked By | Initial by
2011-04-14 | New | New | New | New | Pass | MF | sgw
2011-09-20 | Excellent | Excellent | Excellent | Excellent | Pass | Haley M. | sgw
2012-09-04 | Excellent | Excellent | Excellent | Excellent | Pass | Thomas M. | sgw
2013-09-10 | Excellent | Excellent | Excellent | Excellent | Pass | Abiy G. | sgw
2014-09-19 | Excellent | Excellent | Excellent | Excellent | Pass | MF

## # 50 Sieve
Sieve Calibration Worksheet
| Equipment ID: | # 50 Sieve |  |  | Serial #: | 11147415
| Manufacturer: | Gilson |  |  | UofA ID: | NA
| Model #: | V8SF# 050 |  |  | Location: | ST (218)
[5 empty] | Calibration Interval: | 12 months
| Calibration Item: | Verify physical condition of sieve and critical dimensions for
sieves larger than #4 (4.75 mm)
| Calibration Procedure: | In-House Procedure #00-5
| Calibration Equipment: | Mititoyo Digimatic Calipers
Condition
Date | Frame | Joints | Cloth | Tightness | Pass / Fail | Checked By | Initial by
2011-04-14 | New | New | New | New | Pass | MF | sgw
2011-09-20 | Excellent | Excellent | Excellent | Excellent | Pass | Haley M. | sgw
2012-09-04 | Excellent | Excellent | Excellent | Excellent | Pass | Thomas M. | sgw
2013-09-10 | Excellent | Excellent | Excellent | Excellent | Pass | Abiy G. | sgw
2014-09-19 | Excellent | Excellent | Excellent | Excellent | Pass | MF

## # 100 Sieve
Sieve Calibration Worksheet
| Equipment ID: | # 100 Sieve |  |  | Serial #: | 11147416
| Manufacturer: | Gilson |  |  | UofA ID: | NA
| Model #: | V8SF#100 |  |  | Location: | ST (218)
[5 empty] | Calibration Interval: | 12 months
| Calibration Item: | Verify physical condition of sieve and critical dimensions for
sieves larger than #4 (4.75 mm)
| Calibration Procedure: | In-House Procedure #00-5
| Calibration Equipment: | Mititoyo Digimatic Calipers
Condition
Date | Frame | Joints | Cloth | Tightness | Pass / Fail | Checked By | Initial by
2011-04-14 | New | New | New | New | Pass | MF | sgw
2011-09-20 | Excellent | Excellent | Excellent | Excellent | Pass | Haley M. | sgw
2012-09-04 | Excellent | Excellent | Excellent | Excellent | Pass | Thomas M. | sgw
2013-09-10 | Excellent | Excellent | Excellent | Excellent | Pass | Abiy G. | sgw
2014-09-19 | Excellent | Excellent | Excellent | Excellent | Pass | MF

## # 200 Sieve
Sieve Calibration Worksheet
| Equipment ID: | # 200 Sieve |  |  | Serial #: | 11147417
| Manufacturer: | Gilson |  |  | UofA ID: | NA
| Model #: | V8SF#200 |  |  | Location: | ST (218)
[5 empty] | Calibration Interval: | 12 months
| Calibration Item: | Verify physical condition of sieve and critical dimensions for
sieves larger than #4 (4.75 mm)
| Calibration Procedure: | In-House Procedure #00-5
| Calibration Equipment: | Mititoyo Digimatic Calipers
Condition
Date | Frame | Joints | Cloth | Tightness | Pass / Fail | Checked By | Initial by
2011-04-14 | New | New | New | New | Pass | MF | sgw
2011-09-20 | Excellent | Excellent | Excellent | Excellent | Pass | Haley M. | sgw
2012-09-04 | Excellent | Excellent | Excellent | Excellent | Pass | Thomas M. | sgw
2013-09-10 | Excellent | Excellent | Excellent | Excellent | Pass | Abiy G. | sgw
2014-09-19 | Excellent | Excellent | Excellent | Excellent | Pass | MF

## # 200 Wash Sieve
Sieve Calibration Worksheet
| Equipment ID: | # 200 Wash Sieve |  |  | Serial #: | 5018767
| Manufacturer: | Gilson |  |  | UofA ID: | NA
| Model #: | V8 SF # 200 |  |  | Location: | ST (218)
[5 empty] | Calibration Interval: | 6 months
| Calibration Item: | Verify physical condition of sieve and critical dimensions for
sieves larger than #4 (4.75 mm)
| Calibration Procedure: | In-House Procedure #00-5
| Calibration Equipment: | Mititoyo Digimatic Calipers
Condition
Date | Frame | Joints | Cloth | Tightness | Pass / Fail | Checked By | Initial by
2007-03-22 | New | New | New | New | Pass | SGW | sgw
2007-09-20 | Good | Good | Good | Good | Pass | AM | sgw
2008-03-19 | Good | Good | Good | Good | Pass | AM | sgw
2008-09-10 | OK | OK | OK | OK | Pass | BH | sgw
2009-03-27 | OK | OK | OK | OK | Pass | MF | sgw
2009-09-20 | OK | OK | OK | OK | Pass | MF | sgw
2010-03-15 | OK | OK | OK | OK | Pass | MF | sgw
2010-09-03 | OK | OK | OK | OK | Pass | MF | sgw
2011-03-01 | OK | OK | Small Marks | OK | Pass | MF | sgw
2011-09-20 | OK | OK | Small Marks | OK | Pass | Haley M. | sgw
2012-09-04 | OK | OK | Small Marks | OK | Pass | Thomas M. | sgw
2013-09-10 | OK | OK | Small Marks | OK | Pass | MF | sgw
2014-09-19 | OK | OK | Small Marks | OK | Pass | MF

## Fine Sieve Calibration
Sieve Calibration Worksheet
| Equipment ID: | Fine Sieves |  |  | Serial #:
| Manufacturer: | Gilson |  |  | UofA ID: | NA
| Model #: | V8CF - |  |  | Location: | ST (218)
[5 empty] | Calibration Interval: | 12 months
| Calibration Item: | Verify physical condition of sieve and critical dimensions for
sieves larger than #4 (4.75 mm)
| Calibration Procedure: | In-House Procedure #00-5
| Calibration Equipment: | Mititoyo Digimatic Calipers
March 2010
Condition
Sieve | Frame | Joints | Cloth | Tightness | Comments |  | Initial by
# 8
# 10
# 16
# 30
# 40
# 50
# 100
# 200
# 200 W
The fine sieves are logged on individual sheets in the equipment calibration manual.