## Liquid Limit Device
Liquid Limit Device and Grooving Tool Calibration Worksheet
|  | Equipment ID: | Liquid Limit Device |  | Date: | 3/5/14
|  | Manufacturer: | ELE / Soiltest |  | Performed by: | M. Fleck
|  | Model #: | CL - 206
|  | Serial #: | NA |  | Next Calibration Due: | March 2015
|  | Location: | ST (218) |  | Last Calibration: | March 2013
| Calibration Item: | Verify critical dimensions of the Liquid Limit Device, Grooving Tool, and Gage
| Calibration Procedure: | In-House Procedure #00-12
| Calibration Equipment: | Mititoyo Digimatic Calipers, balance read. to 0.1 g, and ruler read. to 1 mm.
|  | Thickness of Cup: | 1.98 mm | Tolerance = 2 +/- 0.1 mm
|  | Depth of Cup: | 27.24 mm | Tolerance = 27 +/- 1 mm
|  | Cup to Base (at cam follower): | 47.04 mm | Tolerance = 47 +/- 1.5 mm
|  | Thickness of Base: | 50.43 mm | Tolerance = 50 +/- 5 mm
|  | Length of Base: | 149.90 mm | Tolerance = 150 +/- 5 mm
|  | Width of Base: | 125.02 mm | Tolerance = 125 +/- 5 mm
|  | Grooving Tool Thickness: | 9.95 mm | Tolerance = 10 +/- 0.1 mm
|  | Width of Cutting Edge of Tool: | 2.04 mm | Tolerance = 2 +/- 0.1 mm
|  | Width of Grooving Tool: | 13.42 mm | Tolerance = 13.5 +/- 0.1 mm
|  | Depth of Gauge: | 9.96 mm | Tolerance = 10 +/- 0.2 mm
|  | Width of Gauge: | 9.91 mm | Tolerance = 10 +/- 0.2 mm
|  | Length of Gauge: | 15.89 mm | Tolerance = At least 15.9 mm
|  | Height Gauge Block: | 9.98 mm |  |  | Pass/Fail: | Pass
Comments:
Initial By:

## Standard Hammer
Proctor Hammer Calibration Worksheet
|  | Equipment ID: | Stand. Proctor Hammer |  |  | Date: | 3/4/14
|  | Manufacturer: | ELE / Soiltest |  |  | Performed by: | M. Fleck
|  | Model #: | CN - 415
|  | Serial #: | NA |  |  | Next Calibration Due: | March 2015
|  | UofA ID: | NA |  |  | Last Calibration: | March 2013
|  | Location: | ST (218)
| Calibration Item: | Verify critical dimensions of the proctor hammer
| Calibration Procedure: | In-House Procedure #00-6
| Calibration Equipment: | Mititoyo Digimatic Calipers, tape measure, and Scale CL1.
|  | Diameter of Rammer Face: | 50.83 mm | 50.73 mm | Tolerance = 50.80 +/- 0.25 mm
|  | Drop Height of Hammer: | 305 mm |  | Tolerance = 305 +/- 2 mm
|  | Mass of Hammer: | 2.491 kg |  | Tolerance = 2.495 +/- 0.009 kg
|  | Vent Hole Diameters: | 9.41 mm | 9.43 mm | Tolerance = approx. 3/8 inch
|  |  | 9.46 mm | 9.41 mm
|  |  | 9.43 mm | 9.45 mm
|  |  | 9.41 mm | 9.44 mm
[7 empty] | Pass/Fail: | Pass
Comments:
Initial By:

## Modified Hammer
Modified Proctor Hammer Calibration Worksheet
|  | Equipment ID: | Mod. Proctor Hammer |  |  | Date: | 3/4/14
|  | Manufacturer: | ELE / Soiltest |  |  | Performed by: | M. Fleck
|  | Model #: | CN - 416
|  | Serial #: | NA |  |  | Next Calibration Due: | March 2015
|  | UofA ID: | NA |  |  | Last Calibration: | March 2013
|  | Location: | ST (218)
| Calibration Item: | Verify critical dimensions of the modified proctor hammer
| Calibration Procedure: | In-House Procedure #00-6
| Calibration Equipment: | Mititoyo Digimatic Calipers, tape measure, and Scale CL1.
|  | Diameter of Rammer Face: | 50.83 mm | 50.77 mm | Tolerance = 50.80 +/- 0.25 mm
|  | Drop Height of Hammer: | 457 mm |  | Tolerance = 457 +/- 2 mm
|  | Mass of Hammer: | 4.544 kg |  | Tolerance = 4.536 +/- 0.009 kg
|  | Vent Hole Diameters: | 9.53 mm | 9.51 mm | Tolerance = approx. 3/8 inch
|  |  | 9.53 mm | 9.54 mm
|  |  | 9.51 mm | 9.52 mm
|  |  | 9.51 mm | 9.54 mm
[7 empty] | Pass/Fail: | Pass
Comments:
Initial By:

## 4" Mold Calibration
4" Proctor Mold Calibration Worksheet
|  | Equipment ID: | 4" Proctor Mold |  |  | Date: | 3/5/14
|  | Manufacturer: | Gilson |  |  | Performed by: | M. Fleck
|  | Model #: | HMA - 111
|  | Serial #: | NA |  |  | Next Calibration Due: | March 2015
|  | UofA ID: | NA |  |  | Last Calibration: | March 2013
|  | Location: | ST (218)
| Calibration Item: | Verify critical dimensions of the proctor mold
| Calibration Procedure: | In-House Procedure #00-7
| Calibration Equipment: | Mititoyo Digimatic Calipers
|  | Top Inside Diameter of Mold: | 101.88 mm | 101.89 mm | Tolerance = 101.60 +/- 0.41 mm
|  | Bottom Inside Diameter of Mold: | 101.88 mm | 101.78 mm | Tolerance = 101.60 +/- 0.41 mm
|  | Height of Mold: | 116.53 mm | 116.52 mm | Tolerance = 116.43 +/- 0.13 mm
|  | Height of Collar: | 60.40 mm | 60.32 mm | Tolerance = approx. 60 mm
|  | Inside Diameter of Collar: | 101.53 mm | 101.42 mm | Tolerance = 101.60 +/- 0.41 mm
|  | Base Plate Planeness: | < 0.0015" |  | Tolerance = 0.005 in.
(maximum deviation)
[7 empty] | Pass/Fail: | Pass
Comments:
Initial By:

## 6" Mold Calibration
6" Proctor Mold Calibration Worksheet
|  | Equipment ID: | 6" Proctor Mold |  |  | Date: | 3/5/14
|  | Manufacturer: | ELE |  |  | Performed by: | M. Fleck
|  | Model #: | CN - 404
|  | Serial #: | NA |  |  | Next Calibration Due: | March 2015
|  | UofA ID: | NA |  |  | Last Calibration: | March 2013
|  | Location: | ST (218)
| Calibration Item: | Verify critical dimensions of the proctor mold
| Calibration Procedure: | In-House Procedure #00-7
| Calibration Equipment: | Mititoyo Digimatic Calipers
|  | Top Inside Diameter of Mold: | 152.50 mm | 152.60 mm | Tolerance = 152.40 +/- 0.66 mm
|  | Bottom Inside Diameter of Mold: | 152.59 mm | 152.52 mm | Tolerance = 152.40 +/- 0.66 mm
|  | Height of Mold: | 116.34 mm | 116.40 mm | Tolerance = 116.43 +/- 0.13 mm
|  | Height of Collar: | 60.14 mm | 60.06 mm | Tolerance = approx. 60 mm
|  | Inside Diameter of Collar: | 152.50 mm | 152.58 mm | Tolerance = 152.40 +/- 0.66 mm
|  | Base Plate Planeness: | < 0.0015" |  | Tolerance = 0.005 in.
(maximum deviation)
[7 empty] | Pass/Fail: | Pass
Comments:
Initial By:

## Pycnometer Blank
| Date: |  |  | Performed by:
| Next Calibration Due: |  |  | Last Calibration:
Pycnometer # | 1 |  |  | Equipment ID:
Manufacturer:
Pyc | Pyc + Water | Temp, F | Water | Model #:
Serial #:
UofA ID:
Location:
Procedure:
Pycnometer # | G |  |  | Equipment ID:
Manufacturer:
Pyc | Pyc + Water | Temp, F | Water | Model #:
Serial #:
UofA ID:
Location:
Procedure:

## Pycnometer
| Date: |  |  | Performed by: | Mary Fleck
| Next Calibration Due: | April 2013 |  | Last Calibration: | April 2011
Pycnometer # | 1 |  |  | Equipment ID: | 500 mL Pyc. #1
|  |  |  | Manufacturer: | Kimax
Pyc | Pyc + Water | Temp, F | Water | Model #: | NA
|  |  | 0 | Serial #: | NA
0 |  |  | 0 | UofA ID: | NA
0 |  |  | 0 | Location: | ST (218)
0 |  |  | 0 | Procedure: | AASHTO T-100 (P7)
Pycnometer # | G |  |  | Equipment ID: | 500 mL Pyc. #2 (G)
|  |  |  | Manufacturer: | Kimax
Pyc | Pyc + Water | Temp, F | Water | Model #: | NA
|  |  | 0 | Serial #: | NA
0 |  |  | 0 | UofA ID: | NA
0 |  |  | 0 | Location: | ST (216)
0 |  |  | 0 | Procedure: | AAHSTO T-100 (P7)

## Pycnometer Statement
| Date: | 2012-04-12 |  | Performed by: | Mary Fleck
| Next Calibration Due: | April 2013 |  | Last Calibration: | April 2011
Pycnometer # | 1 |  |  | Equipment ID: | 500 mL Pyc. #1
|  |  |  | Manufacturer: | Kimax
Pyc | Pyc + Water | Temp, F | Water | Model #: | NA
|  |  |  | Serial #: | NA
|  |  |  | UofA ID: | NA
|  |  |  | Location: | ST (218)
|  |  |  | Procedure: | AASHTO T-100 (P7)
This calibration is no longer being performed. The calibration of the pycnometer is performed
during each test by bringing the pyc + water to the same test temperature as the
pyc + water + sample before obtaining it's weight. This eliminates the need for this procedure.
Initial By:
Pycnometer # | G |  |  | Equipment ID: | 500 mL Pyc. #2 (G)
|  |  |  | Manufacturer: | Kimax
Pyc | Pyc + Water | Temp, F | Water | Model #: | NA
|  |  |  | Serial #: | NA
|  |  |  | UofA ID: | NA
|  |  |  | Location: | ST (216)
|  |  |  | Procedure: | AAHSTO T-100 (P7)
This calibration is no longer being performed. The calibration of the pycnometer is performed
during each test by bringing the pyc + water to the same test temperature as the
pyc + water + sample before obtaining it's weight. This eliminates the need for this procedure.
Initial By: