## Coarse Agg. Splitter
Mechanical Splitter Calibration Worksheet
|  | Equipment ID: | Coarse Agg. Splitter |  | Date: | 3/5/15
|  | Manufacturer: | Soiltest |  | Performed by: | Mary Fleck
|  | Model #: | CL - 262
|  | Serial #: | NA |  | Next Calibration Due: | March 2016
|  | UofA ID: | NA |  | Last Calibration: | March 2014
|  | Location: | AGG (233)
| Calibration Item: | Verify critical dimensions of the mechanical splitter
| Calibration Procedure: | In-House Procedure #00-11
| Calibration Equipment: | Ruler, readable to 1 mm.
|  | Number of Chutes: | 12 | Tolerance = even number
at least 8 for coarse, 12 for fine
|  | Chute Opening: | 50 mm | Tolerance = approx. 50% larger than
|  |  | 51 mm |  | largest aggregate size
50 mm
50 mm
50 mm
49 mm
50 mm
50 mm
50 mm
51 mm
50 mm
50 mm
[6 empty] | Pass/Fail: | Pass
Comments:
Initial By:

## Fine Agg. Splitter
Mechanical Splitter Calibration Worksheet
|  | Equipment ID: | Fine Agg. Splitter |  | Date: | 3/5/15
|  | Manufacturer: | Soiltest |  | Performed by: | Mary Fleck
|  | Model #: | CL - 283
|  | Serial #: | NA |  | Next Calibration Due: | March 2016
|  | UofA ID: | NA |  | Last Calibration: | March 2014
|  | Location: | AGG (233)
| Calibration Item: | Verify critical dimensions of the mechanical splitter
| Calibration Procedure: | In-House Procedure #00-11
| Calibration Equipment: | Ruler, readable to 1 mm.
|  | Number of Chutes: | 14 | Tolerance = even number
at least 8 for coarse, 12 for fine
|  | Chute Opening: | 19 mm | Tolerance = approx. 50% larger than
|  |  | 19 mm |  | largest aggregate size
19 mm
19 mm
19 mm
19 mm
19 mm
19 mm
19 mm
19 mm
19 mm
19 mm
19 mm
19 mm
[6 empty] | Pass/Fail: | Pass
Comments:
Initial By:

## Sieve Shaker - RoTap 1
Sieve Shaker Calibration Worksheet
|  | Equipment ID: | Ro - Tap 1 |  | Date: | 3/19/15
|  | Manufacturer: | Gilson |  | Performed by: | Brendan Oman
|  | Model #: | RX - 29
|  | Serial #: | 12583 |  | Next Calibration Due: | March 2016
|  | UofA ID: | 238798 |  | Last Calibration: | March 2014
|  | Location: | SAW (235)
| Calibration Item: | Verify sieving thoroughness of sieve shakers
| Calibration Procedure: | In-House Procedure #00-10
| Calibration Equipment: | Sieves, Scale CL1, pan
|  | Original Mass of Sample: | 1847.2
After Hand Sieving:
|  |  |  | Mass Passing | Percent Passing
|  | 3/4" | 0.2 |  | 0.010827
|  | 1/2" | 6.1 |  | 0.33023
|  | #4 | 2 |  | 0.10827
|  | #8 | 2.8 |  | 0.15158
|  | #16 | 0.7 |  | 0.037895
|  | #50 | 0.3 |  | 0.016241
|  | #100 | 0.2 |  | 0.010827
| Tolerance: | No sieve shall have greater than 0.5% passing by total mass of sample.
[6 empty] | Pass/Fail: | Pass
| Comments: | Shake Time = 7 minutes
Initial By:
Maintenance (see mfg. instructions):
√ | Check for cleanliness?
√ | Plug not excessively worn?
√ | Greased if needed?
| Comments: | No maintenance required. M. Fleck.
Initial By:

## Sieve Shaker - RoTap 2
Sieve Shaker Calibration Worksheet
|  | Equipment ID: | Ro - Tap 2 |  | Date: | 3/19/15
|  | Manufacturer: | Gilson |  | Performed by: | Brendan Oman
|  | Model #: | RX - 29
|  | Serial #: | 21741 |  | Next Calibration Due: | March 2016
|  | UofA ID: | NA |  | Last Calibration: | March 2014
|  | Location: | SAW (235)
| Calibration Item: | Verify sieving thoroughness of sieve shakers
| Calibration Procedure: | In-House Procedure #00-10
| Calibration Equipment: | Sieves, Scale CL1, pan
|  | Original Mass of Sample: | 1846.8
After Hand Sieving:
|  |  |  | Mass Passing | Percent Passing
|  | 3/4" | 0.1 |  | 0.0054148
|  | 1/2" | 5.6 |  | 0.30323
|  | #4 | 1.5 |  | 0.081222
|  | #8 | 1.4 |  | 0.075807
|  | #16 | 0.5 |  | 0.027074
|  | #50 | 0.2 |  | 0.01083
|  | #100 | 0.2 |  | 0.01083
| Tolerance: | No sieve shall have greater than 0.5% passing by total mass of sample.
[6 empty] | Pass/Fail: | Pass
| Comments: | Shake Time = 7 minutes
Initial By:
Maintenance (see mfg. instructions):
√ | Check for cleanliness?
√ | Plug not excessively worn?
√ | Greased if needed?
| Comments: | No maintenenace required. M. Fleck
Initial By:

## Sieve Shaker - MaryAnn 1
Sieve Shaker Calibration Worksheet
|  | Equipment ID: | Mary Ann 1 |  | Date: | 3/19/15
|  | Manufacturer: | Rainhart |  | Performed by: | Brendan Oman
|  | Model #: | 637
|  | Serial #: | NA |  | Next Calibration Due: | March 2016
|  | UofA ID: | NA |  | Last Calibration: | March 2014
|  | Location: | SAW (235)
| Calibration Item: | Verify sieving thoroughness of sieve shakers
| Calibration Procedure: | In-House Procedure #00-10
| Calibration Equipment: | Sieves, Scale CL1, pan
|  | Original Mass of Sample: | 1774.2
After Hand Sieving:
|  |  |  | Mass Passing | Percent Passing
|  | 3/4" | 0 |  | 0
|  | 1/2" | 0 |  | 0
|  | #4 | 0.5 |  | 0.028182
|  | #8 | 1.2 |  | 0.067636
|  | #16 | 1.1 |  | 0.062
|  | #50 | 0.2 |  | 0.011273
|  | #100 | 0.4 |  | 0.022545
| Tolerance: | No sieve shall have greater than 0.5% passing by total mass of sample.
[6 empty] | Pass/Fail: | Pass
| Comments: | Shake Time = 7 minutes
Initial By:
Maintenance (see mfg. instructions):
√ | Check for cleanliness?
√ | Paddles do not show excessive wear?
√ | Greased if needed?
| Comments: | No maintenance required.
Initial By:

## Sieve Shaker - MaryAnn 2
Sieve Shaker Calibration Worksheet
|  | Equipment ID: | Mary Ann 2 |  | Date: | 3/19/15
|  | Manufacturer: | Rainhart |  | Performed by: | Brendan Oman
|  | Model #: | 637
|  | Serial #: | NA |  | Next Calibration Due: | March 2016
|  | UofA ID: | NA |  | Last Calibration: | March 2014
|  | Location: | SAW (235)
| Calibration Item: | Verify sieving thoroughness of sieve shakers
| Calibration Procedure: | In-House Procedure #00-10
| Calibration Equipment: | Sieves, Scale CL1, pan
|  | Original Mass of Sample: | 1849.5
After Hand Sieving:
|  |  |  | Mass Passing | Percent Passing
|  | 3/4" | 0 |  | 0
|  | 1/2" | 3.5 |  | 0.18924
|  | #4 | 0.7 |  | 0.037848
|  | #8 | 1.4 |  | 0.075696
|  | #16 | 0.8 |  | 0.043255
|  | #50 | 0.3 |  | 0.016221
|  | #100 | 0.3 |  | 0.016221
| Tolerance: | No sieve shall have greater than 0.5% passing by total mass of sample.
[6 empty] | Pass/Fail: | Pass
| Comments: | Shake Time = 7 minutes
Initial By:
Maintenance (see mfg. instructions):
√ | Check for cleanliness?
√ | Paddles do not show excessive wear?
√ | Greased if needed?
| Comments: | No maintenance required
Initial By:

## Cone and Tamper Calibration
Cone and Tamper Calibration Worksheet
|  | Equipment ID: | Cone & Tamper |  |  | Date: | 3/2/15
|  | Manufacturer: | ELE |  |  | Performed by: | M. Fleck
|  | Model #: | G - 325
|  | Serial #: | NA |  |  | Next Calibration Due: | March 2016
|  | UofA ID: | NA |  |  | Last Calibration: | March 2014
|  | Location: | ST (218)
| Calibration Item: | Verify critical dimensions of the cone and tamper
| Calibration Procedure: | In-House Procedure #00-8
| Calibration Equipment: | Mititoyo Digimatic Calipers, Scale CL1, and ruler read. to 1 mm.
|  | Top Inside Diameter of Cone: | 41.26 mm | 41.23 mm | Tolerance = 40 +/- 3 mm
|  | Bottom Inside Diameter of Cone: | 89.97 mm | 90.30 mm | Tolerance = 90 +/- 3 mm
|  | Height of Cone: | 76.69 mm |  | Tolerance = 75 +/- 3 mm
|  | Thickness of Cone: | 1.43 mm | 1.30 mm | Tolerance = at least 0.8 mm
|  |  | 1.38 mm | 1.20 mm
|  | Diameter of tamper face: | 25.30 mm | 25.33 mm | Tolerance = 25 +/- 3 mm
|  | Mass of tamper: | 345.3 g |  | Tolerance = 340 +/- 15 g
[7 empty] | Pass/Fail: | Pass
Comments:
Initial By:

## Fine Agg. Angularity
Fine Aggregate Angularity Calibration Worksheet
|  | Equipment ID: | FAA # 2 |  | Date: | 3/5/15
|  | Manufacturer: | Gilson |  | Performed by: | M. Fleck
|  | Model #: | H - 1680
|  | Serial #: | NA |  | Next Calibration Due: | March 2016
|  | UofA ID: | NA |  | Last Calibration: | March 2014
|  | Location: | ST (218)
| Calibration Item: | Verify critical dimensions of the fine aggregate angularity device
| Calibration Procedure: | In-House Procedure #00-13
| Calibration Equipment: | Mititoyo Digimatic Calipers, balance read. to 0.1 g, and ruler read. to 1 mm.
|  | Inside Diameter of Cylinder: | 39.38 mm | Tolerance = approx. 39.02 mm
|  | Inside Height of Cylinder: | 82.76 mm | Tolerance = approx. 86 mm
|  | Thickness of Cylinder Bottom: | 6.33 mm | Tolerance = at least 6 mm
|  | Inside Diameter of Funnel: | 12.31 mm | Tolerance = 12.7 +/- 0.6 mm
|  | Height of Funnel: | 50.57 mm | Tolerance = at least 38 mm
|  | Distance from Funnel to Cylinder: | 115 mm | Tolerance = 115 +/- 2 mm
|  | Width of Glass Plate: | 60.06 mm | Tolerance = approx. 60 mm
|  | Length of Glass Plate: | 59.86 mm | Tolerance = approx. 60 mm
|  | Thickness of Glass Plate: | 4.75 mm | Tolerance = at least 4 mm
[6 empty] | Pass/Fail: | Pass
Comments:
Initial By: