## Mitutoyo
Digital Caliper Calibration
| Equipment ID: | Digital Caliper |  |  | Date: | Yearly
| Manufacturer: | Mitutoyo |  |  | Performed by:
| Model #: | 500 - 196
| Serial #: | 528605 |  |  | Tolerance : | ± 0.01 mm
| UofA ID: | N / A |  |  |  | ± 0.0039 in.
| Location: | SR (218)
| Calibration Item: | Verify correct measurements
| Calibration Procedure: | In-House Procedure #01-17
| Calibration Equipment: | Gage blocks included in gyratory compactor calibration kits
Date: | Reading @1.000 in. | Reading @ 2.000 in. | Reading @ 3.000 in. | Reading @ 114.38 mm | Pass/Fail: | Next Calib. Due: | Performed By:
2001-10-25 | 1 | 1.9995 | 3.0005 | 114.35 | Pass | Oct. - 02 | sgw
2002-10-01 | 1 | 2 | 3 | 114.38 | Pass | Oct. - 03 | sgw
2003-10-21 | 1 | 2 | 3 | 114.38 | Pass | Oct. - 04 | sgw
2004-10-19 | 1 | 2 | 3.0002 | 114.38 | Pass | Oct. - 05 | sgw
2005-10-17 | 1 | 2 | 3.0001 | 114.38 | Pass | Oct. - 06 | sgw
2006-10-02 | 1 | 2 | 3 | 114.38 | Pass | Oct. - 07 | sgw
2007-10-18 | 1 | 2 | 3 | 114.36 | Pass | Oct. - 08 | sgw
2008-10-30 | 1 | 2.0005 | 3.0015 | 114.37 | Pass | Oct. - 09 | an
2009-10-01 | 1 | 2 | 3 | 114.37 | Pass | Oct. - 10 | mf
2010-10-05 | 0.9995 | 1.9995 | 3 | 114.36 | Pass | Oct. - 11 | mf
2011-10-13 | 1 | 2 | 3 | 114.36 | Pass | Oct. - 12 | hm
2012-10-04 | 0.9995 | 2 | 3 | 114.38 | Pass | Oct. - 13 | mf
2013-10-04 | 1 | 2 | 3 | 114.37 | Pass | Oct. - 14 | mf
2014-10-02 | 1 | 1.9995 | 3 | 114.36 | Pass | Oct. - 15 | mf
2015-10-12 | 1 | 2 | 2.9995 | 114.37 | Pass | Oct. - 16 | mf
Initial by :