## Oven AG2
Oven Calibration Worksheet
Equipment ID: | AG2 |  | Serial #: | 187253
Manufacturer: | Despatch |  | UofA ID: | NA
Model #: | LBB1-69A-1 |  | Location: |  | Agg Lab (233)
Calibration Item: | Verify temperature readings |  |  | Typical Set Temperature: | 230 ° F
Calibration Procedure: | In-House Procedure #00-1 |  |  | Tolerance : | ± 9 °
Calibration Equipment: | Thermometer ASTM | 67 F
Temp. Readings
Date: | 2015-10-12 | Set Temp. | 1 | 2 | 3
Performed by: | M. Fleck | 236 | 226.5 | 230.3 | 228.7
Next Calibration Due: | Oct. 2016
Last Calibration: | Oct. 2014
Requirement: | ± 9 ° F
| Pass/Fail: | Pass |  |  | Initial by: | sgw
| Comments: | Oven set point at 236°F. (Tag on oven)
Constant temperature of 230°F.
Temp. Readings
Date: |  | Set Temp. | 1 | 2 | 3
Performed by: |  | 236
Next Calibration Due: | Oct. 2017
Last Calibration: | Oct. 2015
Requirement: | ± 9 ° F
| Pass/Fail: |  |  |  | Initial by:
| Comments: | Oven set point at 244°F.
Constant temperature of 230°F.
Temp. Readings
Date: |  | Set Temp. | 1 | 2 | 3
Performed by: |  | 236
Next Calibration Due: | Oct. 2018
Last Calibration: | Oct. 2016
Requirement: | ± 9 ° F
| Pass/Fail: |  |  |  | Initial by:
| Comments: | Oven set point at 244°F.
Constant temperature of 230°F.

## Oven HM1
Oven Calibration Worksheet
Equipment ID: | HM - 1 |  | Serial #: | 170436
Manufacturer: | Despatch |  | UofA ID: | 255801
Model #: | LBB2 - 18 - 1 |  | Location: |  | Hot Mix Lab (225)
Calibration Item: | Verify temperature readings |  |  | Typical Set Temperature: | 300 ° F
Calibration Procedure: | In-House Procedure #00-1 |  |  | Tolerance : | ± 5 ° F
Calibration Equipment: | Thermometer ASTM | 67F / 68F
Temp. Readings
Date: | 2015-10-13 | Set Temp. | 1 | 2 | 3
Performed by: | M/ Fleck | 270 | 268.1 | 269.3 | 269.5
Next Calibration Due: | October 2016 | 300 | 297.3 | 297.9 | 296.9
Last Calibration: | October 2014 | 328 | 324.7 | 326.1 | 326.9
Requirement: | ± 5 ° F
| Pass/Fail: | Pass |  |  | Initial by:
| Comments: | Oven setpoint @ 3° higher. (Tag on oven)
Variable Temperatures (250° - 340° F)
Temp. Readings
Date: |  | Set Temp. | 1 | 2 | 3
Performed by: |  | 230
Next Calibration Due: | October 2017 | 280
Last Calibration: | October 2015 | 315
Requirement: | ± 5 ° F
| Pass/Fail: |  |  |  | Initial by:
| Comments: | Oven setpoint @ 3° higher. (Tag on oven)
Variable Temperatures (250° - 340° F)
Temp. Readings
Date: |  | Set Temp. | 1 | 2 | 3
Performed by: |  | 250
Next Calibration Due: | October 2018 | 290
Last Calibration: | October 2016 | 325
Requirement: | ± 5 ° F
| Pass/Fail: |  |  |  | Initial by:
| Comments: | Oven setpoint @ 3° higher. (Tag on oven)
Variable Temperatures (250° - 340° F)

## Oven MX5
Oven Calibration Worksheet
Equipment ID: | MX - 5 |  | Serial #: | 9040305
Manufacturer: | Sheldon Mfg., Inc. |  | UofA ID: | 259372
Model #: | 1685 |  | Location: |  | Mixing Lab (234)
Calibration Item: | Verify temperature readings |  |  | Typical Set Temperature: | 330 ° F
Calibration Procedure: | In-House Procedure #00-1 |  |  | Tolerance : | ± 5 ° F
Calibration Equipment: | Thermometer ASTM | 68 F
Temp. Readings
Date: | 2014-10-03 | Set Temp. | 1 | 2 | 3
Performed by: | M. Fleck | 250 | 249.6 | 249.7 | 249.1
Next Calibration Due: | October 2015 | 300 | 294.9 | 294.7 | 294.7
Last Calibration: | October 2013 | 350 | 342.1 | 342.5 | 342.7
Requirement: | ± 5 ° F
| Pass/Fail: | Pass |  |  | Initial by: | sgw
| Comments: | Setpoint tag on oven.
Variable Temperatures (200° - 350° F)
Temp. Readings
Date: | 2015-10-13 | Set Temp. | 1 | 2 | 3
Performed by: | M. Fleck | 200 | 203.2 | 202.5 | 203.1
Next Calibration Due: | October 2016 | 290 | 288.4 | 288.3 | 288.9
Last Calibration: | October 2014 | 300 | 297 | 296 | 297.3
Requirement: | ± 5 ° F | 345 | 340.1 | 339.5 | 340
| Pass/Fail: | Pass |  |  | Initial by:
| Comments: | Setpoint tag on oven.
Variable Temperatures (200° - 350° F)
Temp. Readings
Date: |  | Set Temp. | 1 | 2 | 3
Performed by: |  | 240
Next Calibration Due: | October 2017 | 290
Last Calibration: | October 2015 | 340
Requirement: | ± 5 ° F
| Pass/Fail: |  |  |  | Initial by:
| Comments: | Setpoint tag on oven.
Variable Temperatures (200° - 350° F)