## Liquid Limit Device
Liquid Limit Device and Grooving Tool Calibration Worksheet
|  | Equipment ID: | Liquid Limit Device |  | Date: | 3/2/16
|  | Manufacturer: | ELE / Soiltest |  | Performed by: | M. Fleck
|  | Model #: | CL - 206
|  | Serial #: | NA |  | Next Calibration Due: | March 2017
|  | Location: | ST (218) |  | Last Calibration: | March 2015
| Calibration Item: | Verify critical dimensions of the Liquid Limit Device, Grooving Tool, and Gage
| Calibration Procedure: | In-House Procedure #00-12
| Calibration Equipment: | Mititoyo Digimatic Calipers, balance read. to 0.1 g, and ruler read. to 1 mm.
|  | Thickness of Cup: | 2.01 mm | Tolerance = 2 +/- 0.1 mm
|  | Depth of Cup: | 26.88 mm | Tolerance = 27 +/- 1 mm
|  | Cup to Base (at cam follower): | 46.94 mm | Tolerance = 47 +/- 1.5 mm
|  | Thickness of Base: | 50.44 mm | Tolerance = 50 +/- 5 mm
|  | Length of Base: | 149.95 mm | Tolerance = 150 +/- 5 mm
|  | Width of Base: | 124.83 mm | Tolerance = 125 +/- 5 mm
|  | Grooving Tool Thickness: | 9.95 mm | Tolerance = 10 +/- 0.1 mm
|  | Width of Cutting Edge of Tool: | 2.05 mm | Tolerance = 2 +/- 0.1 mm
|  | Width of Grooving Tool: | 13.45 mm | Tolerance = 13.5 +/- 0.1 mm
|  | Depth of Gauge: | 9.97 mm | Tolerance = 10 +/- 0.2 mm
|  | Width of Gauge: | 9.94 mm | Tolerance = 10 +/- 0.2 mm
|  | Length of Gauge: | 15.87 mm | Tolerance = At least 15.9 mm
|  | Height Gauge Block: | 9.98 mm |  |  | Pass/Fail: | Pass
Comments:
Initial By:

## Standard Hammer
Proctor Hammer Calibration Worksheet
|  | Equipment ID: | Stand. Proctor Hammer |  |  | Date: | 3/2/16
|  | Manufacturer: | ELE / Soiltest |  |  | Performed by: | M. Fleck
|  | Model #: | CN - 415
|  | Serial #: | NA |  |  | Next Calibration Due: | March 2017
|  | UofA ID: | NA |  |  | Last Calibration: | March 2015
|  | Location: | ST (218)
| Calibration Item: | Verify critical dimensions of the proctor hammer
| Calibration Procedure: | In-House Procedure #00-6
| Calibration Equipment: | Mititoyo Digimatic Calipers, tape measure, and Scale CL1.
|  | Diameter of Rammer Face: | 50.63 mm | 50.65 mm | Tolerance = 50.80 +/- 0.25 mm
|  | Drop Height of Hammer: | 305 mm |  | Tolerance = 305 +/- 2 mm
|  | Mass of Hammer: | 2.490 kg |  | Tolerance = 2.495 +/- 0.009 kg
|  | Vent Hole Diameters: | 9.43 mm | 9.43 mm | Tolerance = approx. 3/8 inch
|  |  | 9.44 mm | 9.45 mm
|  |  | 9.44 mm | 9.45 mm
|  |  | 9.42 mm | 9.44 mm
[7 empty] | Pass/Fail: | Pass
Comments:
Initial By:

## Modified Hammer
Modified Proctor Hammer Calibration Worksheet
|  | Equipment ID: | Mod. Proctor Hammer |  |  | Date: | 3/2/16
|  | Manufacturer: | ELE / Soiltest |  |  | Performed by: | M. Fleck
|  | Model #: | CN - 416
|  | Serial #: | NA |  |  | Next Calibration Due: | March 2017
|  | UofA ID: | NA |  |  | Last Calibration: | March 2015
|  | Location: | ST (218)
| Calibration Item: | Verify critical dimensions of the modified proctor hammer
| Calibration Procedure: | In-House Procedure #00-6
| Calibration Equipment: | Mititoyo Digimatic Calipers, tape measure, and Scale CL1.
|  | Diameter of Rammer Face: | 50.83 mm | 50.76 mm | Tolerance = 50.80 +/- 0.25 mm
|  | Drop Height of Hammer: | 457 mm |  | Tolerance = 457 +/- 2 mm
|  | Mass of Hammer: | 4.544 kg |  | Tolerance = 4.536 +/- 0.009 kg
|  | Vent Hole Diameters: | 9.53 mm | 9.52 mm | Tolerance = approx. 3/8 inch
|  |  | 9.52 mm | 9.54 mm
|  |  | 9.54 mm | 9.54 mm
|  |  | 9.51 mm | 9.53 mm
[7 empty] | Pass/Fail: | Pass
Comments:
Initial By:

## 4" Mold Calibration
4" Proctor Mold Calibration Worksheet
|  | Equipment ID: | 4" Proctor Mold |  |  |  | Date: | 3/2/16
|  | Manufacturer: | ELE |  |  |  | Performed by: | M. Fleck
|  | Model #: | CN - 405
|  | Serial #: | NA |  |  |  | Next Calibration Due: | March 2017
|  | UofA ID: | NA |  |  |  | Last Calibration: | March 2015
|  | Location: | ST (218)
| Calibration Item: | Verify critical dimensions and volume of the proctor mold
| Calibration Procedure: | In-House Procedure #00-7
| Calibration Equipment: | Mititoyo Digimatic Calipers
|  | Top Inside Diameter of Mold: | 101.66 | 101.73 |  | Tolerance = 101.60 +/- 0.40 mm
|  | Bottom Inside Diameter of Mold: | 101.71 | 101.64 |  | Tolerance = 101.60 +/- 0.40 mm
|  | Height of Mold: | 116.42 | 116.41 |  | Tolerance = 116.40 +/- 0.50 mm
|  | Height of Collar: | 60.33 | 60.43 |  | Tolerance = approx. 60 mm
|  | Inside Diameter of Collar: | 101.53 | 101.52 |  | Tolerance = 101.60 +/- 0.40 mm
|  | Base Plate Planeness: | < 0.0015" |  |  | Tolerance = 0.005 in.
(maximum deviation)
AASHTO T 19 - Density of Water (Interpolated)
Mold Volume: |  |  |  |  | ° F |  | lb/ft³ |  | ° F | lb/ft³
[5 empty] | 65 |  | 62.336 |  | 74 | 62.269
|  | Empty Weight of Mold, Base, |  |  | 66 |  | 62.329 |  | 75 | 62.261
|  | and Glass Plate (g): | 4456.9 |  | 67 |  | 63.322 |  | 76 | 62.252
[5 empty] | 68 |  | 62.315 |  | 77 | 62.243
|  | Weight of Mold, Base, and |  |  | 69 |  | 62.308 |  | 78 | 62.234
|  | Glass Plate filled with Water (g): | 5396 |  | 70 |  | 62.301 |  | 79 | 62.225
[5 empty] | 71 |  | 62.293 |  | 80 | 62.216
|  | Temperature of Water (° F): | 71° F |  | 72 |  | 62.285 |  | 81 | 62.206
[5 empty] | 73 |  | 62.277 |  | 82 | 62.196
|  | Density of Water (lb/ft³): | 62.293
|  | Mold Volume (ft³): | 0.033235 |  |  |  |  | Pass/Fail: | Pass
|  |  | Tolerance = 0.0333 ± 0.0005 ft³ | (0.0328 - 0.0338)
Comments:
Initial By:

## 6" Mold Calibration
6" Proctor Mold Calibration Worksheet
|  | Equipment ID: | 6" Proctor Mold |  |  |  | Date: | 3/2/16
|  | Manufacturer: | ELE |  |  |  | Performed by: | M. Fleck
|  | Model #: | CN - 404
|  | Serial #: | NA |  |  |  | Next Calibration Due: | March 2017
|  | UofA ID: | NA |  |  |  | Last Calibration: | March 2015
|  | Location: | ST (218)
| Calibration Item: | Verify critical dimensions and volume of the proctor mold
| Calibration Procedure: | In-House Procedure #00-7
| Calibration Equipment: | Mititoyo Digimatic Calipers
|  | Top Inside Diameter of Mold: | 152.71 mm | 152.67 mm |  | Tolerance = 152.40 +/- 0.70 mm
|  | Bottom Inside Diameter of Mold: | 152.68 mm | 152.53 mm |  | Tolerance = 152.40 +/- 0.70 mm
|  | Height of Mold: | 116.43 mm | 116.43 mm |  | Tolerance = 116.40 +/- 0.50 mm
|  | Height of Collar: | 60.15 mm | 59.99 mm |  | Tolerance = approx. 60 mm
|  | Inside Diameter of Collar: | 152.55 mm | 152.67 mm |  | Tolerance = 152.40 +/- 0.70 mm
|  | Base Plate Planeness: | < 0.004" |  |  | Tolerance = 0.005 in.
(maximum deviation)
AASHTO T 19 - Density of Water (Interpolated)
Mold Volume: |  |  |  |  | ° F |  | lb/ft³ |  | ° F | lb/ft³
[5 empty] | 65 |  | 62.336 |  | 74 | 62.269
|  | Empty Weight of Mold, Base, |  |  | 66 |  | 62.329 |  | 75 | 62.261
|  | and Glass Plate (g): | 6916.6 |  | 67 |  | 63.322 |  | 76 | 62.252
[5 empty] | 68 |  | 62.315 |  | 77 | 62.243
|  | Weight of Mold, Base, and |  |  | 69 |  | 62.308 |  | 78 | 62.234
|  | Glass Plate filled with Water (g): | 9039.2 |  | 70 |  | 62.301 |  | 79 | 62.225
[5 empty] | 71 |  | 62.293 |  | 80 | 62.216
|  | Temperature of Water (° F): | 71 °F |  | 72 |  | 62.285 |  | 81 | 62.206
[5 empty] | 73 |  | 62.277 |  | 82 | 62.196
|  | Density of Water (lb/ft³): | 62.293
|  | Mold Volume (ft³): | 0.075121 |  |  |  |  | Pass/Fail: | Pass
|  |  | Tolerance = 0.0750 ± 0.0009 ft³ | (0.0741 - 0.0759)
Comments:
Initial By:

## Pycnometer Blank
| Date: |  |  | Performed by:
| Next Calibration Due: |  |  | Last Calibration:
Pycnometer # | 1 |  |  | Equipment ID:
Manufacturer:
Pyc | Pyc + Water | Temp, F | Water | Model #:
Serial #:
UofA ID:
Location:
Procedure:
Pycnometer # | G |  |  | Equipment ID:
Manufacturer:
Pyc | Pyc + Water | Temp, F | Water | Model #:
Serial #:
UofA ID:
Location:
Procedure:

## Pycnometer
| Date: |  |  | Performed by: | Mary Fleck
| Next Calibration Due: | April 2013 |  | Last Calibration: | April 2011
Pycnometer # | 1 |  |  | Equipment ID: | 500 mL Pyc. #1
|  |  |  | Manufacturer: | Kimax
Pyc | Pyc + Water | Temp, F | Water | Model #: | NA
|  |  | 0 | Serial #: | NA
0 |  |  | 0 | UofA ID: | NA
0 |  |  | 0 | Location: | ST (218)
0 |  |  | 0 | Procedure: | AASHTO T-100 (P7)
Pycnometer # | G |  |  | Equipment ID: | 500 mL Pyc. #2 (G)
|  |  |  | Manufacturer: | Kimax
Pyc | Pyc + Water | Temp, F | Water | Model #: | NA
|  |  | 0 | Serial #: | NA
0 |  |  | 0 | UofA ID: | NA
0 |  |  | 0 | Location: | ST (216)
0 |  |  | 0 | Procedure: | AAHSTO T-100 (P7)

## Pycnometer Statement
| Date: | 2012-04-12 |  | Performed by: | Mary Fleck
| Next Calibration Due: | April 2013 |  | Last Calibration: | April 2011
Pycnometer # | 1 |  |  | Equipment ID: | 500 mL Pyc. #1
|  |  |  | Manufacturer: | Kimax
Pyc | Pyc + Water | Temp, F | Water | Model #: | NA
|  |  |  | Serial #: | NA
|  |  |  | UofA ID: | NA
|  |  |  | Location: | ST (218)
|  |  |  | Procedure: | AASHTO T-100 (P7)
This calibration is no longer being performed. The calibration of the pycnometer is performed
during each test by bringing the pyc + water to the same test temperature as the
pyc + water + sample before obtaining it's weight. This eliminates the need for this procedure.
Initial By:
Pycnometer # | G |  |  | Equipment ID: | 500 mL Pyc. #2 (G)
|  |  |  | Manufacturer: | Kimax
Pyc | Pyc + Water | Temp, F | Water | Model #: | NA
|  |  |  | Serial #: | NA
|  |  |  | UofA ID: | NA
|  |  |  | Location: | ST (216)
|  |  |  | Procedure: | AAHSTO T-100 (P7)
This calibration is no longer being performed. The calibration of the pycnometer is performed
during each test by bringing the pyc + water to the same test temperature as the
pyc + water + sample before obtaining it's weight. This eliminates the need for this procedure.
Initial By: