## Volumetric Air Meter
Volumetric Air Meter Calibration Worksheet
|  | Equipment ID: | Volumetric Air Meter |  |  |  |  | Date: | 2008-05-12
|  | Manufacturer: | Test Mark |  |  |  |  | Performed by: | F. Griffith
|  | Model #: | CA - 0600 - P
|  | Serial #: | NA |  |  |  |  | Next Calibration Due: | May 2009
|  | UofA ID: | NA |  |  |  |  | Last Calibration: | 2007-05-01
|  | Location: | ST (218)
| Calibration Item: | Verify measuring bowl, air content gradations, and calibrated cup volume
| Calibration Procedure: | In-House Procedure #07-4
| Calibration Equipment: | balance read. to 0.1% of weight of water to fill meter bowl
Verify the measuring bowl volume:
|  | Wt. of bowl + glass plate: | 2.1 | lb |  |  |  |  | Temperature of water (T): | 69.0° F
|  | Wt. of bowl + glass plate + water: | 6.86 | lb |  |  |  |  | Density of Water at T (C 29): | 62.307 | lb/ft³
|  | Wt. of Water in bowl (w): | 4.76 | lb
|  | Volume of bowl (mass/density): | 0.076396 | ft³ |  |  | Tolerance = at least 0.075 cf
Verify that the top assembly volume is at least 20% > bowl volume:
|  | Wt. of meter filled with water: | 17.04 | lb |  |  |  |  | Wt. of meter: | 2.96 | lb
|  | Volume of water in top assembly: | 0.1027 | ft³ |  |  |  |  | Wt. of water in meter: | 6.4 | lb
(Wt. water meter - Wt. water bowl) / Density
|  | Ratio of vol. of top assembly to bowl: | 1.344 |  |  |  | Tolerance = at least 1.20
Bowl diameter should be 1.00 - 1.25 x the height, and capacity of at least 0.075 ft3
|  | Diameter of bowl: | 5.36 | in |  | (outside dimension)
|  | Height of bowl: | 5.125 | in |  | (outside dimension)
|  | Diameter/Ht. Ratio: | 1.046 |  |  |  | Tolerance = 1.0 to 1.25
Check air content gradations on the Neck: | [6 empty] | Tolerance = within 0.1%
at 8%___ | at 7%___ | at 6%___ | at 5%___ |  | at 4%___
at 3%___ | at 2%___ | at 1%___ | at 0%___
Check volume of calibrated cup:
|  | Mass of calibrated cup and glass: |  | lb |  |  |  |  | Density of Water at 70 deg. (C 29): | 62.301 | lb/ft³
|  | Mass of cup, glass and water: |  | lb
|  | Volume of cup(mass/density): | 0 | lb/ft³
|  | Volume of Bowl/Volume of cup: | 0 |  | Tolerance = 1.00 ± 0.04% bowl |  |  |  |  | Pass/Fail:
Comments:
[10 empty] | Initial By: | ftg

## VAM Strike Off Bar
Volumetric Air Meter Plastic Strike Off Bar Calibration Worksheet
|  | Equipment ID: | VAM Strike Off Bar |  | Date: | May
|  | Manufacturer: | Test Mark |  | Performed by: | Various
|  | Model #: | CA - 0600 - SO Bar
|  | Serial #: | NA
|  | UofA ID: | NA |  | Calibration Interval: | 12 months
|  | Location: | ST (218)
| Calibration Item: | Length, width, and thickness of bar
| Calibration Procedure: | ASTM C 231, Section 4.11
| Calibration Equipment: | Mitutoyo Digital calipers
Date | Length (approx. 12") | Width (approx. 3/4") | Thickness (at least 1/4") | Comments |  |  | Pass/Fail | Initial
2007-05-14 | 12" | 1.5" | 0.25" | New Bar |  |  | Pass | sgw
2008-05-12 | 12" | 1.5" | 0.24" |  |  |  | Pass | ftg
| Comments: | Tolerance : 6 mm = 0.236 inches
Initial By:

## VAM Tamping Rod
VAM Tamping Rod
|  | Equipment ID: | VAM Tamping Rod |  | Date: | May
|  | Manufacturer: | Test Mark |  | Performed by: | Various
|  | Model #: | CA - 0600 - TR
|  | Serial #: | NA
|  | UofA ID: | NA |  | Calibration Interval: | 12 months
|  | Location: | ST (218)
| Calibration Item: | Length, diameter, and shape of rod
| Calibration Procedure: | In-House #07-2
| Calibration Equipment: | Mitutoyo Digital calipers, tape measure
Date | Length (approx. 12") | Diameter (5/8" ± 1/16") | Hemispherical Tip? | Comments |  |  | Pass/Fail | Initial
2007-05-14 | 12" | 0.6025" | Yes |  |  |  | Pass | sgw
2008-05-12 | 12" | 0.6360" | Yes |  |  |  | Pass | ftg
| Comments: | Tolerance = 0.5625" to 0.6875"
Initial By: