## Slump Cone
Slump Cone
|  | Equipment ID: | Slump Cone |  |  | Date: | 2009-05-12
|  | Manufacturer: | Gilson |  |  | Performed by: | ftg
|  | Model #: | HM - 45
|  | Serial #: | NA |  |  | Next Calibration Due: | May 2010
|  | UofA ID: | NA |  |  | Last Calibration: | 5/12/2008
|  | Location: | ST (218)
| Calibration Item: | Verify critical dimensions of the cone
| Calibration Procedure: | In-House Procedure #07-1
| Calibration Equipment: | Mititoyo Digimatic Calipers, and ruler read. to 1 mm.
|  | Thickness of cone: | 0.0615 | 0.066 | Tolerance = Not less than 0.060 inches
|  | Bottom Inside Diameter of Cone: | 8.0445" | 8.005" | Tolerance = 8 +/- 1/8 inches
|  | Top Inside Diameter of Cone: | 4.0295" | 4.0090" | Tolerance = 4 +/- 1/8 inches
|  | Height of Cone: | 12.120" | 12.083" | Tolerance = 12 +/- 1/8 inches
|  | Cone Interior Smooth: | Yes | Yes | Pass/Fail
|  | Base and top opening parallel: | Yes | Yes | Pass/Fail
[7 empty] | Pass/Fail: | Pass
Comments:
[8 empty] | Initial By: | ftg

## Cylinder Molds
Single-Use Cylinder Mold
|  | Equipment ID: | Cylinder Molds |  |  |  | Date: | May
|  | Manufacturer: | Test Mark |  |  |  | Performed by: | Various
|  | Model #: | MA Industries
|  | Serial #: | NA
|  | UofA ID: | NA |  |  |  | Next Calibration Due: | May 2010
|  | Location: | ST(218) |  |  |  | Last Calibration: | May 2008
| Calibration Item: | Verify critical dimensions of the cylinder
| Calibration Procedure: | ASTM C 31 - C 470
Resistance to damage under use and water leakage
| Calibration Equipment: | Mititoyo Digimatic Calipers, and ruler read. to 1 mm.
Date | Length | Diameter | L/D ratio (approx. 2) |  | Comments |  |  |  | Pass/Fail | Initial
2009-05-12 | 12.0655" | 6.0320" | 2 | Brand: |  | Test Mark
Shipment:
|  |  |  | Water Leakage: |  |  |  |  | Pass | ftg
|  |  |  | Resistance to Damage: |  |  |  |  | Pass | ftg
2009-09-15 | 12.08" | 6.03" | 2 | Brand: |  | MA Industries
|  |  |  | Shipment: |  | 2008-07-01
|  |  |  | Water Leakage: |  |  |  |  | Pass | ftg
|  |  |  | Resistance to Damage: |  |  |  |  | Pass | ftg
2009-09-15 | 12.13" | 6.12" | 1.98 | Brand: |  | MA Industries
|  |  |  | Shipment: |  | 2008-07-01
|  |  |  | Water Leakage: |  |  |  |  | Pass | ftg
|  |  |  | Resistance to Damage: |  |  |  |  | Pass | ftg
2009-09-15 | 12.24" | 6.12" | 2 | Brand: |  | MA Industries
|  |  |  | Shipment: |  | 2008-07-01
|  |  |  | Water Leakage: |  |  |  |  | Pass | ftg
|  |  |  | Resistance to Damage: |  |  |  |  | Pass | ftg
Brand:
Shipment:
Water Leakage:
Resistance to Damage:
| Comments: | [7 empty] | Initial by: | sgw