## Air Pressure Meter (1)
Air Pressure Meter Calibration Worksheet
|  | Equipment ID: | Air Pressure Meter |  |  |  | Date: | 2010-02-23
|  | Manufacturer: | Gilson |  |  |  | Performed by: | F. Griffith
|  | Model #: | HM - 30
|  | Serial #: | NA |  |  |  | Next Calibration Due: | May 2010
|  | UofA ID: | NA |  |  |  | Last Calibration: | November 2009
|  | Location: | ST (218)
| Calibration Item: | Verify calibration vessel, measuring bowl, and air readings
| Calibration Procedure: | In-House Procedure #07-3
| Calibration Equipment: | balance read. to 0.01 lb, calibration tubing
Verify the calibration vessel:
|  | Wt. of vessel: |  | lbs
|  | Wt. of vessel and water: |  | lbs
|  | Wt. of Water in vessel (w): |  | lbs
|  | Wt. of pressure bowl + plate: |  | lbs
|  | Wt. of water + pressure bowl + plate: |  | lbs
|  | Wt. of Water in pressure bowl (W): |  | lbs
|  | Effective Volume (R = w/W): |  | % | Tolerance = Approximately 5.0%
Air pressure bowl diameter should be 0.75 - 1.25 x the height, and capacity of at least 0.20 ft3
|  | Inside Diameter of bowl: |  | in
|  | Inside Height of bowl: |  | in
|  | Diameter/Ht. Ratio: |  |  |  | Tolerance = 0.75 to 1.25
|  | Volume of Bowl: |  | ft³ |  | At least 0.20 ft3
Check of Allowance for Expansion Factor (D):
|  | Initial Pressure: |  | %
|  | Final Pressure: |  | % |  | Final Pressure should be 0%
Check air content gradations on the Pressure Gage:
|  | Initial Pressure: | 3 | %
|  | % Water Removed: | 5 | %
|  | Gage Reading: | 5 | % |  | Tolerance = ± 0.1 %
[8 empty] | Pass/Fail: | Pass
Comments:
Initial By:

## Air Pressure Meter (2)
Air Pressure Meter Calibration Worksheet
|  | Equipment ID: | Air Pressure Meter |  |  |  | Date: | 2010-05-18
|  | Manufacturer: | Gilson |  |  |  | Performed by: | DAP
|  | Model #: | HM - 30
|  | Serial #: | NA |  |  |  | Next Calibration Due: | August 2010
|  | UofA ID: | NA |  |  |  | Last Calibration: | February 2010
|  | Location: | ST (218)
| Calibration Item: | Verify calibration vessel, measuring bowl, and air readings
| Calibration Procedure: | In-House Procedure #07-3
| Calibration Equipment: | balance read. to 0.01 lb, calibration tubing
Verify the calibration vessel:
|  | Wt. of vessel: | 0.33 | lbs
|  | Wt. of vessel and water: | 1.101 | lbs
|  | Wt. of Water in vessel (w): | 0.771 | lbs
|  | Wt. of pressure bowl + plate: | 7.444 | lbs
|  | Wt. of water + pressure bowl + plate: | 22.97 | lbs
|  | Wt. of Water in pressure bowl (W): | 15.526 | lbs
|  | Effective Volume (R = w/W): | 4.9659 | % | Tolerance = Approximately 5.0%
Air pressure bowl diameter should be 0.75 - 1.25 x the height, and capacity of at least 0.20 ft3
|  | Inside Diameter of bowl: | 7.5 | in
|  | Inside Height of bowl: | 8.83 | in
|  | Diameter/Ht. Ratio: | 0.85 |  |  | Tolerance = 0.75 to 1.25
|  | Volume of Bowl: | 0.226 | ft³ |  | At least 0.20 ft3
Check of Allowance for Expansion Factor (D):
|  | Initial Pressure: | 3 | %
|  | Final Pressure: | 0 | % |  | Final Pressure should be 0%
Check air content gradations on the Pressure Gage:
|  | Initial Pressure: | 3 | %
|  | % Water Removed: | 5 | %
|  | Gage Reading: | 5 | % |  | Tolerance = ± 0.1 %
[8 empty] | Pass/Fail: | Pass
Comments:
[9 empty] | Initial By: | ftg

## Air Pressure Meter (3)
Air Pressure Meter Calibration Worksheet
|  | Equipment ID: | Air Pressure Meter |  |  |  | Date: | 2010-07-30
|  | Manufacturer: | Gilson |  |  |  | Performed by: | F. Griffith
|  | Model #: | HM - 30
|  | Serial #: | NA |  |  |  | Next Calibration Due: | October 2010
|  | UofA ID: | NA |  |  |  | Last Calibration: | May 2010
|  | Location: | ST (218)
| Calibration Item: | Verify calibration vessel, measuring bowl, and air readings
| Calibration Procedure: | In-House Procedure #07-3
| Calibration Equipment: | balance read. to 0.01 lb, calibration tubing
Verify the calibration vessel:
|  | Wt. of vessel: |  | lbs
|  | Wt. of vessel and water: |  | lbs
|  | Wt. of Water in vessel (w): |  | lbs
|  | Wt. of pressure bowl + plate: |  | lbs
|  | Wt. of water + pressure bowl + plate: |  | lbs
|  | Wt. of Water in pressure bowl (W): |  | lbs
|  | Effective Volume (R = w/W): |  | % | Tolerance = Approximately 5.0%
Air pressure bowl diameter should be 0.75 - 1.25 x the height, and capacity of at least 0.20 ft3
|  | Inside Diameter of bowl: |  | in
|  | Inside Height of bowl: |  | in
|  | Diameter/Ht. Ratio: |  |  |  | Tolerance = 0.75 to 1.25
|  | Volume of Bowl: |  | ft³ |  | At least 0.20 ft3
Check of Allowance for Expansion Factor (D):
|  | Initial Pressure: |  | %
|  | Final Pressure: |  | % |  | Final Pressure should be 0%
Check air content gradations on the Pressure Gage:
|  | Initial Pressure: | 3 | %
|  | % Water Removed: | 5 | %
|  | Gage Reading: | 5 | % |  | Tolerance = ± 0.1 %
[8 empty] | Pass/Fail: | Pass
Comments:
[9 empty] | Initial By: | ftg

## Air Pressure Meter (4)
Air Pressure Meter Calibration Worksheet
|  | Equipment ID: | Air Pressure Meter |  |  |  | Date: | 2010-11-23
|  | Manufacturer: | Gilson |  |  |  | Performed by: | F. Griffith
|  | Model #: | HM - 30
|  | Serial #: | NA |  |  |  | Next Calibration Due: | February 2010
|  | UofA ID: | NA |  |  |  | Last Calibration: | July 2010
|  | Location: | ST (218)
| Calibration Item: | Verify calibration vessel, measuring bowl, and air readings
| Calibration Procedure: | In-House Procedure #07-3
| Calibration Equipment: | balance read. to 0.01 lb, calibration tubing
Verify the calibration vessel:
|  | Wt. of vessel: |  | lbs
|  | Wt. of vessel and water: |  | lbs
|  | Wt. of Water in vessel (w): |  | lbs
|  | Wt. of pressure bowl + plate: |  | lbs
|  | Wt. of water + pressure bowl + plate: |  | lbs
|  | Wt. of Water in pressure bowl (W): |  | lbs
|  | Effective Volume (R = w/W): |  | % | Tolerance = Approximately 5.0%
Air pressure bowl diameter should be 0.75 - 1.25 x the height, and capacity of at least 0.20 ft3
|  | Inside Diameter of bowl: |  | in
|  | Inside Height of bowl: |  | in
|  | Diameter/Ht. Ratio: |  |  |  | Tolerance = 0.75 to 1.25
|  | Volume of Bowl: |  | ft³ |  | At least 0.20 ft3
Check of Allowance for Expansion Factor (D):
|  | Initial Pressure: |  | %
|  | Final Pressure: |  | % |  | Final Pressure should be 0%
Check air content gradations on the Pressure Gage:
|  | Initial Pressure: | 3 | %
|  | % Water Removed: | 5 | %
|  | Gage Reading: | 5 | % |  | Tolerance = ± 0.1 %
[8 empty] | Pass/Fail: | Pass
Comments:
[9 empty] | Initial By: | ftg

## APM Strike Off Plate
Air Pressure Meter Strike Off Plate Calibration Worksheet
|  | Equipment ID: | Strike Off Plate |  | Date: | 2007-05-14
|  | Manufacturer: | Gilson |  | Performed by: | F. Griffith
|  | Model #: | HM SO Plate
|  | Serial #: | NA
|  | UofA ID: | NA |  | Calibration Interval: | 12 months
|  | Location: | ST (218)
| Calibration Item: | Length, width, and thickness of bar
| Calibration Procedure: | ASTM C 231, Section 4.11
| Calibration Equipment: | Mitutoyo Digital calipers
Date | Length (2"≥ bowl dia.) | Width (2"≥ bowl dia.) | Thickness (≥ 13 mm) | Comments |  |  | Pass/Fail | Initial
2007-05-14 | 14" | 14" | 1/2" |  |  |  | Pass | sgw
2008-05-12 | 14" | 14" | 1/2" |  |  |  | Pass | ftg
2009-05-12 | 14.024" | 13.793" | 0.512" |  |  |  | Pass | ftg
2010-05-18 | 14" | 14" | 0.51" |  |  |  | Pass | ftg
Comments:

## APM Strike Off Bar
Air Pressure Meter Strike Off Bar Calibration Worksheet
|  | Equipment ID: | Stike Off Bar |  | Date: | 2007-05-14
|  | Manufacturer: | Gilson |  | Performed by: | F. Griffith
|  | Model #: | RPHM-30KEY34R
|  | Serial #: | NA
|  | UofA ID: | NA |  | Calibration Interval: | 12 months
|  | Location: | ST (218)
| Calibration Item: | Length, width, and thickness of bar
| Calibration Procedure: | ASTM C 231, Section 4.11
| Calibration Equipment: | Mitutoyo Digital calipers
Date | Length (approx. 12") | Width (approx. 3/4") | Thickness ( ≥ 1/8") | Comments |  |  | Pass/Fail | Initial
2007-05-14 | 12 1/16" | 1" | 3/16" |  |  |  | Pass | sgw
2008-05-12 | 12.0" | 1" | 3/16" |  |  |  | Pass | ftg
2009-05-12 | 12.0615" | 1.00" | 0.1905" |  |  |  | Pass | ftg
2010-05-27 | 12.059" | 1.00" | 0.1879" |  |  |  | Pass | ftg
Comments: