## Slump Cone
Slump Cone
|  | Equipment ID: | Slump Cone |  |  | Date: | 2012-05-21
|  | Manufacturer: | Gilson |  |  | Performed by: | M. Fleck
|  | Model #: | HM - 45
|  | Serial #: | NA |  |  | Next Calibration Due: | May 2013
|  | UofA ID: | NA |  |  | Last Calibration: | May 2011
|  | Location: | ST (218)
| Calibration Item: | Verify critical dimensions of the cone
| Calibration Procedure: | In-House Procedure #07-1
| Calibration Equipment: | Mititoyo Digimatic Calipers, and ruler read. to 1 mm.
|  | Thickness of cone: | 0.705" | 0.705" | Tolerance = Not less than 0.060 inches
|  | Bottom Inside Diameter of Cone: | 7 15/16" | 7 15/16" | Tolerance = 8 +/- 1/8 inches
|  | Top Inside Diameter of Cone: | 4" | 4" | Tolerance = 4 +/- 1/8 inches
|  | Height of Cone: | 12" | 12" | Tolerance = 12 +/- 1/8 inches
|  | Cone Interior Smooth: | Yes | Yes | Pass/Fail
|  | Base and top opening parallel: | Yes | Yes | Pass/Fail
[7 empty] | Pass/Fail: | Pass
Comments:
[8 empty] | Initial By: | sgw

## Cylinder Molds
Single-Use Cylinder Mold
|  | Equipment ID: | Cylinder Molds |  |  |  | Date: | 2012-05-30
|  | Manufacturer: | Jatco Inc. |  |  |  | Performed by:
|  | Model #: | Forney
Serial #:
|  | UofA ID: |  |  |  |  | Next Calibration Due: | May 2013
|  | Location: |  |  |  |  | Last Calibration: | May 2011
| Calibration Item: | Verify critical dimensions of the cylinder
| Calibration Procedure: | ASTM C 31 - C 470
Resistance to damage under use and water leakage
| Calibration Equipment: | Mititoyo Digimatic Calipers, and ruler read. to 1 mm.
Date | Length | Diameter | L/D ratio (approx. 2) |  | Comments |  |  |  | Pass/Fail | Initial
2012-05-30 | 12.03" | 6.02" | 2 | Brand: |  | Jatco
|  |  |  | Shipment: |  | 2011-06-01
|  |  |  | Water Leakage: |  |  |  |  | Pass | ftg
|  |  |  | Resistance to Damage: |  |  |  |  | Pass | ftg
2012-05-30 | 12.04" | 5.59" | 2.15 | Brand: |  | Jatco
|  |  |  | Shipment: |  | 2011-06-01
|  |  |  | Water Leakage: |  |  |  |  | Pass | ftg
|  |  |  | Resistance to Damage: |  |  |  |  | Pass | ftg
2012-05-30 | 12.05" | 6.00" | 2.01" | Brand: |  | Jatco
|  |  |  | Shipment: |  | 2011-06-01
|  |  |  | Water Leakage: |  |  |  |  | Pass | ftg
|  |  |  | Resistance to Damage: |  |  |  |  | Pass | ftg
Brand:
Shipment:
Water Leakage:
Resistance to Damage:
Brand:
Shipment:
Water Leakage:
Resistance to Damage:
| Comments: | New shipment replacing previous brand of molds May 2012. | [6 empty] | Initial by: | sgw