## Scale (G1-3)
Scale Calibration Worksheet
| Equipment ID: | G1 - 3 |  |  |  |  | Date: | 2008-09-09
| Manufacturer: | Ohaus |  |  |  |  | Performed by: | M. Fleck
| Model #: | GT 400
| Serial #: | 5376 |  |  |  |  | Next Calibration Due: | September 2009
| UofA ID: | NA |  |  |  |  | Last Calibration: | NA (new)
| Location: | STO (218)
| Calibration Item: | Verify mass readings for at least 5 points in the range of the balance
Check sensitivity (0.1%) for at least 5 points in the range of the balance
| Calibration Procedure: | AASHTO M-231
| Calibration Equipment: | Class F, Class I, and 10-kg Calibration Weights
Tolerances:
| Class |  | Readability and Sensitivity |  |  |  | Accuracy
| G1 |  |  | 0.01 g |  | 0.02 g or 0.1 percent
| G2 |  |  | 0.1 g |  | 0.2 or 0.1 percent
| G5 |  |  | 1 g |  | 2 g or 0.1 percent
| G20 |  |  | 5 g |  | 5 g or 0.1 percent
| G100 |  |  | 20 g |  | 20 g or 0.1 percent
| Weight Applied (g) |  | Scale Reading (g) |  | Tolerance (g) |  |  | Pass/Fail
| 0.01 |  | 0.01 |  | 0.02
| 1 |  | 1 |  | 0.02
| 5 |  | 5 |  | 0.02
| 10 |  | 10 |  | 0.02
| 10.01 |  | 10.01 |  | 0.02
| 50 |  | 50 |  | 0.05
| 50.01 |  | 50.01 |  | 0.05
| 100 |  | 100 |  | 0.1
| 100.01 |  | 100.01 |  | 0.1
| 200 |  | 200 |  | 0.2
| 200.01 |  | 200.01 |  | 0.2
| 300 |  | 300 |  | 0.3
| 400 |  | 400 |  | 0.4
[9 empty] | Pass/Fail: | Pass
Comments:
[9 empty] | Initial By: | sgw

## Scale (G1-4)
Scale Calibration Worksheet
| Equipment ID: | G1 - 4 |  |  |  |  | Date: | 2008-09-09
| Manufacturer: | Ohaus |  |  |  |  | Performed by: | M. Fleck
| Model #: | GT 400
| Serial #: | 5373 |  |  |  |  | Next Calibration Due: | September 2010
| UofA ID: | NA |  |  |  |  | Last Calibration: | NA (new)
| Location: | STO (218)
| Calibration Item: | Verify mass readings for at least 5 points in the range of the balance
Check sensitivity (0.1%) for at least 5 points in the range of the balance
| Calibration Procedure: | AASHTO M-231
| Calibration Equipment: | Class F, Class I, and 10-kg Calibration Weights
Tolerances:
| Class |  | Readability and Sensitivity |  |  |  | Accuracy
| G1 |  |  | 0.01 g |  | 0.02 g or 0.1 percent
| G2 |  |  | 0.1 g |  | 0.2 or 0.1 percent
| G5 |  |  | 1 g |  | 2 g or 0.1 percent
| G20 |  |  | 5 g |  | 5 g or 0.1 percent
| G100 |  |  | 20 g |  | 20 g or 0.1 percent
| Weight Applied (g) |  | Scale Reading (g) |  | Tolerance (g) |  |  | Pass/Fail
| 0.01 |  | 0.01 |  | 0.02
| 1 |  | 1 |  | 0.02
| 5 |  | 5 |  | 0.02
| 10 |  | 10 |  | 0.02
| 10.01 |  | 10.01 |  | 0.02
| 50 |  | 50 |  | 0.05
| 50.01 |  | 50.01 |  | 0.05
| 100 |  | 100 |  | 0.1
| 100.01 |  | 100.01 |  | 0.1
| 200 |  | 200 |  | 0.2
| 200.01 |  | 200.01 |  | 0.2
| 300 |  | 300 |  | 0.3
| 400 |  | 400 |  | 0.4
[9 empty] | Pass/Fail: | Pass
Comments:
[9 empty] | Initial By: | sgw

## Scale (G2-A)
Scale Calibration Worksheet
| Equipment ID: | AG - A |  |  |  |  | Date: | 2008-09-09
| Manufacturer: | Ohaus |  |  |  |  | Performed by: | M. Fleck
| Model #: | EP12001
| Serial #: | 1126012086 |  |  |  |  | Next Calibration Due: | September 2009
| UofA ID: | NA |  |  |  |  | Last Calibration: | September 2007
| Location: | AG (233) Rear Table (West)
| Calibration Item: | Verify mass readings for at least 5 points in the range of the balance
Check sensitivity (0.1%) for at least 5 points in the range of the balance
| Calibration Procedure: | AASHTO M-231
| Calibration Equipment: | Class F, Class I, and 10-kg Calibration Weights
Tolerances:
| Class |  | Readability and Sensitivity |  |  |  | Accuracy
| G1 |  |  | 0.01 g |  | 0.02 g or 0.1 percent
| G2 |  |  | 0.1 g |  | 0.2 or 0.1 percent
| G5 |  |  | 1 g |  | 2 g or 0.1 percent
| G20 |  |  | 5 g |  | 5 g or 0.1 percent
| G100 |  |  | 20 g |  | 20 g or 0.1 percent
| Weight Applied (g) |  | Scale Reading (g) |  | Tolerance (g) |  |  | Pass/Fail
| 0.1 |  | 0.1 |  | 0.2
| 1 |  | 1 |  | 0.2
| 10 |  | 10 |  | 0.2
| 100 |  | 100 |  | 0.2
| 100.1 |  | 100.1 |  | 0.2
| 200 |  | 200 |  | 0.2
| 500 |  | 500 |  | 0.5
| 500.1 |  | 500.1 |  | 0.5
| 1000 |  | 1000 |  | 1
| 2000 |  | 2000 |  | 2
| 2000.1 |  | 2000.1 |  | 2
| 5000 |  | 5000.1 |  | 5
| 5000.1 |  | 5000.2 |  | 5
| 10000 |  | 10000 |  | 10
| 11000 |  | 11000 |  | 11
[9 empty] | Pass/Fail: | Pass
| Comments: | Span recalibrated.
[9 empty] | Initial By: | sgw

## Scale (G2-B)
Scale Calibration Worksheet
| Equipment ID: | AG - B |  |  |  |  | Date: | 2008-09-09
| Manufacturer: | Ohaus |  |  |  |  | Performed by: | M. Fleck
| Model #: | EP12001
| Serial #: | 1125430302 |  |  |  |  | Next Calibration Due: | September 2009
| UofA ID: | NA |  |  |  |  | Last Calibration: | September 2007
| Location: | AG (233) Front Table (East)
| Calibration Item: | Verify mass readings for at least 5 points in the range of the balance
Check sensitivity (0.1%) for at least 5 points in the range of the balance
| Calibration Procedure: | AASHTO M-231
| Calibration Equipment: | Class F, Class I, and 10-kg Calibration Weights
Tolerances:
| Class |  | Readability and Sensitivity |  |  |  | Accuracy
| G1 |  |  | 0.01 g |  | 0.02 g or 0.1 percent
| G2 |  |  | 0.1 g |  | 0.2 or 0.1 percent
| G5 |  |  | 1 g |  | 2 g or 0.1 percent
| G20 |  |  | 5 g |  | 5 g or 0.1 percent
| G100 |  |  | 20 g |  | 20 g or 0.1 percent
| Weight Applied (g) |  | Scale Reading (g) |  | Tolerance (g) |  |  | Pass/Fail
| 0.1 |  | 0.1 |  | 0.2
| 1 |  | 1 |  | 0.2
| 10 |  | 10 |  | 0.2
| 100 |  | 100 |  | 0.2
| 100.1 |  | 100.1 |  | 0.2
| 200 |  | 200.1 |  | 0.2
| 500 |  | 500 |  | 0.5
| 500.1 |  | 500.1 |  | 0.5
| 1000 |  | 1000 |  | 1
| 2000 |  | 2000 |  | 2
| 2000.1 |  | 2000.1 |  | 2
| 5000 |  | 5000 |  | 5
| 5000.1 |  | 5000.1 |  | 5
| 10000 |  | 10000 |  | 10
| 11000 |  | 11000.1 |  | 11
[9 empty] | Pass/Fail: | Pass
Comments:
[9 empty] | Initial By: | sgw

## Scale (G2-E)
Scale Calibration Worksheet
| Equipment ID: | AG - E |  |  |  |  | Date: | 2008-09-09
| Manufacturer: | Ohaus |  |  |  |  | Performed by: | M. Fleck
| Model #: | EP12001
| Serial #: | 1125330661 |  |  |  |  | Next Calibration Due: | September 2009
| UofA ID: | NA |  |  |  |  | Last Calibration: | September 2007
| Location: | AG (233) Front SpG (East)
| Calibration Item: | Verify mass readings for at least 5 points in the range of the balance
Check sensitivity (0.1%) for at least 5 points in the range of the balance
| Calibration Procedure: | AASHTO M-231
| Calibration Equipment: | Class F, Class I, and 10-kg Calibration Weights
Tolerances:
| Class |  | Readability and Sensitivity |  |  |  | Accuracy
| G1 |  |  | 0.01 g |  | 0.02 g or 0.1 percent
| G2 |  |  | 0.1 g |  | 0.2 or 0.1 percent
| G5 |  |  | 1 g |  | 2 g or 0.1 percent
| G20 |  |  | 5 g |  | 5 g or 0.1 percent
| G100 |  |  | 20 g |  | 20 g or 0.1 percent
| Weight Applied (g) |  | Scale Reading (g) |  | Tolerance (g) |  |  | Pass/Fail
| 0.1 |  | 0.1 |  | 0.2
| 1 |  | 1 |  | 0.2
| 10 |  | 10 |  | 0.2
| 100 |  | 100 |  | 0.2
| 100.1 |  | 100.1 |  | 0.2
| 200 |  | 200 |  | 0.2
| 500 |  | 500 |  | 0.5
| 500.1 |  | 500.1 |  | 0.5
| 1000 |  | 1000 |  | 1
| 2000 |  | 2000 |  | 2
| 2000.1 |  | 2000.1 |  | 2
| 5000 |  | 5000 |  | 5
| 5000.1 |  | 5000.1 |  | 5
| 10000 |  | 10000 |  | 10
| 11000 |  | 11000 |  | 11
[9 empty] | Pass/Fail: | Pass
Comments:
[9 empty] | Initial By: | sgw

## Scale (G2-F)
Scale Calibration Worksheet
| Equipment ID: | AG - F |  |  |  |  | Date: | 2008-09-09
| Manufacturer: | Ohaus |  |  |  |  | Performed by: | M. Fleck
| Model #: | EP12001
| Serial #: | 1126051479 |  |  |  |  | Next Calibration Due: | September 2009
| UofA ID: | NA |  |  |  |  | Last Calibration: | September 2007
| Location: | AG (233) Rear SpG (West)
| Calibration Item: | Verify mass readings for at least 5 points in the range of the balance
Check sensitivity (0.1%) for at least 5 points in the range of the balance
| Calibration Procedure: | AASHTO M-231
| Calibration Equipment: | Class F, Class I, and 10-kg Calibration Weights
Tolerances:
| Class |  | Readability and Sensitivity |  |  |  | Accuracy
| G1 |  |  | 0.01 g |  | 0.02 g or 0.1 percent
| G2 |  |  | 0.1 g |  | 0.2 or 0.1 percent
| G5 |  |  | 1 g |  | 2 g or 0.1 percent
| G20 |  |  | 5 g |  | 5 g or 0.1 percent
| G100 |  |  | 20 g |  | 20 g or 0.1 percent
| Weight Applied (g) |  | Scale Reading (g) |  | Tolerance (g) |  |  | Pass/Fail
| 0.1 |  | 0.1 |  | 0.2
| 1 |  | 1 |  | 0.2
| 10 |  | 10 |  | 0.2
| 100 |  | 100 |  | 0.2
| 100.1 |  | 100.1 |  | 0.2
| 200 |  | 200 |  | 0.2
| 500 |  | 500 |  | 0.5
| 500.1 |  | 500.1 |  | 0.5
| 1000 |  | 1000 |  | 1
| 2000 |  | 2000 |  | 2
| 2000.1 |  | 2000.1 |  | 2
| 5000 |  | 5000 |  | 5
| 5000.1 |  | 5000.1 |  | 5
| 10000 |  | 10000 |  | 10
| 11000 |  | 11000 |  | 11
[9 empty] | Pass/Fail: | Pass
Comments:
[9 empty] | Initial By: | sgw

## Scale (G2-CL1)
Scale Calibration Worksheet
| Equipment ID: | CL - 1 |  |  |  |  | Date: | 2008-09-09
| Manufacturer: | Ohaus |  |  |  |  | Performed by: | M. Fleck
| Model #: | EP12001
| Serial #: | 1128061460 |  |  |  |  | Next Calibration Due: | September 2009
| UofA ID: | NA |  |  |  |  | Last Calibration: | December 2007
| Location: | CL (220) Left SpG
| Calibration Item: | Verify mass readings for at least 5 points in the range of the balance
Check sensitivity (0.1%) for at least 5 points in the range of the balance
| Calibration Procedure: | AASHTO M-231
| Calibration Equipment: | Class F, Class I, and 10-kg Calibration Weights
Tolerances:
| Class |  | Readability and Sensitivity |  |  |  | Accuracy
| G1 |  |  | 0.01 g |  | 0.02 g or 0.1 percent
| G2 |  |  | 0.1 g |  | 0.2 or 0.1 percent
| G5 |  |  | 1 g |  | 2 g or 0.1 percent
| G20 |  |  | 5 g |  | 5 g or 0.1 percent
| G100 |  |  | 20 g |  | 20 g or 0.1 percent
| Weight Applied (g) |  | Scale Reading (g) |  | Tolerance (g) |  |  | Pass/Fail
| 0.1 |  | 0.1 |  | 0.2
| 1 |  | 1 |  | 0.2
| 10 |  | 10 |  | 0.2
| 100 |  | 100 |  | 0.2
| 100.1 |  | 100.1 |  | 0.2
| 200 |  | 200 |  | 0.2
| 500 |  | 500 |  | 0.5
| 500.1 |  | 500.1 |  | 0.5
| 1000 |  | 1000 |  | 1
| 2000 |  | 2000 |  | 2
| 2000.1 |  | 2000.1 |  | 2
| 5000 |  | 5000.1 |  | 5
| 5000.1 |  | 5000.2 |  | 5
| 10000 |  | 10000.1 |  | 10
| 11000 |  | 11000 |  | 11
[9 empty] | Pass/Fail: | Pass
Comments:
[9 empty] | Initial By: | sgw

## Scale (G2-CL2)
Scale Calibration Worksheet
| Equipment ID: | CL - 2 |  |  |  |  | Date: | 2008-09-09
| Manufacturer: | Ohaus |  |  |  |  | Performed by: | M. Fleck
| Model #: | EP12001
| Serial #: | 1128061463 |  |  |  |  | Next Calibration Due: | September 2009
| UofA ID: | NA |  |  |  |  | Last Calibration: | September 2007
| Location: | CL (220) Right SpG
| Calibration Item: | Verify mass readings for at least 5 points in the range of the balance
Check sensitivity (0.1%) for at least 5 points in the range of the balance
| Calibration Procedure: | AASHTO M-231
| Calibration Equipment: | Class F, Class I, and 10-kg Calibration Weights
Tolerances:
| Class |  | Readability and Sensitivity |  |  |  | Accuracy
| G1 |  |  | 0.01 g |  | 0.02 g or 0.1 percent
| G2 |  |  | 0.1 g |  | 0.2 or 0.1 percent
| G5 |  |  | 1 g |  | 2 g or 0.1 percent
| G20 |  |  | 5 g |  | 5 g or 0.1 percent
| G100 |  |  | 20 g |  | 20 g or 0.1 percent
| Weight Applied (g) |  | Scale Reading (g) |  | Tolerance (g) |  |  | Pass/Fail
| 0.1 |  | 0.1 |  | 0.2
| 1 |  | 1 |  | 0.2
| 10 |  | 10 |  | 0.2
| 100 |  | 100 |  | 0.2
| 100.1 |  | 100.1 |  | 0.2
| 200 |  | 200 |  | 0.2
| 500 |  | 500 |  | 0.5
| 500.1 |  | 500.1 |  | 0.5
| 1000 |  | 1000 |  | 1
| 2000 |  | 2000 |  | 2
| 2000.1 |  | 2000.1 |  | 2
| 5000 |  | 5000 |  | 5
| 5000.1 |  | 5000.1 |  | 5
| 10000 |  | 10000 |  | 10
| 11000 |  | 11000.1 |  | 11
[9 empty] | Pass/Fail: | Pass
Comments:
[9 empty] | Initial By: | sgw

## Scale (G2-IO)
Scale Calibration Worksheet
| Equipment ID: | Ignition Oven |  |  |  |  | Date: | 2008-09-09
| Manufacturer: | Ohaus |  |  |  |  | Performed by: | M. Fleck
| Model #: | GT 8000
| Serial #: | 945960260909 |  |  |  |  | Next Calibration Due: | September 2009
| UofA ID: | NA |  |  |  |  | Last Calibration: | September 2007
| Location: | HM (225) Ignition Oven
| Calibration Item: | Verify mass readings for at least 5 points in the range of the balance
Check sensitivity (0.1%) for at least 5 points in the range of the balance
| Calibration Procedure: | AASHTO M-231
| Calibration Equipment: | Class F, Class I, and 10-kg Calibration Weights
Tolerances:
| Class |  | Readability and Sensitivity |  |  |  | Accuracy
| G1 |  |  | 0.01 g |  | 0.02 g or 0.1 percent
| G2 |  |  | 0.1 g |  | 0.2 or 0.1 percent
| G5 |  |  | 1 g |  | 2 g or 0.1 percent
| G20 |  |  | 5 g |  | 5 g or 0.1 percent
| G100 |  |  | 20 g |  | 20 g or 0.1 percent
| Weight Applied (g) |  | Scale Reading (g) |  | Tolerance (g) |  |  | Pass/Fail
| 0.1 |  | 0.1 |  | 0.2
| 1 |  | 1 |  | 0.2
| 10 |  | 10 |  | 0.2
| 100 |  | 100 |  | 0.2
| 100.1 |  | 100.1 |  | 0.2
| 200 |  | 200 |  | 0.2
| 500 |  | 500 |  | 0.5
| 500.1 |  | 500.1 |  | 0.5
| 1000 |  | 1000 |  | 1
| 2000 |  | 2000 |  | 2
| 2000.1 |  | 2000.2 |  | 2
| 5000 |  | 5000.2 |  | 5
| 5000.1 |  | 5000.3 |  | 5
| 8000 |  | 8000.3 |  | 8
[9 empty] | Pass/Fail: | Pass
Comments:
[9 empty] | Initial By: | sgw

## Scale (G5-MX1)
Scale Calibration Worksheet
| Equipment ID: | MX - 1 |  |  |  |  | Date: | 2008-09-09
| Manufacturer: | Mettler |  |  |  |  | Performed by: | M. Fleck
| Model #: | SG 32000
| Serial #: | 2115275579 |  |  |  |  | Next Calibration Due: | September 2009
| UofA ID: | NA |  |  |  |  | Last Calibration: | September 2007
| Location: | MX (234) (Right)
| Calibration Item: | Verify mass readings for at least 5 points in the range of the balance
Check sensitivity (0.1%) for at least 5 points in the range of the balance
| Calibration Procedure: | AASHTO M-231
| Calibration Equipment: | Class F, Class I, and 10-kg Calibration Weights
Tolerances:
| Class |  | Readability and Sensitivity |  |  |  | Accuracy
| G1 |  |  | 0.01 g |  | 0.02 g or 0.1 percent
| G2 |  |  | 0.1 g |  | 0.2 or 0.1 percent
| G5 |  |  | 1 g |  | 2 g or 0.1 percent
| G20 |  |  | 5 g |  | 5 g or 0.1 percent
| G100 |  |  | 20 g |  | 20 g or 0.1 percent
| Weight Applied (g) |  | Scale Reading (g) |  | Tolerance (g) |  |  | Pass/Fail
| 1 |  | 1 |  | 2
| 5 |  | 5 |  | 2
| 10 |  | 10 |  | 2
| 50 |  | 50 |  | 2
| 100 |  | 100 |  | 2
| 500 |  | 500 |  | 2
| 501 |  | 501 |  | 2
| 1000 |  | 1000 |  | 2
| 5000 |  | 5000 |  | 5
| 5001 |  | 5001 |  | 5
| 10000 |  | 10000 |  | 10
| 10001 |  | 10001 |  | 10
| 20000 |  | 20000 |  | 20
| 30000 |  | 30002 |  | 30
| 30001 |  | 30003 |  | 30
[9 empty] | Pass/Fail: | Pass
Comments:
[9 empty] | Initial By: | sgw

## Scale (G5-MX2)
Scale Calibration Worksheet
| Equipment ID: | MX - 2 |  |  |  |  | Date: | 2008-09-09
| Manufacturer: | Mettler |  |  |  |  | Performed by: | M. Fleck
| Model #: | SG 32000
| Serial #: | 1120303652 |  |  |  |  | Next Calibration Due: | September 2009
| UofA ID: | NA |  |  |  |  | Last Calibration: | September 2007
| Location: | MX (234) (Left)
| Calibration Item: | Verify mass readings for at least 5 points in the range of the balance
Check sensitivity (0.1%) for at least 5 points in the range of the balance
| Calibration Procedure: | AASHTO M-231
| Calibration Equipment: | Class F, Class I, and 10-kg Calibration Weights
Tolerances:
| Class |  | Readability and Sensitivity |  |  |  | Accuracy
| G1 |  |  | 0.01 g |  | 0.02 g or 0.1 percent
| G2 |  |  | 0.1 g |  | 0.2 or 0.1 percent
| G5 |  |  | 1 g |  | 2 g or 0.1 percent
| G20 |  |  | 5 g |  | 5 g or 0.1 percent
| G100 |  |  | 20 g |  | 20 g or 0.1 percent
| Weight Applied (g) |  | Scale Reading (g) |  | Tolerance (g) |  |  | Pass/Fail
| 1 |  | 1 |  | 2
| 5 |  | 5 |  | 2
| 10 |  | 10 |  | 2
| 50 |  | 50 |  | 2
| 100 |  | 100 |  | 2
| 500 |  | 500 |  | 2
| 501 |  | 501 |  | 2
| 1000 |  | 1000 |  | 2
| 5000 |  | 5000 |  | 5
| 5001 |  | 5001 |  | 5
| 10000 |  | 10001 |  | 10
| 10001 |  | 10002 |  | 10
| 20000 |  | 20000 |  | 20
| 30000 |  | 30000 |  | 30
| 30001 |  | 30001 |  | 30
[9 empty] | Pass/Fail: | Pass
Comments:
[9 empty] | Initial By: | sgw

## Scale (G100-C1)
Scale Calibration Worksheet
| Equipment ID: | C - 1 |  |  |  |  | Date: | 2008-09-09
| Manufacturer: | AND |  |  |  |  | Performed by: | M. Fleck
| Model #: | HV60KGL
| Serial #: | EM7306611 |  |  |  |  | Next Calibration Due: | September 2009
| UofA ID: | NA |  |  |  |  | Last Calibration: | December 2007
| Location: | CL (260)
| Calibration Item: | Verify mass readings for at least 5 points in the range of the balance
Check sensitivity (0.1%) for at least 5 points in the range of the balance
| Calibration Procedure: | AASHTO M-231
| Calibration Equipment: | Class F, Class I, and 10-kg Calibration Weights
Tolerances:
| Class |  | Readability and Sensitivity |  |  |  | Accuracy
| G1 |  |  | 0.01 g |  | 0.02 g or 0.1 percent
| G2 |  |  | 0.1 g |  | 0.2 or 0.1 percent
| G5 |  |  | 1 g |  | 2 g or 0.1 percent
| G20 |  |  | 5 g |  | 5 g or 0.1 percent
| G100 |  |  | 20 g |  | 20 g or 0.1 percent
| Weight Applied (g) |  | Scale Reading (g) |  | Tolerance (g) |  |  | Pass/Fail
| 20 |  | 20 |  | 20
| 500 |  | 500 |  | 20
| 520 |  | 520 |  | 20
| 1000 |  | 1000 |  | 20
| 5000 |  | 5000 |  | 20
| 10000 |  | 10000 |  | 20
| 10020 |  | 10020 |  | 20
| 15000 |  | 15000 |  | 20
| 30000 |  | 30000 |  | 30
| 30020 |  | 30020 |  | 30
| 45000 |  | 45000 |  | 45
| 45020 |  | 45020 |  | 45
| 60000 |  | 60000 |  | 60
[9 empty] | Pass/Fail: | Pass
Comments:
[9 empty] | Initial By: | sgw

## Scale (G100-C2)
Scale Calibration Worksheet
| Equipment ID: | C - 2 |  |  |  |  | Date: | 2008-09-09
| Manufacturer: | AND |  |  |  |  | Performed by: | M. Fleck
| Model #: | FG60KAM
| Serial #: | Q1911021 |  |  |  |  | Next Calibration Due: | September 2009
| UofA ID: | NA |  |  |  |  | Last Calibration: | NA (new)
| Location: | STO (216)
| Calibration Item: | Verify mass readings for at least 5 points in the range of the balance
Check sensitivity (0.1%) for at least 5 points in the range of the balance
| Calibration Procedure: | AASHTO M-231
| Calibration Equipment: | Class F, Class I, and 10-kg Calibration Weights
Tolerances:
| Class |  | Readability and Sensitivity |  |  |  | Accuracy
| G1 |  |  | 0.01 g |  | 0.02 g or 0.1 percent
| G2 |  |  | 0.1 g |  | 0.2 or 0.1 percent
| G5 |  |  | 1 g |  | 2 g or 0.1 percent
| G20 |  |  | 5 g |  | 5 g or 0.1 percent
| G100 |  |  | 20 g |  | 20 g or 0.1 percent
| Weight Applied (g) |  | Scale Reading (g) |  | Tolerance (g) |  |  | Pass/Fail
| 20 |  | 20 |  | 20
| 500 |  | 500 |  | 20
| 520 |  | 520 |  | 20
| 1000 |  | 1000 |  | 20
| 5000 |  | 5000 |  | 20
| 10000 |  | 10000 |  | 20
| 10020 |  | 10020 |  | 20
| 15000 |  | 15000 |  | 20
| 30000 |  | 30000 |  | 30
| 30020 |  | 30020 |  | 30
| 45000 |  | 45000 |  | 45
| 45020 |  | 45020 |  | 45
| 60000 |  | 60000 |  | 60
[9 empty] | Pass/Fail: | Pass
Comments:
[9 empty] | Initial By: | sgw