## WB WS
THERMOMETER CALIBRATION WORKSHEET
THERM. | RANGE | ACC. | LOCATION | CALIB | READ | CALIB | READ | CALIB | READ | CALIB | READ | CALIB | READ
WB 1 | 0 - 120 °F | 1° | STO(216)
WB 2 | 0 - 120 °F | 1° | STO(216)
WB 3 | 0 - 120 °F | 1° | CL(220) #1
WB 4 | 0 - 120 °F | 1° | CL(220) #2
WB 5 | 0 - 120 °F | 1° | AGG(233) #1
WB 6 | 0 - 120 °F | 1° | AGG(233) #2
SP 7 | 30 - 300 °F | 2° | STO(216)
CA 2 | 25 - 125 °F | 1° | STO (218)
WB | 63 F |  | CALIBRATION THERMOMETER |  |  |  | 0.2 ° F
SP | 63F/65F/67F |  | CALIBRATION THERMOMETER |  |  |  | 0.2 / 0.5 ° F |  | (must have one reading over 265° F)
CA 2 | 63F / 64F |  | CALIBRATION THERMOMETER |  |  |  | 0.2 ° F

## Fluke D1
Thermometer Calibration Worksheet
Equipment ID: | Thermometer D 1 |  | Date: | 4/25/12
Manufacturer: | Fluke |  | Performed by: | Mary Fleck
Model #: | 54 - 2
Serial #: | 99340004 |  | Next Calibration Due: | April 2013
UofA ID: | Research Digital |  | Last Calibration: | April 2011
Location: | Stor (218)
|  |  | Typical Set Temperature: | 70 - 350° F | deg.
| Calibration Item: | Verify temperature readings
| Calibration Procedure: | In-House Procedure #00-9
| Calibration Equipment: | NIST Traceable Partial Immersion Calibration Thermometer
Temperature Measurement (deg.)
| Temp. Setting | Reading | Correction
| 46.2 | 47.9 | -1.7
| 70.6 | 71.1 | -0.5
| 178.5 | 180.6 | -2.1
| 244.6 | 246.4 | -1.8
| 324 | 327.3 | -3.3
| Pass/Fail: | Pass
| Comments: | No correction needed.
Initial By:

## Fluke D2
Thermometer Calibration Worksheet
Equipment ID: | Thermometer D 2 |  | Date: | 4/25/12
Manufacturer: | Fluke |  | Performed by: | Mary Fleck
Model #: | 54 - 2
Serial #: | 12790001 |  | Next Calibration Due: | April 2013
UofA ID: | Roselie Digital |  | Last Calibration: | April 2011
Location: | Roselie
|  |  | Typical Set Temperature: | 70 - 350° F | deg.
| Calibration Item: | Verify temperature readings
| Calibration Procedure: | In-House Procedure #00-9
| Calibration Equipment: | NIST Traceable Partial Immersion Calibration Thermometer
Temperature Measurement (deg.)
| Temp. Setting | Reading | Correction
| 46.2 | 46 | 0.2
| 70.6 | 70 | 0.6
| 178.5 | 179.3 | -0.8
| 244.6 | 244.3 | 0.3
| 324 | 324.6 | -0.6
| Pass/Fail: | Pass
| Comments: | No correction needed.
Initial By:

## Kessler 1
Thermometer Calibration Worksheet
Equipment ID: | Thermometer WB 1 |  | Date: | 4/24/12
Manufacturer: | Kessler |  | Performed by: | Deshinka
Model #: | MA 242
Serial #: | NA |  | Next Calibration Due: | April 2013
UofA ID: | 0 - 120° F (1° F) |  | Last Calibration: | April 2011
Location: | SpG 2 (CL 220)
|  |  | Typical Set Temperature: | 73.4° F | deg.
| Calibration Item: | Verify temperature readings
| Calibration Procedure: | In-House Procedure #00-9
| Calibration Equipment: | NIST Traceable Calibration Thermometer, ASTM 63F, cat#1322
Temperature Measurement (deg.)
| Temp. Setting | Reading | Correction
| 72.5 | 73 | -0.5
| 96.1 | 96.5 | -0.4
| 105.7 | 108 | -2.3
| 89.6 | 91 | -1.4
| 67.5 | 68.5 | -1
| Pass/Fail: | Pass
| Comments: | No correction needed.
Initial By:

## Kessler 2
Thermometer Calibration Worksheet
Equipment ID: | Thermometer WB 2 |  | Date: | 4/24/12
Manufacturer: | Kessler |  | Performed by: | Deshinka
Model #: | MA 242
Serial #: | NA |  | Next Calibration Due: | April 2013
UofA ID: | 0 - 120° F (1° F) |  | Last Calibration: | April 2011
Location: | Stor (216)
|  |  | Typical Set Temperature: | 73.4° F | deg.
| Calibration Item: | Verify temperature readings
| Calibration Procedure: | In-House Procedure #00-9
| Calibration Equipment: | NIST Traceable Calibration Thermometer, ASTM 63F, cat#1322
Temperature Measurement (deg.)
| Temp. Setting | Reading | Correction
| 72.4 | 73 | -0.6
| 96.8 | 98 | -1.2
| 106.2 | 108 | -1.8
| 89.6 | 91 | -1.4
| 67 | 68 | -1
| Pass/Fail: | Pass
| Comments: | No correction needed.
Initial By:

## Kessler 3
Thermometer Calibration Worksheet
Equipment ID: | Thermometer WB 3 |  | Date: | 4/24/12
Manufacturer: | Kessler |  | Performed by: | Deshinka
Model #: | MA 242
Serial #: | NA |  | Next Calibration Due: | April 2013
UofA ID: | 0 - 120° F (1° F) |  | Last Calibration: | April 2011
Location: | SpG 1 (CL 220)
|  |  | Typical Set Temperature: | 77° F | deg.
| Calibration Item: | Verify temperature readings
| Calibration Procedure: | In-House Procedure #00-9
| Calibration Equipment: | NIST Traceable Calibration Thermometer, ASTM 63F, cat#1322
Temperature Measurement (deg.)
| Temp. Setting | Reading | Correction
| 72.3 | 73.3 | -1
| 96.8 | 97.5 | -0.7
| 106.8 | 108 | -1.2
| 89.6 | 91.5 | -1.9
| 67 | 68.5 | -1.5
| Pass/Fail: | Pass
| Comments: | No correction needed.
Initial By:

## Kessler 4
Thermometer Calibration Worksheet
Equipment ID: | Thermometer WB 4 |  | Date: | 4/24/12
Manufacturer: | Kessler |  | Performed by: | Deshinka
Model #: | MA 242
Serial #: | NA |  | Next Calibration Due: | April 2013
UofA ID: | 0 - 120° F (1° F) |  | Last Calibration: | April 2011
Location: | SpG 2 (CL 220)
|  |  | Typical Set Temperature: | 77° F | deg.
| Calibration Item: | Verify temperature readings
| Calibration Procedure: | In-House Procedure #00-9
| Calibration Equipment: | NIST Traceable Calibration Thermometer, ASTM 63F, cat#1322
Temperature Measurement (deg.)
| Temp. Setting | Reading | Correction
0
0
0
0
0
Pass/Fail:
| Comments: | Broken - replaced by
Kessler WB 1
Initial By:

## Kessler 5
Thermometer Calibration Worksheet
Equipment ID: | Thermometer WB 5 |  | Date: | 4/24/12
Manufacturer: | Kessler |  | Performed by: | Deshinka
Model #: | MA 242
Serial #: | NA |  | Next Calibration Due: | April 2013
UofA ID: | 0 - 120° F (1° F) |  | Last Calibration: | April 2011
Location: | SpG 1 (Agg 233)
|  |  | Typical Set Temperature: | 73.4° F | deg.
| Calibration Item: | Verify temperature readings
| Calibration Procedure: | In-House Procedure #00-9
| Calibration Equipment: | NIST Traceable Calibration Thermometer, ASTM 63F, cat#1322
Temperature Measurement (deg.)
| Temp. Setting | Reading | Correction
| 72.2 | 73 | -0.8
| 96.2 | 96.5 | -0.3
| 106 | 106.5 | -0.5
| 89.5 | 90.5 | -1
| 67.5 | 68 | -0.5
| Pass/Fail: | Pass
| Comments: | No correction needed.
Initial By:

## Kessler 6
Thermometer Calibration Worksheet
Equipment ID: | Thermometer WB 6 |  | Date: | 4/24/12
Manufacturer: | Kessler |  | Performed by: | Deshinka
Model #: | MA 242
Serial #: | NA |  | Next Calibration Due: | April 2013
UofA ID: | 0 - 120° F (1° F) |  | Last Calibration: | April 2011
Location: | SpG 1 (Agg 233)
|  |  | Typical Set Temperature: | 73.4° F | deg.
| Calibration Item: | Verify temperature readings
| Calibration Procedure: | In-House Procedure #00-9
| Calibration Equipment: | NIST Traceable Calibration Thermometer, ASTM 63F, cat#1322
Temperature Measurement (deg.)
| Temp. Setting | Reading | Correction
| 72.2 | 72 | 0.2
| 96.4 | 97 | -0.6
| 105.7 | 107 | -1.3
| 89.3 | 90.5 | -1.2
| 67.2 | 67 | 0.2
| Pass/Fail: | Pass
| Comments: | No correction needed.
Initial By:

## Gilson (Dial)(P3)
Thermometer Calibration Worksheet
Equipment ID: | Thermometer (Asphalt) P3 |  |  | Date:
Manufacturer: | Gilson MA - 123 |  |  | Performed by:
Model #: | Gilson MA - 123
Serial #: | NA |  |  | Next Calibration Due:
UofA ID: | 50 - 550° F (5°) |  |  | Last Calibration:
Location: | STO (218)
|  |  |  | Typical Set Temperature: | 300 °F | deg.
| Calibration Item: |  | Verify temperature readings
| Calibration Procedure: |  | In-House Procedure #00-9
| Calibration Equipment: |  | NIST Traceable Calibration Thermometers, ASTM 63F, 67F, 68F
Temperature Measurement (deg.)
| Temp. Setting |  | Reading | Correction
0
0
0
.
Pass/Fail:
| Comments: |  | Not in Use
Initial By:

## Kessler (Dial)(CA 2)
Thermometer Calibration Worksheet
Equipment ID: | Concrete Thermometer 2 |  |  | Date: | 4/24/12
Manufacturer: | Gilson (Kessler) |  |  | Performed by: | Deshinka
Model #: | MA 121F
Serial #: | NA |  |  | Next Calibration Due: | April 2013
UofA ID: | 25 - 125° F (1 °) |  |  | Last Calibration: | April 2011
Location: | Stor (218)
|  |  |  | Typical Set Temperature: | 70 ° F | deg.
| Calibration Item: |  | Verify temperature readings
| Calibration Procedure: |  | In-House Procedure #00-9
| Calibration Equipment: |  | NIST Traceable Calibration Thermometers, ASTM 63F, cat# 1322
Temperature Measurement (deg.)
| Temp. Setting |  | Reading | Correction
| 44.2 |  | 44 | 0.2
| 72.5 |  | 72 | 0.5
| 107.6 |  | 107.5 | 0.1
.
| Pass/Fail: |  | Pass
Comments:
Initial By:

## FPC 62F
Thermometer Calibration for Freezing Point
| Equipment ID: | Thermometer 62F |  | Date: | 2000-10-02
| Manufacturer: | Kessler |  | Initialed By:
| Model #: | SP - 3273
| Serial #: | 397382
| Catalog #: | 1320
| UofA ID: | TCAL - 1
| Location: | STO (218)
| Calibration Item: | Verify freezing point for calibration thermometers
| Calibration Procedure: | ASTM E 77
| Calibration Equipment: | Dewar vessel, distilled ice
Date: | Temp. Reading at Freezing Point: |  | Comments: |  | Performed By: | Initial By:
2000-10-02 | 32.0° F |  |  |  | S. Williams | sgw
2001-10-22 | 32.0° F |  |  |  | S. Williams | sgw
2002-10-11 | 32.0° F |  |  |  | S. Williams | sgw
2003-10-17 | 32.0° F |  |  |  | S. Williams | sgw
2004-10-19 | 32.0° F |  |  |  | S. Williams | sgw
2005-10-17 | 32.0° F |  |  |  | S. Williams | sgw
2006-10-30 | 32.0° F |  |  |  | S. Williams | sgw
2007-10-25 | 32.1° F |  |  |  | S. Williams | sgw
2008-10-30 | 32.1° F |  | Separated - Fixed |  | S. Williams | sgw
2009-10-01 | 32.0° F |  |  |  | M. Fleck | sgw
2010-10-06 | 32.0° F |  |  |  | M. Fleck | sgw
2011-10-31 | 32.0° F |  |  |  | M. Fleck | sgw
2012-10-04 | 32.0° F |  |  |  | M. Fleck

## FPC 63F
Thermometer Calibration for Freezing Point
| Equipment ID: | Thermometer 63F |  | Date: | 2000-10-02
| Manufacturer: | Kessler |  | Initialed By:
| Model #: | SP - 3274
| Serial #: | 393879
| Catalog #: | 1322
| UofA ID: | TCAL - 2
| Location: | STO (218)
| Calibration Item: | Verify freezing point for calibration thermometers
| Calibration Procedure: | ASTM E 77
| Calibration Equipment: | Dewar vessel, distilled ice
Date: | Temp. Reading at Freezing Point: |  | Comments: |  | Performed By: | Initial By:
2000-10-02 | 32.0° F |  |  |  | S. Williams | sgw
2001-10-22 | 32.0° F |  |  |  | S. Williams | sgw
2002-10-11 | 32.0° F |  |  |  | S. Williams | sgw
2003-10-17 | 32.0° F |  |  |  | S. Williams | sgw
2004-10-19 | 32.0° F |  |  |  | S. Williams | sgw
2005-10-17 | 32.0° F |  |  |  | S. Williams | sgw
2006-10-30 | 32.0° F |  |  |  | S. Williams | sgw
2007-10-25 | 32.0° F |  |  |  | S. Williams | sgw
2008-10-30 | 32.0° F |  |  |  | S. Williams | sgw
2009-10-01 | 32.0° F |  |  |  | M. Fleck | sgw
2010-10-06 | 32.0° F |  |  |  | M. Fleck | sgw
2011-10-31 | 32.0° F |  |  |  | M. Fleck | sgw
2012-10-04 | 32.0° F |  |  |  | M. Fleck

## FPC 64F
Thermometer Calibration for Freezing Point
| Equipment ID: | Thermometer 64F |  | Date: | 2003-10-17
| Manufacturer: | Miller & Weber, Inc. |  | Initialed By:
| Model #: | T - 3750
| Serial #: | 3N0464
| Catalog #: | N / A
| UofA ID: | TCAL - 3
| Location: | STO (218)
| Calibration Item: | Verify freezing point for calibration thermometers
| Calibration Procedure: | ASTM E 77
| Calibration Equipment: | Dewar vessel, distilled ice
Date: | Temp. Reading at Freezing Point: |  | Comments: |  | Performed By: | Initial By:
2003-10-17 | 32.0° F |  |  |  | S. Williams | sgw
2004-10-19 | 32.0° F |  |  |  | S. Williams | sgw
2005-10-17 | 32.0° F |  |  |  | S. Williams | sgw
2006-10-30 | 32.0° F |  |  |  | S. Williams | sgw
2007-10-25 | 32.0° F |  |  |  | S. Williams | sgw
2008-10-30 | 32.0° F |  |  |  | S. Williams | sgw
2009-10-01 | 32.0° F |  |  |  | M. Fleck | sgw
2010-10-06 | 32.0° F |  |  |  | M. Fleck | sgw
2011-10-31 | 32.0° F |  |  |  | M. Fleck | sgw
2012-10-04 | 32.0° F |  |  |  | M. Fleck

## FPC 65F
Thermometer Calibration for Freezing Point
| Equipment ID: | Thermometer 65F |  | Date: | 2003-10-17
| Manufacturer: | Miller & Weber, Inc. |  | Initialed By:
| Model #: | T - 3750
| Serial #: | 3P8908
| Catalog #: | N / A
| UofA ID: | TCAL - 4
| Location: | STO (218)
| Calibration Item: | Verify freezing point for calibration thermometers
| Calibration Procedure: | ASTM E 77
| Calibration Equipment: | Dewar vessel, distilled ice
Date: | Temp. Reading at Freezing Point: |  | Comments: |  | Performed By: | Initial By:
2003-10-17 | 32.0° F |  |  |  | S. Williams | sgw
2004-10-19 | 32.0° F |  |  |  | S. Williams | sgw
2005-10-17 | 32.0° F |  |  |  | S. Williams | sgw
2006-10-30 | 32.0° F |  |  |  | S. Williams | sgw
2007-10-25 | 32.0° F |  |  |  | S. Williams | sgw
2008-10-30 | 32.0° F |  |  |  | S. Williams | sgw
2009-10-01 | 32.0° F |  |  |  | M. Fleck | sgw
2010-10-06 | 32.0° F |  |  |  | M. Fleck | sgw
2011-10-31 | 32.0° F |  |  |  | M. Fleck | sgw
2012-10-04 | 32.0° F |  |  |  | M. Fleck

## FPC 66F
Thermometer Calibration for Freezing Point
| Equipment ID: | Thermometer 66F |  | Date: | 2001-10-22
| Manufacturer: | Miller & Weber, Inc. |  | Initialed By:
| Model #: | ASTM 66F
| Serial #: | 3G4726
| Catalog #: | N / A
| UofA ID: | TCAL - 5
| Location: | STO (218)
| Calibration Item: | Verify freezing point for calibration thermometers
| Calibration Procedure: | ASTM E 77
| Calibration Equipment: | Dewar vessel, distilled ice
Date: | Temp. Reading at Freezing Point: |  | Comments: |  | Performed By: | Initial By:
2001-10-22 | 32.0° F |  |  |  | S. Williams | sgw
2002-10-11 | 32.0° F |  |  |  | S. Williams | sgw
2003-10-17 | 32.0° F |  |  |  | S. Williams | sgw
2004-10-19 | 32.0° F |  |  |  | S. Williams | sgw
2005-10-17 | 32.0° F |  |  |  | S. Williams | sgw
2006-10-30 | 32.0° F |  |  |  | S. Williams | sgw
2007-10-25 | 32.0° F |  |  |  | S. Williams | sgw
2008-10-30 | 32.0° F |  |  |  | S. Williams | sgw
2009-10-01 | 32.0° F |  |  |  | M. Fleck | sgw
2010-10-06 | 32.0° F |  |  |  | M. Fleck | sgw
2011-10-31 | 32.0° F |  |  |  | M. Fleck | sgw
2012-10-04 | 32.0° F |  |  |  | M. Fleck

## FPC 67Fa
Thermometer Calibration for Freezing Point
| Equipment ID: | Thermometer 67F |  | Date: | 2000-10-02
| Manufacturer: | Kessler |  | Initialed By:
| Model #: | SP - 3804
| Serial #: | 396898
| Catalog #: | 1330
| UofA ID: | TCAL - 6a
| Location: | SR (218)
| Calibration Item: | Verify freezing point for calibration thermometers
| Calibration Procedure: | ASTM E 77
| Calibration Equipment: | Dewar vessel, distilled ice
Date: | Temp. Reading at Freezing Point: |  | Comments: |  | Performed By: | Initial By:
2000-10-02 | 32.0° F |  |  |  | S. Williams | sgw
2001-10-22 | 32.0° F |  |  |  | S. Williams | sgw
2002-10-11 | 32.0° F |  |  |  | S. Williams | sgw
2003-10-17 | 32.0° F |  |  |  | S. Williams | sgw
2004-10-19 | 32.0° F |  |  |  | S. Williams | sgw
2005-10-17 | 32.0° F |  |  |  | S. Williams | sgw
2006-10-30 | 32.0° F |  |  |  | S. Williams | sgw
2007-10-25 | 32.0° F |  |  |  | S. Williams | sgw
2008-10-30 | 32.5° F |  | -0.5 Correction Tag |  | S. Williams | sgw
2009-10-01 | 32.5° F |  | -0.5 Correction Tag |  | M. Fleck | sgw
2010-10-06 | 32.5° F |  | -0.5 Correction Tag |  | M. Fleck | sgw
2011-10-31 | 32.5° F |  | -0.5 Correction Tag |  | M. Fleck | sgw
2012-10-04 | 32.5° F |  | -0.5 Correction Tag |  | M. Fleck

## FPC 67Fb
Thermometer Calibration for Freezing Point
| Equipment ID: | Thermometer 67F |  | Date: | 2003-10-17
| Manufacturer: | Kessler |  | Initialed By:
| Model #: | MA - 435FT - C
| Serial #: | 479390
| Catalog #: | N / A
| UofA ID: | TCAL - 6b
| Location: | SR (218)
| Calibration Item: | Verify freezing point for calibration thermometers
| Calibration Procedure: | ASTM E 77
| Calibration Equipment: | Dewar vessel, distilled ice
Date: | Temp. Reading at Freezing Point: |  | Comments: |  | Performed By: | Initial By:
2003-10-17 | 32.0° F |  |  |  | S. Williams | sgw
2004-10-19 | 32.0° F |  |  |  | S. Williams | sgw
2005-10-17 | 32.0° F |  |  |  | S. Williams | sgw
2006-10-30 | 32.0° F |  |  |  | S. Williams | sgw
2007-10-25 | 32.0° F |  |  |  | S. Williams | sgw
2008-10-30 | 32.0° F |  |  |  | S. Williams | sgw
2009-10-01 | 32.0° F |  |  |  | M. Fleck | sgw
2010-10-06 | 32.0° F |  |  |  | M. Fleck | sgw
2011-10-31 | 32.0° F |  |  |  | M. Fleck | sgw
2012-10-04 | 32.0° F |  |  |  | M. Fleck

## FPC 68Fb
Thermometer Calibration for Freezing Point
| Equipment ID: | Thermometer 68F |  | Date: | 2010-03-01
| Manufacturer: | Kessler |  | Initialed By:
| Model #: | SP - 3860 |  |  | Replaces TCAL - 7
| Serial #: | 583222
| Catalog #: | 1332C
| UofA ID: | TCAL - 7b
| Location: | SR (218)
| Calibration Item: | Verify freezing point for calibration thermometers
| Calibration Procedure: | ASTM E 77
| Calibration Equipment: | Dewar vessel, distilled ice
Date: | Temp. Reading at Freezing Point: |  | Comments: |  | Performed By: | Initial By:
2010-03-01 | 32.0° F |  |  |  | M. Fleck | sgw
2010-10-06 | 32.0° F |  |  |  | M. Fleck | sgw
2011-10-31 | 32.0° F |  |  |  | M. Fleck | sgw
2012-10-04 | 32.0° F |  |  |  | M. Fleck

## FPC Partial (1)
Thermometer Calibration for Freezing Point
| Equipment ID: | Part. Imm. Therm.(1) |  | Date: | 2010-12-01
| Manufacturer: | H - B Instrument |  | Initialed By:
| Model #: | WB Thermometer |  |  | New
| Serial #: | 0789
| Catalog #: | 89095-962
| UofA ID: | TCAL - 8
| Location: | SW (218)
| Calibration Item: | Verify freezing point for calibration thermometers
| Calibration Procedure: | ASTM E 77
| Calibration Equipment: | Dewar vessel, distilled ice
Date: | Temp. Reading at Freezing Point: |  | Comments: |  | Performed By: | Initial By:
2010-11-12 | 32.0° F |  | New - See NIST Certificate |  | Leah Smith | sgw
2011-10-31 | 32.0° F |  |  |  | M. Fleck | sgw
2012-10-04 | 32.0° F |  |  |  | M. Fleck

## FPC Partial (2)
Thermometer Calibration for Freezing Point
| Equipment ID: | Part. Imm. Therm.(2) |  | Date: | 2010-11-11
| Manufacturer: | H - B Instrument |  | Initialed By:
| Model #: | HB 20497 |  |  | New
| Serial #: | 7540
| Catalog #: | 61027-341
| UofA ID: | TCAL - 9
| Location: | SW (218)
| Calibration Item: | Verify freezing point for calibration thermometers
| Calibration Procedure: | ASTM E 77
| Calibration Equipment: | Dewar vessel, distilled ice
Date: | Temp. Reading at Freezing Point: |  | Comments: |  | Performed By: | Initial By:
2010-11-11 | 32.0° F |  | New - See NIST Certificate |  | Leah Smith | sgw
2011-10-31 | 32.0° F |  |  |  | M. Fleck | sgw
2012-10-04 | 32.0° F |  |  |  | M. Fleck

## ASTM
ASTM Thermometer Guide
Model # |  | Range | Accuracy | CTTP #
ASTM | 62 F | -36° TO 35° F | 0.2° F | T-CAL | 1
ASTM | 63 F | 18° TO 89° F | 0.2° F | T-CAL | 2
ASTM | 64 F | 77° TO 131° F | 0.2° F | T-CAL | 3
ASTM | 65 F | 122° TO 176° F | 0.2° F | T-CAL | 4
ASTM | 66 F | 167° TO 221° F | 0.2° F | T-CAL | 5
ASTM | 67 F | 203° TO 311° F | 0.5° F | T-CAL | 6A* & 6B
ASTM | 68 F | 293° TO 401° F | 0.5° F | T-CAL | 7b
H-B | Partial | 30° - 124° F | 0.2° F | T-CAL | 8
H-B | Partial | 30° - 395° F | 0.5° F | T-CAL | 9
* T - CAL 6A reads 0.5° high

## FP WS
Freezing Point Calibration WS
Tested By : ___________ |  | Date :
| Thermometer | ASTM # | Reading
| T-CAL 1 | 62 F
| T-CAL 2 | 63 F
| T-CAL 3 | 64 F
| T-CAL 4 | 65 F
| T-CAL 5 | 66 F
| T-CAL 6 A | 67 F
| T-CAL 6 B | 67 F
| T-CAL 7 B | 68 F
| T-CAL 8 | Part (1)
| T-CAL 9 | Part (2)