## Oven HM1
Oven Calibration Worksheet
Equipment ID: | HM - 1 |  | Serial #: | 170436
Manufacturer: | Despatch |  | UofA ID: | 255801
Model #: | LBB2 - 18 - 1 |  | Location: |  | Hot Mix Lab (225)
Calibration Item: | Verify temperature readings |  |  | Typical Set Temperature: | 300 ° F
Calibration Procedure: | In-House Procedure #00-1 |  |  | Tolerance : | ± 5 ° F
Calibration Equipment: | Thermometer ASTM | 67F / 68F
Temp. Readings
Date: | 2012-10-31 | Set Temp. | 1 | 2 | 3
Performed by: | M. Fleck | 270 | 270.4 | 274.9 | 271.1
Next Calibration Due: | October 2013 | 300 | 301.9 | 302.4 | 302.8
Last Calibration: | October 2011 | 328 | 325.5 | 326 | 326.5
Requirement: | ± 5 ° F
| Pass/Fail: | Pass |  |  | Initial by: | sgw
| Comments: | Digital setpoints.
3° F SP Offset Note added to Oven
Temp. Readings
Date: | 2013-10-11 | Set Temp. | 1 | 2 | 3
Performed by: | M. Fleck | 230 | 233.5 | 233.7 | 234
Next Calibration Due: | October 2014 | 280 | 282.4 | 282.7 | 282
Last Calibration: | October 2012 | 315 | 313.1 | 314.1 | 313.8
Requirement: | ± 5 ° F
| Pass/Fail: | Pass |  |  | Initial by: | sgw
| Comments: | Digital setpoints.
Temp. Readings
Date: |  | Set Temp. | 1 | 2 | 3
Performed by: |  | 250
Next Calibration Due: | October 2015 | 290
Last Calibration: | October 2013 | 325
Requirement: | ± 5 ° F
| Pass/Fail: |  |  |  | Initial by:
| Comments: | Digital setpoints.

## Oven HM-A
Oven Calibration Worksheet
Equipment ID: | HM - A |  | Serial #: | 9602244
Manufacturer: | Quincy Lab |  | UofA ID: | 234501
Model #: | 21 - 350 |  | Location: |  | Hot Mix Lab (225)
Calibration Item: | Verify temperature readings |  |  | Typical Set Temperature: | 200 ° F
Calibration Procedure: | In-House Procedure #00-1 |  |  | Tolerance : | ± 9 °
Calibration Equipment: | Thermometer ASTM | 67 F
Temp. Readings
Date: | 2011-10-31 | Set Temp. | 1 | 2 | 3
Performed by: | M. Fleck | 200 | 213 | 212 | 198.4
Next Calibration Due: | October 2012
Last Calibration: | October 2010
Requirement: | ± 9 ° F
| Pass/Fail: | Pass |  |  | Initial by: | sgw
| Comments: | Dial Set @ 200° mark.
Constant temperature.
Temp. Readings
Date: | 2012-10-31 | Set Temp. | 1 | 2 | 3
Performed by: | M. Fleck | 200 | 188.3 | 197.2 | 205.8
Next Calibration Due: | October 2013
Last Calibration: | October 2011
Requirement: | ± 9 ° F
| Pass/Fail: | Pass |  |  | Initial by: | sgw
| Comments: | Dial Set @ 200° mark.
Constant temperature.
Temp. Readings
Date: | 2013-10-14 | Set Temp. | 1 | 2 | 3
Performed by: | Charles S. | 200 | 187.4 | 183.2 | 187.6
Next Calibration Due: | October 2014
Last Calibration: | October 2012
Requirement: | ± 9 ° F
| Pass/Fail: | Fail |  |  | Initial by: | sgw
| Comments: | Dial Set @ 200° mark.
Constant temperature.
Due to increased variability from oven dial setpoints, temperature setpoints
and required accuracy must be verified by thermometer before each use.
Therefore, yearly calibration of this oven will be suspended.

## Oven HM-B
Oven Calibration Worksheet
Equipment ID: | HM - B |  | Serial #: | 9602251
Manufacturer: | Quincy Lab |  | UofA ID: | 234504
Model #: | 21 - 350 |  | Location: |  | Hot Mix Lab (225)
Calibration Item: | Verify temperature readings |  |  | Typical Set Temperature: | 125 ° F
Calibration Procedure: | In-House Procedure #00-1 |  |  | Tolerance : | ± 9 ° F
Calibration Equipment: | Thermometer ASTM | 65 F
Temp. Readings
Date: | 2012-10-23 | Set Temp. | 1 | 2 | 3
Performed by: | M. Fleck | 125 | 130.5 | 124.7 | 119.5
Next Calibration Due: | October 2013
Last Calibration: | October 2011
Requirement: | ± 9 ° F
| Pass/Fail: | Pass |  |  | Initial by: | sgw
| Comments: | Dial set @ 125° F mark.
Constant temperature.
Temp. Readings
Date: | 2013-10-14 | Set Temp. | 1 | 2 | 3
Performed by: | Charles S. | 125 | 138.1 | 136.1 | 131
Next Calibration Due: | October 2014
Last Calibration: | October 2012
Requirement: | ± 9 ° F
| Pass/Fail: | Fail |  |  | Initial by: | sgw
| Comments: | Dial set @ 125° F mark.
Constant temperature.
Temp. Readings
Date: |  | Set Temp. | 1 | 2 | 3
Performed by: |  | 125
Next Calibration Due:
Last Calibration:
Requirement:
| Pass/Fail: |  |  |  | Initial by:
Comments:
Due to increased variability from oven dial setpoints, temperature setpoints
and required accuracy must be verified by thermometer before each use.
Therefore, yearly calibration of this oven will be suspended.

## Oven HM-C
Oven Calibration Worksheet
Equipment ID: | HM - C |  | Serial #: | 9602245
Manufacturer: | Quincy Lab |  | UofA ID: | 234299
Model #: | 21 - 350 |  | Location: |  | Hot Mix Lab (225)
Calibration Item: | Verify temperature readings |  |  | Typical Set Temperature: | 200 ° F
Calibration Procedure: | In-House Procedure #00-1 |  |  | Tolerance : | ± 9 ° F
Calibration Equipment: | Thermometer ASTM | 66 F
Temp. Readings
Date: | 2012-10-24 | Set Temp. | 1 | 2 | 3
Performed by: | M. Fleck | 200 | 196.5 | 205.5 | 214.5
Next Calibration Due: | October 2013
Last Calibration: | October 2011
Requirement: | ± 9 ° F
| Pass/Fail: | Pass |  |  | Initial by: | sgw
| Comments: | Dial Set @ 200° mark.
Constant temperature.
Temp. Readings
Date: | 2013-10-15 | Set Temp. | 1 | 2 | 3
Performed by: | M. Fleck | 200 | 212.8 | 210.1 | 215.9
Next Calibration Due: | October 2014
Last Calibration: | October 2012
Requirement: | ± 9 ° F
| Pass/Fail: | Fail |  |  | Initial by: | sgw
| Comments: | Dial Set @ 200° mark.
Constant temperature.
Temp. Readings
Date: |  | Set Temp. | 1 | 2 | 3
Performed by: |  | 200
Next Calibration Due:
Last Calibration:
Requirement:
| Pass/Fail: |  |  |  | Initial by:
Comments:
Due to increased variability from oven dial setpoints, temperature setpoints
and required accuracy must be verified by thermometer before each use.
Therefore, yearly calibration of this oven will be suspended.

## Oven HM-D
Oven Calibration Worksheet
Equipment ID: | HM - D |  | Serial #: | 9602247
Manufacturer: | Quincy Lab |  | UofA ID: | 234297
Model #: | 21 - 350 |  | Location: |  | Hot Mix Lab (225)
Calibration Item: | Verify temperature readings |  |  | Typical Set Temperature: | 125 ° F
Calibration Procedure: | In-House Procedure #00-1 |  |  | Tolerance : | ± 9 ° F
Calibration Equipment: | Thermometer ASTM | 65 F
Temp. Readings
Date: | 2012-10-03 | Set Temp. | 1 | 2 | 3
Performed by: | M. Fleck
Next Calibration Due: | October 2013
Last Calibration: | October 2011
Requirement: | ± 9 ° F
| Pass/Fail: |  |  |  | Initial by: | sgw
| Comments: | Out of Service
Temp. Readings
Date: | 2013-10-09 | Set Temp. | 1 | 2 | 3
Performed by: | M. Fleck | 125 | 121.7 | 129.5 | 125.2
Next Calibration Due: | October 2014
Last Calibration: | October 2012
Requirement: | ± 9 ° F
| Pass/Fail: | Pass |  |  | Initial by: | sgw
| Comments: | Dial set @ 125 ° F mark.
Constant temperature.
Temp. Readings
Date: |  | Set Temp. | 1 | 2 | 3
Performed by:
Next Calibration Due:
Last Calibration:
Requirement:
| Pass/Fail: |  |  |  | Initial by:
Comments:
Due to increased variability from oven dial setpoints, temperature setpoints
and required accuracy must be verified by thermometer before each use.
Therefore, yearly calibration of this oven will be suspended.

## Oven MX1
Oven Calibration Worksheet
Equipment ID: | MX - 1 |  | Serial #: | 9602246
Manufacturer: | Quincy Lab |  | UofA ID: | 234300
Model #: | 21 - 350 |  | Location: |  | Mixing Lab (234)
Calibration Item: | Verify temperature readings |  |  | Typical Set Temperature: | 230 ° F
Calibration Procedure: | In-House Procedure #00-1 |  |  | Tolerance : | ± 9 ° F
Calibration Equipment: | Thermometer ASTM | 67 F
Temp. Readings
Date: | 2012-10-22 | Set Temp. | 1 | 2 | 3
Performed by: | M. Fleck
Next Calibration Due: | October 2013 | 230 | 224.5 | 228 | 233.4
Last Calibration: | October 2011
Requirement: | ± 9 ° F
| Pass/Fail: | Pass |  |  | Initial by: | sgw
| Comments: | Dial set @ 230° mark.
Constant temperature.
Temp. Readings
Date: | 2013-10-09 | Set Temp. | 1 | 2 | 3
Performed by: | M. Fleck
Next Calibration Due: | October 2014 | 230 | 232.1 | 235.9 | 225.6
Last Calibration: | October 2012
Requirement: | ± 9 ° F
| Pass/Fail: | Pass |  |  | Initial by:
| Comments: | Dial set @ 230° mark.
Constant temperature.
Temp. Readings
Date: |  | Set Temp. | 1 | 2 | 3
Performed by:
Next Calibration Due: | October 2015 | 230
Last Calibration: | October 2013
Requirement: | ± 9 ° F
| Pass/Fail: |  |  |  | Initial by:
| Comments: | Dial set @ 230° mark.
Constant temperature.
Due to increased variability from oven dial setpoints, temperature setpoints
and required accuracy must be verified by thermometer before each use.
Therefore, yearly calibration of this oven will be suspended.

## Oven MX4
Oven Calibration Worksheet
Equipment ID: | MX - 4 |  | Serial #: | 3064606
Manufacturer: | Sheldon Mfg., Inc. |  | UofA ID: | N / A
Model #: | 1370FZZ |  | Location: |  | Mixing Lab (234)
Calibration Item: | Verify temperature readings |  |  | Typical Set Temperature: | 330 ° F
Calibration Procedure: | In-House Procedure #00-1 |  |  | Tolerance : | ± 5 ° F
Calibration Equipment: | Thermometer ASTM | 68 F
Temp. Readings
Date: | 2012-10-22 | Set Temp. | 1 | 2 | 3
Performed by: | M. Fleck | 200 | 202.7 | 198.1 | 195.1
Next Calibration Due: | October 2013 | 280 | 277.7 | 274 | 271.3
Last Calibration: | October 2011 | 330 | 320.3 | 316 | 312.9
Requirement: | ± 5 ° F
| Pass/Fail: | Pass |  |  | Initial by: | sgw
| Comments: | Calibration sheet added to oven door.
Temp. Readings
Date: | 2013-10-10 | Set Temp. | 1 | 2 | 3
Performed by: | M. Fleck | 300 | 302.9 | 304 | 301.5
Next Calibration Due: | October 2014 | 257 | 254.4 | 256.5 | 258.9
Last Calibration: | October 2012 | 202 | 201.7 | 199.1 | 197.3
Requirement: | ± 5 ° F
| Pass/Fail: | Pass |  |  | Initial by: | sgw
| Comments: | Set temperature is according to calibration sheet
based on set point for oven.
Temp. Readings
Date: |  | Set Temp. | 1 | 2 | 3
Performed by:
Next Calibration Due: | October 2015
Last Calibration: | October 2013
Requirement: | ± 5 ° F
| Pass/Fail: |  |  |  | Initial by:
Comments:

## Oven MX5
Oven Calibration Worksheet
Equipment ID: | MX - 5 |  | Serial #: | 9040305
Manufacturer: | Sheldon Mfg., Inc. |  | UofA ID: | 259372
Model #: | 1685 |  | Location: |  | Mixing Lab (234)
Calibration Item: | Verify temperature readings |  |  | Typical Set Temperature: | 330 ° F
Calibration Procedure: | In-House Procedure #00-1 |  |  | Tolerance : | ± 5 ° F
Calibration Equipment: | Thermometer ASTM | 68 F
Temp. Readings
Date: | 2011-10-21 | Set Temp. | 1 | 2 | 3
Performed by: | Haley M. | 250 | 254.6 | 253.4 | 248.8
Next Calibration Due: | October 2012 | 290 | 291 | 286.3 | 289.4
Last Calibration: | October 2010 | 330 | 328.5 | 330.4 | 332.1
Requirement: | ± 5 ° F
| Pass/Fail: | Pass |  |  | Initial by: | sgw
| Comments: | Digital Set @ 250, 290, 330 °F.
Temp. Readings
Date: | 2012-10-23 | Set Temp. | 1 | 2 | 3
Performed by: | M. Fleck | 250 | 247.7 | 246.9 | 247.3
Next Calibration Due: | October 2013 | 290 | 286.1 | 285.7 | 285.1
Last Calibration: | October 2011 | 330 | 322.7 | 323.2 | 323.8
Requirement: | ± 5 ° F
| Pass/Fail: | Pass |  |  | Initial by: | sgw
| Comments: | Digital Set @ 250, 290, 330 °F.
5° F Correction Tag Added to Oven
Temp. Readings
Date: | 2013-10-09 | Set Temp. | 1 | 2 | 3
Performed by: | M. Fleck | 254 | 252.3 | 252.3 | 252.5
Next Calibration Due: | October 2014 | 300 | 293.6 | 294.1 | 294.7
Last Calibration: | October 2012 | 350 | 340.5 | 340.8 | 340.7
Requirement: | ± 5 ° F
| Pass/Fail: | Pass |  |  | Initial by: | sgw
| Comments: | Digital Set @ 254, 300, 350 °F.
5° F Correction Tag Added to Oven