## Instructions
Scale Calibration Instructions
Ohaus EP12001 |  |  |  |  | AND HV60KGL
| → Main Menu |  |  |  |  | Remove back plate
| Enter |  |  |  |  | Press & hold Cal button
| ↕↔ Calibration |  |  |  |  | Enter (Pan = 0)
| Enter |  |  |  |  | Load 60 kg
| ↕ Span |  |  |  |  | Enter
| Enter |  |  |  |  | Remove mass
| Enter |  |  |  |  | Press Cal button
| Place 10 kg on scales |  |  |  |  | Replace plate
Enter
| Press any key to exit |  |  |  | AND FG 60KAL
Remove back plate
I/O Ohaus GT 8000 | [5 empty] | Press & hold Cal button
| I/O Off |  |  |  |  | Press mode (SPAN)
| Open door (leave open) |  |  |  |  | Load 60 kg
| Remove ceramic plate & put calibration |  |  |  |  | Press Mode
| on ceramic pegs |  |  |  |  | Remove mass
| Hold #4 and turn on I/O until beep |  |  |  |  | Replace plate
Use paper clip to press ON button
| (SPAN) |  |  |  | Metttler SG32000
| Press ON |  |  |  |  | Hold Cal button
| Press ON (load 8 kg) |  |  |  |  | (Int. Cal Set)
| Press ON |  |  |  |  | Scale does it's thing
| Power OFF |  |  |  |  | (this stops cal flashing)
Replace ceramic plate
| . | . | . | .
| Print | Mode | Off | On/Tare

## Scale List
Current Scale List 9/9/09
Scale | Mfg | Model # | Serial # | U of A ID | Class | Capacity | Readability | Date In | Date Out
[6 empty] | (g) | (g) | Service | of Service
G1 - 2 | OHAUS | EOD 120 | E2411119231431 | 250795 | G1 | 4100 | 0.01 | ?
G1 - 3 | OHAUS | GT 400 | 5376 | NA | G1 | 400 | 0.01 | 2008-09-05
G1 - 4 | OHAUS | GT 400 | 5373 | NA | G1 | 400 | 0.01 | 2008-09-05
AG - A | OHAUS | EP12001 | 1126012086 | NA | G2 | 12000 | 0.1 | ?
AG - B | OHAUS | EP12001 | 1125430302 | NA | G2 | 12000 | 0.1 | ?
AG - E | OHAUS | EP12001 | 1125330661 | NA | G2 | 12000 | 0.1 | ?
AG - F | OHAUS | EP12001 | 1126051479 | NA | G2 | 12000 | 0.1 | ?
CL - 1 | OHAUS | EP12001 | 1128061460 | NA | G2 | 12000 | 0.1 | ?
CL - 2 | OHAUS | EP12001 | 1128061463 | NA | G2 | 12000 | 0.1
IGN. OVEN | OHAUS | GT 8000 | 945960260909 | 234502 | G2 | 8000 | 0.01 | ?
MX - 1 | METTLER | SG32000 | 2115275579 | NA | G5 | 32000 | 1 | ?
MX - 2 | METTLER | SG32000 | 1120303652 | NA | G5 | 32000 | 1 | ?
C - 1 | AND | HV60KGL | EM7306611 | NA | G100 | 60000 | 20 | ?
C - 2 | AND | FG60KAM | Q1911021 | NA | G100 | 60000 | 20 | 2008-09-05
Calibrations should cover at least 90% of the scale's capacity. Calibration weights should be verified every 12 years.
Extra Scales in Storage:
| OHAUS | EP12001 | 1126051480 | NA | G2 | 12000 | 0.1
| OHAUS | EP12001 | 1125373087 | NA | G2 | 12000 | 0.1
| OHAUS | EP12001 | 1125373345 | NA | G2 | 12000 | 0.1
| OHAUS | EP12001 | 1126012088 | NA | G2 | 12000 | 0.1

## Scale (G1-2)
Scale Calibration Worksheet
| Equipment ID: | G1 - 2 |  |  |  |  | Date: | 2013-09-24
| Manufacturer: | Ohaus |  |  |  |  | Performed by: | M. Fleck
| Model #: | EOD 120
| Serial #: | E2411119231431 |  |  |  |  | Next Calibration Due: | September 2014
| UofA ID: | [5 empty] | Last Calibration: | September 2012
| Location: | STO (218)
| Calibration Item: | Verify mass readings for at least 5 points in the range of the balance
Check sensitivity (0.1%) for at least 5 points in the range of the balance
| Calibration Procedure: | AASHTO M-231
| Calibration Equipment: | Class F, Class I, and 10-kg Calibration Weights
Tolerances:
| Class |  | Readability and Sensitivity |  |  |  | Accuracy
| G1 |  |  | 0.01 g |  | 0.02 g or 0.1 percent
| G2 |  |  | 0.1 g |  | 0.2 or 0.1 percent
| G5 |  |  | 1 g |  | 2 g or 0.1 percent
| G20 |  |  | 5 g |  | 5 g or 0.1 percent
| G100 |  |  | 20 g |  | 20 g or 0.1 percent
| Weight Applied (g) |  | Scale Reading (g) |  | Tolerance (g) |  |  | Pass/Fail
| 1 |  | 1 |  | 0.02
| 1.01 |  | 1.01 |  | 0.02
| 5 |  | 5 |  | 0.02
| 10 |  | 10 |  | 0.02
| 20 |  | 20 |  | 0.02
| 50 |  | 50 |  | 0.05
| 100 |  | 100.01 |  | 0.1
| 500 |  | 500.01 |  | 0.5
| 500.01 |  | 500.02 |  | 0.5
| 1000 |  | 1000 |  | 1
| 1000.01 |  | 1000.01 |  | 1
| 4000 |  | 4000 |  | 4
| 4000.01 |  | 4000.01 |  | 4
[9 empty] | Pass/Fail: | Pass
Comments:
Initial By:

## Scale (G1-3)
Scale Calibration Worksheet
| Equipment ID: | G1 - 3 |  |  |  |  | Date: | 2013-09-24
| Manufacturer: | Ohaus |  |  |  |  | Performed by: | M. Fleck
| Model #: | GT 400
| Serial #: | 5376 |  |  |  |  | Next Calibration Due: | September 2014
| UofA ID: | NA |  |  |  |  | Last Calibration: | September 2012
| Location: | STO (218)
| Calibration Item: | Verify mass readings for at least 5 points in the range of the balance
Check sensitivity (0.1%) for at least 5 points in the range of the balance
| Calibration Procedure: | AASHTO M-231
| Calibration Equipment: | Class F, Class I, and 10-kg Calibration Weights
Tolerances:
| Class |  | Readability and Sensitivity |  |  |  | Accuracy
| G1 |  |  | 0.01 g |  | 0.02 g or 0.1 percent
| G2 |  |  | 0.1 g |  | 0.2 or 0.1 percent
| G5 |  |  | 1 g |  | 2 g or 0.1 percent
| G20 |  |  | 5 g |  | 5 g or 0.1 percent
| G100 |  |  | 20 g |  | 20 g or 0.1 percent
| Weight Applied (g) |  | Scale Reading (g) |  | Tolerance (g) |  |  | Pass/Fail
| 0.01 |  | 0.01 |  | 0.02
| 1 |  | 1 |  | 0.02
| 5 |  | 5 |  | 0.02
| 10 |  | 10 |  | 0.02
| 10.01 |  | 10.01 |  | 0.02
| 50 |  | 49.99 |  | 0.05
| 50.01 |  | 50 |  | 0.05
| 100 |  | 100.02 |  | 0.1
| 100.01 |  | 100.03 |  | 0.1
| 200 |  | 200.03 |  | 0.2
| 200.01 |  | 200.04 |  | 0.2
| 300 |  | 300.03 |  | 0.3
| 400 |  | 400.05 |  | 0.4
[9 empty] | Pass/Fail: | Pass
Comments:
Initial By:

## Scale (G1-4)
Scale Calibration Worksheet
| Equipment ID: | G1 - 4 |  |  |  |  | Date: | 2013-09-24
| Manufacturer: | Ohaus |  |  |  |  | Performed by: | M. Fleck
| Model #: | GT 400
| Serial #: | 5373 |  |  |  |  | Next Calibration Due: | September 2014
| UofA ID: | NA |  |  |  |  | Last Calibration: | September 2012
| Location: | STO (218)
| Calibration Item: | Verify mass readings for at least 5 points in the range of the balance
Check sensitivity (0.1%) for at least 5 points in the range of the balance
| Calibration Procedure: | AASHTO M-231
| Calibration Equipment: | Class F, Class I, and 10-kg Calibration Weights
Tolerances:
| Class |  | Readability and Sensitivity |  |  |  | Accuracy
| G1 |  |  | 0.01 g |  | 0.02 g or 0.1 percent
| G2 |  |  | 0.1 g |  | 0.2 or 0.1 percent
| G5 |  |  | 1 g |  | 2 g or 0.1 percent
| G20 |  |  | 5 g |  | 5 g or 0.1 percent
| G100 |  |  | 20 g |  | 20 g or 0.1 percent
| Weight Applied (g) |  | Scale Reading (g) |  | Tolerance (g) |  |  | Pass/Fail
| 0.01 |  | 0.01 |  | 0.02
| 1 |  | 1 |  | 0.02
| 5 |  | 5 |  | 0.02
| 10 |  | 10 |  | 0.02
| 10.01 |  | 10.01 |  | 0.02
| 50 |  | 49.99 |  | 0.05
| 50.01 |  | 50 |  | 0.05
| 100 |  | 100.02 |  | 0.1
| 100.01 |  | 100.03 |  | 0.1
| 200 |  | 200.03 |  | 0.2
| 200.01 |  | 200.04 |  | 0.2
| 300 |  | 300.04 |  | 0.3
| 400 |  | 400.05 |  | 0.4
[9 empty] | Pass/Fail: | Pass
Comments:
Initial By:

## Scale (G2-A)
Scale Calibration Worksheet
| Equipment ID: | AG - A |  |  |  |  | Date: | 2013-09-13
| Manufacturer: | Ohaus |  |  |  |  | Performed by: | Abiy G.
| Model #: | EP12001
| Serial #: | 1126012086 |  |  |  |  | Next Calibration Due: | September 2014
| UofA ID: | NA |  |  |  |  | Last Calibration: | September 2012
| Location: | AG (233) Rear Table (West)
| Calibration Item: | Verify mass readings for at least 5 points in the range of the balance
Check sensitivity (0.1%) for at least 5 points in the range of the balance
| Calibration Procedure: | AASHTO M-231
| Calibration Equipment: | Class F, Class I, and 10-kg Calibration Weights
Tolerances:
| Class |  | Readability and Sensitivity |  |  |  | Accuracy
| G1 |  |  | 0.01 g |  | 0.02 g or 0.1 percent
| G2 |  |  | 0.1 g |  | 0.2 or 0.1 percent
| G5 |  |  | 1 g |  | 2 g or 0.1 percent
| G20 |  |  | 5 g |  | 5 g or 0.1 percent
| G100 |  |  | 20 g |  | 20 g or 0.1 percent
| Weight Applied (g) |  | Scale Reading (g) |  | Tolerance (g) |  |  | Pass/Fail
| 0.1 |  | 0.1 |  | 0.2
| 1 |  | 1 |  | 0.2
| 10 |  | 10 |  | 0.2
| 100 |  | 100 |  | 0.2
| 100.1 |  | 100.1 |  | 0.2
| 200 |  | 200 |  | 0.2
| 500 |  | 500 |  | 0.5
| 500.1 |  | 500.1 |  | 0.5
| 1000 |  | 1000 |  | 1
| 2000 |  | 2000 |  | 2
| 2000.1 |  | 2000.1 |  | 2
| 5000 |  | 4999.9 |  | 5
| 5000.1 |  | 5000.1 |  | 5
| 10000 |  | 10000 |  | 10
| 11000 |  | 11000.1 |  | 11
[9 empty] | Pass/Fail: | Pass
Comments:
Initial By:

## Scale (G2-B)
Scale Calibration Worksheet
| Equipment ID: | AG - B |  |  |  |  | Date: | 2013-09-13
| Manufacturer: | Ohaus |  |  |  |  | Performed by: | Abiy G.
| Model #: | EP12001
| Serial #: | 1125430302 |  |  |  |  | Next Calibration Due: | September 2014
| UofA ID: | NA |  |  |  |  | Last Calibration: | September 2012
| Location: | AG (233) Front Table (East)
| Calibration Item: | Verify mass readings for at least 5 points in the range of the balance
Check sensitivity (0.1%) for at least 5 points in the range of the balance
| Calibration Procedure: | AASHTO M-231
| Calibration Equipment: | Class F, Class I, and 10-kg Calibration Weights
Tolerances:
| Class |  | Readability and Sensitivity |  |  |  | Accuracy
| G1 |  |  | 0.01 g |  | 0.02 g or 0.1 percent
| G2 |  |  | 0.1 g |  | 0.2 or 0.1 percent
| G5 |  |  | 1 g |  | 2 g or 0.1 percent
| G20 |  |  | 5 g |  | 5 g or 0.1 percent
| G100 |  |  | 20 g |  | 20 g or 0.1 percent
| Weight Applied (g) |  | Scale Reading (g) |  | Tolerance (g) |  |  | Pass/Fail
| 0.1 |  | 0.1 |  | 0.2
| 1 |  | 1 |  | 0.2
| 10 |  | 10 |  | 0.2
| 100 |  | 100 |  | 0.2
| 100.1 |  | 100.1 |  | 0.2
| 200 |  | 200 |  | 0.2
| 500 |  | 500 |  | 0.5
| 500.1 |  | 500.1 |  | 0.5
| 1000 |  | 1000 |  | 1
| 2000 |  | 2000 |  | 2
| 2000.1 |  | 2000.1 |  | 2
| 5000 |  | 5000 |  | 5
| 5000.1 |  | 5000.1 |  | 5
| 10000 |  | 10000.1 |  | 10
| 11000 |  | 11000.2 |  | 11
[9 empty] | Pass/Fail: | Pass
Comments:
Initial By:

## Scale (G2-E)
Scale Calibration Worksheet
| Equipment ID: | AG - E |  |  |  |  | Date: | 2013-09-13
| Manufacturer: | Ohaus |  |  |  |  | Performed by: | Abiy G.
| Model #: | EP12001
| Serial #: | 1125330661 |  |  |  |  | Next Calibration Due: | September 2014
| UofA ID: | NA |  |  |  |  | Last Calibration: | September 2012
| Location: | AG (233) Front SpG (East)
| Calibration Item: | Verify mass readings for at least 5 points in the range of the balance
Check sensitivity (0.1%) for at least 5 points in the range of the balance
| Calibration Procedure: | AASHTO M-231
| Calibration Equipment: | Class F, Class I, and 10-kg Calibration Weights
Tolerances:
| Class |  | Readability and Sensitivity |  |  |  | Accuracy
| G1 |  |  | 0.01 g |  | 0.02 g or 0.1 percent
| G2 |  |  | 0.1 g |  | 0.2 or 0.1 percent
| G5 |  |  | 1 g |  | 2 g or 0.1 percent
| G20 |  |  | 5 g |  | 5 g or 0.1 percent
| G100 |  |  | 20 g |  | 20 g or 0.1 percent
| Weight Applied (g) |  | Scale Reading (g) |  | Tolerance (g) |  |  | Pass/Fail
| 0.1 |  | 0.1 |  | 0.2
| 1 |  | 1 |  | 0.2
| 10 |  | 10 |  | 0.2
| 100 |  | 100 |  | 0.2
| 100.1 |  | 100.1 |  | 0.2
| 200 |  | 200 |  | 0.2
| 500 |  | 500 |  | 0.5
| 500.1 |  | 500.1 |  | 0.5
| 1000 |  | 1000 |  | 1
| 2000 |  | 2000 |  | 2
| 2000.1 |  | 2000.1 |  | 2
| 5000 |  | 5000 |  | 5
| 5000.1 |  | 5000.1 |  | 5
| 10000 |  | 10000 |  | 10
| 11000 |  | 11000 |  | 11
[9 empty] | Pass/Fail: | Pass
Comments:
Initial By:

## Scale (G2-F)
Scale Calibration Worksheet
| Equipment ID: | AG - F |  |  |  |  | Date: | 2013-09-13
| Manufacturer: | Ohaus |  |  |  |  | Performed by: | Abiy G.
| Model #: | EP12001
| Serial #: | 1126051479 |  |  |  |  | Next Calibration Due: | September 2014
| UofA ID: | NA |  |  |  |  | Last Calibration: | September 2012
| Location: | AG (233) Rear SpG (West)
| Calibration Item: | Verify mass readings for at least 5 points in the range of the balance
Check sensitivity (0.1%) for at least 5 points in the range of the balance
| Calibration Procedure: | AASHTO M-231
| Calibration Equipment: | Class F, Class I, and 10-kg Calibration Weights
Tolerances:
| Class |  | Readability and Sensitivity |  |  |  | Accuracy
| G1 |  |  | 0.01 g |  | 0.02 g or 0.1 percent
| G2 |  |  | 0.1 g |  | 0.2 or 0.1 percent
| G5 |  |  | 1 g |  | 2 g or 0.1 percent
| G20 |  |  | 5 g |  | 5 g or 0.1 percent
| G100 |  |  | 20 g |  | 20 g or 0.1 percent
| Weight Applied (g) |  | Scale Reading (g) |  | Tolerance (g) |  |  | Pass/Fail
| 0.1 |  | 0.1 |  | 0.2
| 1 |  | 1 |  | 0.2
| 10 |  | 10 |  | 0.2
| 100 |  | 100 |  | 0.2
| 100.1 |  | 100.1 |  | 0.2
| 200 |  | 200 |  | 0.2
| 500 |  | 500 |  | 0.5
| 500.1 |  | 500.1 |  | 0.5
| 1000 |  | 1000 |  | 1
| 2000 |  | 2000 |  | 2
| 2000.1 |  | 2000.1 |  | 2
| 5000 |  | 5000 |  | 5
| 5000.1 |  | 5000.1 |  | 5
| 10000 |  | 10000 |  | 10
| 11000 |  | 11000.1 |  | 11
[9 empty] | Pass/Fail: | Pass
Comments:
Initial By:

## Scale (G2-CL1)
Scale Calibration Worksheet
| Equipment ID: | CL - 1 |  |  |  |  | Date: | 2013-09-13
| Manufacturer: | Ohaus |  |  |  |  | Performed by: | Abiy G.
| Model #: | EP12001
| Serial #: | 1128061460 |  |  |  |  | Next Calibration Due: | September 2014
| UofA ID: | NA |  |  |  |  | Last Calibration: | September 2012
| Location: | CL (220) Left SpG
| Calibration Item: | Verify mass readings for at least 5 points in the range of the balance
Check sensitivity (0.1%) for at least 5 points in the range of the balance
| Calibration Procedure: | AASHTO M-231
| Calibration Equipment: | Class F, Class I, and 10-kg Calibration Weights
Tolerances:
| Class |  | Readability and Sensitivity |  |  |  | Accuracy
| G1 |  |  | 0.01 g |  | 0.02 g or 0.1 percent
| G2 |  |  | 0.1 g |  | 0.2 or 0.1 percent
| G5 |  |  | 1 g |  | 2 g or 0.1 percent
| G20 |  |  | 5 g |  | 5 g or 0.1 percent
| G100 |  |  | 20 g |  | 20 g or 0.1 percent
| Weight Applied (g) |  | Scale Reading (g) |  | Tolerance (g) |  |  | Pass/Fail
| 0.1 |  | 0.1 |  | 0.2
| 1 |  | 1 |  | 0.2
| 10 |  | 10 |  | 0.2
| 100 |  | 100 |  | 0.2
| 100.1 |  | 100.1 |  | 0.2
| 200 |  | 200 |  | 0.2
| 500 |  | 500 |  | 0.5
| 500.1 |  | 500.1 |  | 0.5
| 1000 |  | 1000 |  | 1
| 2000 |  | 2000 |  | 2
| 2000.1 |  | 2000.1 |  | 2
| 5000 |  | 5000 |  | 5
| 5000.1 |  | 5000 |  | 5
| 10000 |  | 10000 |  | 10
| 11000 |  | 11000 |  | 11
[9 empty] | Pass/Fail: | Pass
Comments:
Initial By:

## Scale (G2-CL2)
Scale Calibration Worksheet
| Equipment ID: | CL - 2 |  |  |  |  | Date: | 2013-09-13
| Manufacturer: | Ohaus |  |  |  |  | Performed by: | Abiy G.
| Model #: | EP12001
| Serial #: | 1128061463 |  |  |  |  | Next Calibration Due: | September 2014
| UofA ID: | NA |  |  |  |  | Last Calibration: | September 2012
| Location: | CL (220) Right SpG
| Calibration Item: | Verify mass readings for at least 5 points in the range of the balance
Check sensitivity (0.1%) for at least 5 points in the range of the balance
| Calibration Procedure: | AASHTO M-231
| Calibration Equipment: | Class F, Class I, and 10-kg Calibration Weights
Tolerances:
| Class |  | Readability and Sensitivity |  |  |  | Accuracy
| G1 |  |  | 0.01 g |  | 0.02 g or 0.1 percent
| G2 |  |  | 0.1 g |  | 0.2 or 0.1 percent
| G5 |  |  | 1 g |  | 2 g or 0.1 percent
| G20 |  |  | 5 g |  | 5 g or 0.1 percent
| G100 |  |  | 20 g |  | 20 g or 0.1 percent
| Weight Applied (g) |  | Scale Reading (g) |  | Tolerance (g) |  |  | Pass/Fail
| 0.1 |  | 0.1 |  | 0.2
| 1 |  | 1 |  | 0.2
| 10 |  | 10 |  | 0.2
| 100 |  | 100 |  | 0.2
| 100.1 |  | 100.1 |  | 0.2
| 200 |  | 200 |  | 0.2
| 500 |  | 500 |  | 0.5
| 500.1 |  | 500.1 |  | 0.5
| 1000 |  | 1000 |  | 1
| 2000 |  | 2000.1 |  | 2
| 2000.1 |  | 2000.2 |  | 2
| 5000 |  | 5000.1 |  | 5
| 5000.1 |  | 5000.2 |  | 5
| 10000 |  | 10000.1 |  | 10
| 11000 |  | 11000.1 |  | 11
[9 empty] | Pass/Fail: | Pass
Comments:
Initial By:

## Scale (G2-IO)
Scale Calibration Worksheet
| Equipment ID: | Ignition Oven |  |  |  |  | Date: | 2013-09-24
| Manufacturer: | Ohaus |  |  |  |  | Performed by: | M. Fleck
| Model #: | GT 8000
| Serial #: | 945960260909 |  |  |  |  | Next Calibration Due: | September 2014
| UofA ID: | NA |  |  |  |  | Last Calibration: | September 2012
| Location: | HM (225) Ignition Oven
| Calibration Item: | Verify mass readings for at least 5 points in the range of the balance
Check sensitivity (0.1%) for at least 5 points in the range of the balance
| Calibration Procedure: | AASHTO M-231
| Calibration Equipment: | Class F, Class I, and 10-kg Calibration Weights
Tolerances:
| Class |  | Readability and Sensitivity |  |  |  | Accuracy
| G1 |  |  | 0.01 g |  | 0.02 g or 0.1 percent
| G2 |  |  | 0.1 g |  | 0.2 or 0.1 percent
| G5 |  |  | 1 g |  | 2 g or 0.1 percent
| G20 |  |  | 5 g |  | 5 g or 0.1 percent
| G100 |  |  | 20 g |  | 20 g or 0.1 percent
| Weight Applied (g) |  | Scale Reading (g) |  | Tolerance (g) |  |  | Pass/Fail
| 0.1 |  | 0.1 |  | 0.2
| 1 |  | 1 |  | 0.2
| 10 |  | 10 |  | 0.2
| 100 |  | 100 |  | 0.2
| 100.1 |  | 100.1 |  | 0.2
| 200 |  | 200 |  | 0.2
| 500 |  | 500 |  | 0.5
| 500.1 |  | 500.1 |  | 0.5
| 1000 |  | 1000.1 |  | 1
| 2000 |  | 2000.1 |  | 2
| 2000.1 |  | 2000.2 |  | 2
| 5000 |  | 5000.2 |  | 5
| 5000.1 |  | 5000.3 |  | 5
| 8000 |  | 8000.4 |  | 8
[9 empty] | Pass/Fail: | Pass
Comments:
Initial By:

## Scale (G5-MX1)
Scale Calibration Worksheet
| Equipment ID: | MX - 1 |  |  |  |  | Date: | 2013-09-24
| Manufacturer: | Mettler |  |  |  |  | Performed by: | M. Fleck
| Model #: | SG 32000
| Serial #: | 2115275579 |  |  |  |  | Next Calibration Due: | September 2014
| UofA ID: | NA |  |  |  |  | Last Calibration: | September 2012
| Location: | MX (234) (Right)
| Calibration Item: | Verify mass readings for at least 5 points in the range of the balance
Check sensitivity (0.1%) for at least 5 points in the range of the balance
| Calibration Procedure: | AASHTO M-231
| Calibration Equipment: | Class F, Class I, and 10-kg Calibration Weights
Tolerances:
| Class |  | Readability and Sensitivity |  |  |  | Accuracy
| G1 |  |  | 0.01 g |  | 0.02 g or 0.1 percent
| G2 |  |  | 0.1 g |  | 0.2 or 0.1 percent
| G5 |  |  | 1 g |  | 2 g or 0.1 percent
| G20 |  |  | 5 g |  | 5 g or 0.1 percent
| G100 |  |  | 20 g |  | 20 g or 0.1 percent
| Weight Applied (g) |  | Scale Reading (g) |  | Tolerance (g) |  |  | Pass/Fail
| 1 |  | 1 |  | 2
| 5 |  | 5 |  | 2
| 10 |  | 10 |  | 2
| 50 |  | 50 |  | 2
| 100 |  | 100 |  | 2
| 500 |  | 500 |  | 2
| 501 |  | 501 |  | 2
| 1000 |  | 1000 |  | 2
| 5000 |  | 5000 |  | 5
| 5001 |  | 5001 |  | 5
| 10000 |  | 10000 |  | 10
| 10001 |  | 10001 |  | 10
| 20000 |  | 20001 |  | 20
| 30000 |  | 30000 |  | 30
| 30001 |  | 30001 |  | 30
[9 empty] | Pass/Fail: | Pass
Comments:
Initial By:

## Scale (G5-MX2)
Scale Calibration Worksheet
| Equipment ID: | MX - 2 |  |  |  |  | Date: | 2013-09-24
| Manufacturer: | Mettler |  |  |  |  | Performed by: | M. Fleck
| Model #: | SG 32000
| Serial #: | 1120303652 |  |  |  |  | Next Calibration Due: | September 2014
| UofA ID: | NA |  |  |  |  | Last Calibration: | September 2012
| Location: | MX (234) (Left)
| Calibration Item: | Verify mass readings for at least 5 points in the range of the balance
Check sensitivity (0.1%) for at least 5 points in the range of the balance
| Calibration Procedure: | AASHTO M-231
| Calibration Equipment: | Class F, Class I, and 10-kg Calibration Weights
Tolerances:
| Class |  | Readability and Sensitivity |  |  |  | Accuracy
| G1 |  |  | 0.01 g |  | 0.02 g or 0.1 percent
| G2 |  |  | 0.1 g |  | 0.2 or 0.1 percent
| G5 |  |  | 1 g |  | 2 g or 0.1 percent
| G20 |  |  | 5 g |  | 5 g or 0.1 percent
| G100 |  |  | 20 g |  | 20 g or 0.1 percent
| Weight Applied (g) |  | Scale Reading (g) |  | Tolerance (g) |  |  | Pass/Fail
| 1 |  | 1 |  | 2
| 5 |  | 5 |  | 2
| 10 |  | 10 |  | 2
| 50 |  | 50 |  | 2
| 100 |  | 100 |  | 2
| 500 |  | 500 |  | 2
| 501 |  | 501 |  | 2
| 1000 |  | 1000 |  | 2
| 5000 |  | 5000 |  | 5
| 5001 |  | 5001 |  | 5
| 10000 |  | 10000 |  | 10
| 10001 |  | 10001 |  | 10
| 20000 |  | 20001 |  | 20
| 30000 |  | 30001 |  | 30
| 30001 |  | 30002 |  | 30
[9 empty] | Pass/Fail: | Pass
Comments:
Initial By:

## Scale (G100-C1)
Scale Calibration Worksheet
| Equipment ID: | C - 1 |  |  |  |  | Date: | 2013-09-24
| Manufacturer: | AND |  |  |  |  | Performed by: | M. Fleck
| Model #: | HV60KGL
| Serial #: | EM7306611 |  |  |  |  | Next Calibration Due: | September 2014
| UofA ID: | NA |  |  |  |  | Last Calibration: | September 2012
| Location: | CL (260)
| Calibration Item: | Verify mass readings for at least 5 points in the range of the balance
Check sensitivity (0.1%) for at least 5 points in the range of the balance
| Calibration Procedure: | AASHTO M-231
| Calibration Equipment: | Class F, Class I, and 10-kg Calibration Weights
Tolerances:
| Class |  | Readability and Sensitivity |  |  |  | Accuracy
| G1 |  |  | 0.01 g |  | 0.02 g or 0.1 percent
| G2 |  |  | 0.1 g |  | 0.2 or 0.1 percent
| G5 |  |  | 1 g |  | 2 g or 0.1 percent
| G20 |  |  | 5 g |  | 5 g or 0.1 percent
| G100 |  |  | 20 g |  | 20 g or 0.1 percent
| Weight Applied (g) |  | Scale Reading (g) |  | Tolerance (g) |  |  | Pass/Fail
| 20 |  | 20 |  | 20
| 500 |  | 500 |  | 20
| 520 |  | 520 |  | 20
| 1000 |  | 1000 |  | 20
| 5000 |  | 4995 |  | 20
| 10000 |  | 10000 |  | 20
| 10020 |  | 10010 |  | 20
| 15000 |  | 14990 |  | 20
| 30000 |  | 29980 |  | 30
| 30020 |  | 30000 |  | 30
| 45000 |  | 44980 |  | 45
| 45020 |  | 45000 |  | 45
| 60000 |  | 59980 |  | 60
[9 empty] | Pass/Fail: | Pass
Comments:
Initial By:

## Scale (G100-C2)
Scale Calibration Worksheet
| Equipment ID: | C - 2 |  |  |  |  | Date: | 2013-09-24
| Manufacturer: | AND |  |  |  |  | Performed by: | M. Fleck
| Model #: | FG60KAM
| Serial #: | Q1911021 |  |  |  |  | Next Calibration Due: | September 2014
| UofA ID: | NA |  |  |  |  | Last Calibration: | September 2012
| Location: | STO (216)
| Calibration Item: | Verify mass readings for at least 5 points in the range of the balance
Check sensitivity (0.1%) for at least 5 points in the range of the balance
| Calibration Procedure: | AASHTO M-231
| Calibration Equipment: | Class F, Class I, and 10-kg Calibration Weights
Tolerances:
| Class |  | Readability and Sensitivity |  |  |  | Accuracy
| G1 |  |  | 0.01 g |  | 0.02 g or 0.1 percent
| G2 |  |  | 0.1 g |  | 0.2 or 0.1 percent
| G5 |  |  | 1 g |  | 2 g or 0.1 percent
| G20 |  |  | 5 g |  | 5 g or 0.1 percent
| G100 |  |  | 20 g |  | 20 g or 0.1 percent
| Weight Applied (g) |  | Scale Reading (g) |  | Tolerance (g) |  |  | Pass/Fail
| 20 |  | 20 |  | 20
| 500 |  | 500 |  | 20
| 520 |  | 520 |  | 20
| 1000 |  | 1000 |  | 20
| 5000 |  | 4995 |  | 20
| 10000 |  | 9990 |  | 20
| 10020 |  | 10010 |  | 20
| 15000 |  | 14985 |  | 20
| 30000 |  | 29980 |  | 30
| 30020 |  | 30000 |  | 30
| 45000 |  | 45000 |  | 45
| 45020 |  | 45020 |  | 45
| 60000 |  | 60000 |  | 60
[9 empty] | Pass/Fail: | Pass
Comments:
Initial By: