## .75" Sieve (12")
Sieve Calibration Worksheet
|  | Equipment ID: | 3/4" Sieve |  |  |  |  | Serial #: | 09239286
|  | Manufacturer: | Gilson |  |  |  |  | UofA ID: | NA
|  | Model #: | V12CF"075 |  |  |  |  | Location: | ST(218)
| Calibration Item: | Verify physical condition and critical dimensions for sieves (#4 and larger)
| Calibration Procedure: | In-House Procedure #00-5
| Calibration Equipment: | Mititoyo Digimatic Calipers
| Requirement (according to AASHTO M 92) = |  |  |  |  | 19 ± 0.60 mm
(18.4mm - 19.6mm)
|  | Date: | 9/12/13 |  |  |  |  | Date: | 9/16/14
|  | Performed by: | Abiy G. |  |  |  |  | Performed by: | Cardia R.
|  | Next Calib. Due: | September 2014 |  |  |  |  | Next Calib. Due: | September 2015
|  | Last Calibration: | September 2012 |  |  |  |  | Last Calibration: | September 2013
|  | 1 | 19.01 mm |  | 19.02 mm |  |  | 1 | 19.04 |  | 19.06
|  | 2 | 19.00 mm |  | 18.92 mm |  |  | 2 | 19.12 |  | 19.06
|  | 3 | 18.91 mm |  | 19.02 mm |  |  | 3 | 19.08 |  | 19.06
|  | 4 | 18.97 mm |  | 18.89 mm |  |  | 4 | 18.95 |  | 18.94
|  | 5 | 19.00 mm |  | 19.02 mm |  |  | 5 | 19.1 |  | 19.05
|  | 6 | 19.01 mm |  | 19.05 mm |  |  | 6 | 18.994 |  | 19.03
|  | 7 | 18.93 mm |  | 18.97 mm |  |  | 7 | 19.03 |  | 18.92
|  | 8 | 18.89 mm |  | 18.97 mm |  |  | 8 | 18.83 |  | 18.88
|  | 9 | 18.98 mm |  | 18.89 mm |  |  | 9 | 19.01 |  | 19.03
|  | 10 | 18.94 mm |  | 19.06 mm |  |  | 10 | 19.02 |  | 18.95
|  | Frame: | Excellent |  |  |  |  | Frame: | Excellent
|  | Joints: | Excellent |  |  |  |  | Joints: | Excellent
|  | Cloth: | Excellent |  |  |  |  | Cloth: | Excellent
|  | Tightness: | Excellent |  |  |  |  | Tightness: | Excellent
|  | Other: | [5 empty] | Other:
| Comments: | [5 empty] | Comments:
|  |  | Pass/Fail: | Pass |  |  |  |  | Pass/Fail: | Pass
|  |  | Initial By: | [5 empty] | Initial By:

## .5" Sieve (12")
Sieve Calibration Worksheet
|  | Equipment ID: | 1/2" Sieve |  |  |  |  | Serial #: | 09228594
|  | Manufacturer: | Gilson |  |  |  |  | UofA ID: | NA
|  | Model #: | V12CF"050 |  |  |  |  | Location: | ST(218)
| Calibration Item: | Verify physical condition and critical dimensions for sieves (#4 and larger)
| Calibration Procedure: | In-House Procedure #00-5
| Calibration Equipment: | Mititoyo Digimatic Calipers
| Requirement (according to AASHTO M 92) = |  |  |  |  | 12.5 ± 0.39 mm
(12.11mm-12.89mm)
|  | Date: | 9/12/13 |  |  |  |  | Date: | 9/16/14
|  | Performed by: | Abiy G. |  |  |  |  | Performed by: | Cardia R.
|  | Next Calib. Due: | September 2014 |  |  |  |  | Next Calib. Due: | September 2015
|  | Last Calibration: | September 2012 |  |  |  |  | Last Calibration: | September 2013
|  | 1 | 12.50 mm |  | 12.57 mm |  |  | 1 | 12.58 |  | 12.55
|  | 2 | 12.73 mm |  | 12.68 mm |  |  | 2 | 12.57 |  | 12.56
|  | 3 | 12.60 mm |  | 12.45 mm |  |  | 3 | 12.66 |  | 12.54
|  | 4 | 12.64 mm |  | 12.59 mm |  |  | 4 | 12.57 |  | 12.61
|  | 5 | 12.55 mm |  | 12.65 mm |  |  | 5 | 12.6 |  | 12.67
|  | 6 | 12.63 mm |  | 12.58 mm |  |  | 6 | 12.69 |  | 12.57
|  | 7 | 12.59 mm |  | 12.56 mm |  |  | 7 | 12.63 |  | 12.65
|  | 8 | 12.60 mm |  | 12.55 mm |  |  | 8 | 12.65 |  | 12.68
|  | 9 | 12.66 mm |  | 12.64 mm |  |  | 9 | 12.61 |  | 12.56
|  | 10 | 12.58 mm |  | 12.61 mm |  |  | 10 | 12.68 |  | 12.62
|  | Frame: | Excellent |  |  |  |  | Frame: | Excellent
|  | Joints: | Excellent |  |  |  |  | Joints: | Excellent
|  | Cloth: | Excellent |  |  |  |  | Cloth: | Excellent
|  | Tightness: | Excellent |  |  |  |  | Tightness: | Excellent
|  | Other: | [5 empty] | Other:
| Comments: | [5 empty] | Comments:
|  |  | Pass/Fail: | Pass |  |  |  |  | Pass/Fail: | Pass
|  |  | Initial By: | [5 empty] | Initial By:

## 38" Sieve (12")
Sieve Calibration Worksheet
|  | Equipment ID: | 3/8" Sieve |  |  |  |  | Serial #: | 09179282
|  | Manufacturer: | Gilson |  |  |  |  | UofA ID: | NA
|  | Model #: | V12CF"038 |  |  |  |  | Location: | ST(218)
| Calibration Item: | Verify physical condition and critical dimensions for sieves (#4 and larger)
| Calibration Procedure: | In-House Procedure #00-5
| Calibration Equipment: | Mititoyo Digimatic Calipers
| Requirement (according to AASHTO M 92) = |  |  |  |  | 9.5 ± 0.30 mm
(9.20mm-9.80mm)
|  | Date: | 9/12/13 |  |  |  |  | Date: | 9/16/14
|  | Performed by: | Abiy G. |  |  |  |  | Performed by: | Cardia R.
|  | Next Calib. Due: | September 2014 |  |  |  |  | Next Calib. Due: | September 2015
|  | Last Calibration: | September 2012 |  |  |  |  | Last Calibration: | September 2013
|  | 1 | 9.50 mm |  | 9.44 mm |  |  | 1 | 9.4 |  | 9.38
|  | 2 | 9.54 mm |  | 9.44 mm |  |  | 2 | 9.58 |  | 9.55
|  | 3 | 9.47 mm |  | 9.56 mm |  |  | 3 | 9.51 |  | 9.53
|  | 4 | 9.53 mm |  | 9.55 mm |  |  | 4 | 9.57 |  | 9.45
|  | 5 | 9.44 mm |  | 9.48 mm |  |  | 5 | 9.47 |  | 9.53
|  | 6 | 9.43 mm |  | 9.49 mm |  |  | 6 | 9.43 |  | 9.51
|  | 7 | 9.52 mm |  | 9.42 mm |  |  | 7 | 9.46 |  | 9.56
|  | 8 | 9.55 mm |  | 9.49 mm |  |  | 8 | 9.48 |  | 9.53
|  | 9 | 9.52 mm |  | 9.50 mm |  |  | 9 | 9.4 |  | 9.47
|  | 10 | 9.55 mm |  | 9.42 mm |  |  | 10 | 9.47 |  | 9.58
|  | Frame: | Excellent |  |  |  |  | Frame: | Excellent
|  | Joints: | Excellent |  |  |  |  | Joints: | Excellent
|  | Cloth: | Excellent |  |  |  |  | Cloth: | Excellent
|  | Tightness: | Excellent |  |  |  |  | Tightness: | Excellent
|  | Other: | [5 empty] | Other:
| Comments: | [5 empty] | Comments:
|  |  | Pass/Fail: | Pass |  |  |  |  | Pass/Fail: | Pass
|  |  | Initial By: | [5 empty] | Initial By:

## #4 Sieve (12")
Sieve Calibration Worksheet
|  | Equipment ID: | #4 Sieve |  |  |  |  | Serial #: | 1186910
|  | Manufacturer: | Gilson |  |  |  |  | UofA ID: | NA
|  | Model #: | V12CI#004 |  |  |  |  | Location: | ST(218)
| Calibration Item: | Verify physical condition and critical dimensions for sieves (#4 and larger)
| Calibration Procedure: | In-House Procedure #00-5
| Calibration Equipment: | Mititoyo Digimatic Calipers
| Requirement (according to AASHTO M 92) = |  |  |  |  | 4.75 ± 0.15 mm
(4.60mm - 4.90mm)
|  | Date: | 9/12/13 |  |  |  |  | Date: | 9/16/14
|  | Performed by: | Abiy G. |  |  |  |  | Performed by: | Cardia R.
|  | Next Calib. Due: | September 2014 |  |  |  |  | Next Calib. Due: | September 2015
|  | Last Calibration: | September 2012 |  |  |  |  | Last Calibration: | September 2013
|  | 1 | 4.71 mm |  | 4.68 mm |  |  | 1 | 4.73 |  | 4.8
|  | 2 | 4.86 mm |  | 4.87 mm |  |  | 2 | 4.72 |  | 4.66
|  | 3 | 4.72 mm |  | 4.73 mm |  |  | 3 | 4.77 |  | 4.79
|  | 4 | 4.77 mm |  | 4.76 mm |  |  | 4 | 4.64 |  | 4.75
|  | 5 | 4.81 mm |  | 4.80 mm |  |  | 5 | 4.78 |  | 4.79
|  | 6 | 4.88 mm |  | 4.86 mm |  |  | 6 | 4.77 |  | 4.79
|  | 7 | 4.76 mm |  | 4.94 mm |  |  | 7 | 4.71 |  | 4.75
|  | 8 | 4.81 mm |  | 4.80 mm |  |  | 8 | 4.76 |  | 4.8
|  | 9 | 4.88 mm |  | 4.77 mm |  |  | 9 | 4.78 |  | 4.79
|  | 10 | 4.81 mm |  | 4.80 mm |  |  | 10 | 4.81 |  | 4.83
|  | Frame: | Excellent |  |  |  |  | Frame: | Excellent
|  | Joints: | Excellent |  |  |  |  | Joints: | Excellent
|  | Cloth: | Excellent |  |  |  |  | Cloth: | Excellent
|  | Tightness: | Excellent |  |  |  |  | Tightness: | Excellent
|  | Other: | [5 empty] | Other:
| Comments: | [5 empty] | Comments:
|  |  | Pass/Fail: | Pass |  |  |  |  | Pass/Fail: | Pass
|  |  | Initial By: | [5 empty] | Initial By:

## # 8 Sieve (12")
Sieve Calibration Worksheet
| Equipment ID: | # 8 Sieve |  |  | Serial #: | 09269019
| Manufacturer: | Gilson |  |  | UofA ID: | NA
| Model #: | V12CI#008 |  |  | Location: | ST (218)
[5 empty] | Calibration Interval: | 12 months
| Calibration Item: | Verify physical condition of sieve and critical dimensions for
sieves larger than #4 (4.75 mm)
| Calibration Procedure: | In-House Procedure #00-5
| Calibration Equipment: | Mititoyo Digimatic Calipers
Condition
Date | Frame | Joints | Cloth | Tightness | Pass / Fail | Checked By | Initial by
2011-09-20 | New | New | New | New | Pass | Haley M. | sgw
2012-09-04 | Excellent | Excellent | Excellent | Excellent | Pass | Thomas Moss | sgw
2013-09-12 | Excellent | Excellent | Excellent | Excellent | Pass | Abiy G. | sgw
2014-09-17 | Excellent | Excellent | Excellent | Excellent | Pass | M. Fleck

## # 16 Sieve (12")
Sieve Calibration Worksheet
| Equipment ID: | # 16 Sieve |  |  | Serial #: | 09289020
| Manufacturer: | Gilson |  |  | UofA ID: | NA
| Model #: | V12CI# 016 |  |  | Location: | ST (218)
[5 empty] | Calibration Interval: | 12 months
| Calibration Item: | Verify physical condition of sieve and critical dimensions for
sieves larger than #4 (4.75 mm)
| Calibration Procedure: | In-House Procedure #00-5
| Calibration Equipment: | Mititoyo Digimatic Calipers
Condition
Date | Frame | Joints | Cloth | Tightness | Pass / Fail | Checked By | Initial by
2011-09-20 | New | New | New | New | Pass | Haley M. | sgw
2012-09-04 | Excellent | Excellent | Excellent | Excellent | Pass | Thomas Moss | sgw
2013-09-12 | Excellent | Excellent | Excellent | Excellent | Pass | Abiy G. | sgw
2014-09-17 | Excellent | Excellent | Excellent | Excellent | Pass | M. Fleck

## # 30 Sieve (12")
Sieve Calibration Worksheet
| Equipment ID: | # 30 Sieve |  |  | Serial #: | 09289021
| Manufacturer: | Gilson |  |  | UofA ID: | NA
| Model #: | V12CI#030 |  |  | Location: | ST (218)
[5 empty] | Calibration Interval: | 12 months
| Calibration Item: | Verify physical condition of sieve and critical dimensions for
sieves larger than #4 (4.75 mm)
| Calibration Procedure: | In-House Procedure #00-5
| Calibration Equipment: | Mititoyo Digimatic Calipers
Condition
Date | Frame | Joints | Cloth | Tightness | Pass / Fail | Checked By | Initial by
2011-09-20 | New | New | New | New | Pass | Haley M. | sgw
2012-09-04 | Excellent | Excellent | Excellent | Excellent | Pass | Thomas Moss | sgw
2013-09-12 | Excellent | Excellent | Excellent | Excellent | Pass | Abiy G. | sgw
2014-09-17 | Excellent | Excellent | Excellent | Excellent | Pass | M. Fleck

## # 50 Sieve (12")
Sieve Calibration Worksheet
| Equipment ID: | # 50 Sieve |  |  | Serial #: | 09289022
| Manufacturer: | Gilson |  |  | UofA ID: | NA
| Model #: | V12CI# 050 |  |  | Location: | ST (218)
[5 empty] | Calibration Interval: | 12 months
| Calibration Item: | Verify physical condition of sieve and critical dimensions for
sieves larger than #4 (4.75 mm)
| Calibration Procedure: | In-House Procedure #00-5
| Calibration Equipment: | Mititoyo Digimatic Calipers
Condition
Date | Frame | Joints | Cloth | Tightness | Pass / Fail | Checked By | Initial by
2011-09-20 | New | New | New | New | Pass | Haley M. | sgw
2012-09-04 | Excellent | Excellent | Excellent | Excellent | Pass | Thomas Moss | sgw
2013-09-12 | Excellent | Excellent | Excellent | Excellent | Pass | Abiy G. | sgw
2014-09-17 | Excellent | Excellent | Excellent | Excellent | Pass | M. Fleck

## # 100 Sieve (12")
Sieve Calibration Worksheet
| Equipment ID: | # 100 Sieve |  |  | Serial #: | 09289435
| Manufacturer: | Gilson |  |  | UofA ID: | NA
| Model #: | V12CI#100 |  |  | Location: | ST (218)
[5 empty] | Calibration Interval: | 12 months
| Calibration Item: | Verify physical condition of sieve and critical dimensions for
sieves larger than #4 (4.75 mm)
| Calibration Procedure: | In-House Procedure #00-5
| Calibration Equipment: | Mititoyo Digimatic Calipers
Condition
Date | Frame | Joints | Cloth | Tightness | Pass / Fail | Checked By | Initial by
2011-09-20 | New | New | New | New | Pass | Haley M. | sgw
2012-09-04 | Excellent | Excellent | Excellent | Excellent | Pass | Thomas Moss | sgw
2013-09-12 | Excellent | Excellent | Excellent | Excellent | Pass | Abiy G. | sgw
2014-09-17 | Excellent | Excellent | Excellent | Excellent | Pass | M. Fleck

## # 200 Sieve (12")
Sieve Calibration Worksheet
| Equipment ID: | # 200 Sieve |  |  | Serial #: | 09289024
| Manufacturer: | Gilson |  |  | UofA ID: | NA
| Model #: | V12CI#200 |  |  | Location: | ST (218)
[5 empty] | Calibration Interval: | 12 months
| Calibration Item: | Verify physical condition of sieve and critical dimensions for
sieves larger than #4 (4.75 mm)
| Calibration Procedure: | In-House Procedure #00-5
| Calibration Equipment: | Mititoyo Digimatic Calipers
Condition
Date | Frame | Joints | Cloth | Tightness | Pass / Fail | Checked By | Initial by
2011-09-20 | New | New | New | New | Pass | Haley M. | sgw
2012-09-04 | Excellent | Excellent | Excellent | Excellent | Pass | Thomas Moss | sgw
2013-09-12 | Excellent | Excellent | Excellent | Excellent | Pass | Abiy G. | sgw
2014-09-17 | Excellent | Excellent | Excellent | Excellent | Pass | M. Fleck

## Fine Sieve Calibration (12")
Sieve Calibration Worksheet
| Equipment ID: | Fine Sieves |  |  | Serial #:
| Manufacturer: | Gilson |  |  | UofA ID: | NA
| Model #: | V12CI - |  |  | Location: | ST (218)
[5 empty] | Calibration Interval: | 12 months
| Calibration Item: | Verify physical condition of sieve and critical dimensions for
sieves larger than #4 (4.75 mm)
| Calibration Procedure: | In-House Procedure #00-5
| Calibration Equipment: | Mititoyo Digimatic Calipers
March 2010
Condition
Sieve | Frame | Joints | Cloth | Tightness | Comments |  | Initial by
# 8
# 16
# 30
# 50
# 100
# 200
The fine sieves are logged on individual sheets in the equipment calibration manual.