## Gauge Calibration
Vacuum Gauge Calibration Worksheet
| Equipment ID: | Digital Manometer |  |  |  | Serial #: | 344
| Manufacturer: | Karol-Warner |  |  |  | UofA ID: | N/A
| Model #: | 9210 |  |  |  | Location: | CL 220
|  | Calibration Item: | Verify pressure readings
|  | Calibration Procedure: | In-House Procedure #00-2
|  | Calibration Equipment: | Vacuum Verification Gauge
|  | Calibration Interval: | 12 months
Date | Vacuum Verification Reading (Torr) = (mm) |  | Manometer Reading (mm Hg) | Allowable Manometer Reading (mm Hg) |  |  | Pass Fail | Tested By
| Range of 25.00-30.00 |  |  | V V R ± 5 mm
2010-02-10 | 28 |  | 25.8 | 23 | to | 33 | Pass | MF
2011-02-08 | 29 |  | 27.5 | 24 | to | 34 | Pass | MF
2012-02-13 | 27.5 |  | 24.5 | 22.5 | to | 32.5 | Pass | MF
2013-02-04 | 27.5 |  | 25.1 | 22.5 | to | 32.5 | Pass | MF
2014-02-07 | 27 |  | 22 | 22 | to | 32 | Pass | MF
to
to
to
to
to
to
to
to
to
to
to
to
to
to
to
to
Notes : | 2014-02-07 | Calibration factor of (-5.0) tag placed on equipment in lab.
Initial :

## Pump Maintenance
Vacuum Pump Maintenance Schedule Worksheet
| Equipment ID: | High Vacuum Pump |  | Serial #: | None
| Manufacturer: | Fischer |  | UofA ID: | None
| Model #: | LAV - 3 |  | Location: | CL (220)
|  | Maintenance Item: | Verify pressure readings
|  | Maintenance Procedure: | see manufacturer's instructions
|  | Maintenance Equipment: | fresh vacuum pump oil, proper disposal equipment
|  | Maintenance Interval: | 12 months
Date: | Changed by: |  | Comments:
2009-04-15 | Mary Fleck |  | Changed Oil
2010-02-05 | David Peachee |  | Changed Oil
2011-02-07 | David Peachee |  | Drained water from oil.
2012-02-13 | David Peachee |  | Changed Oil
2013-02-05 | David Peachee |  | Drain water from oil, dried dri-rite crystals
2013-06-25 | David Peachee |  | Changed Oil
2014-01-14 | Rosalie Conley |  | Drained water from oil.
2014-07-25 | David Peachee |  | Changed Oil
Initial :