## Coarse Agg. Splitter
Mechanical Splitter Calibration Worksheet
|  | Equipment ID: | Coarse Agg. Splitter |  | Date: | 3/1/16
|  | Manufacturer: | Soiltest |  | Performed by: | Mary Fleck
|  | Model #: | CL - 262
|  | Serial #: | NA |  | Next Calibration Due: | March 2017
|  | UofA ID: | NA |  | Last Calibration: | March 2015
|  | Location: | AGG (233)
| Calibration Item: | Verify critical dimensions of the mechanical splitter
| Calibration Procedure: | In-House Procedure #00-11
| Calibration Equipment: | Ruler, readable to 1 mm.
|  | Number of Chutes: | 16 | Tolerance = even number
at least 8 for coarse, 12 for fine
|  | Chute Opening: | 38 mm | Tolerance = approx. 50% larger than
|  |  | 38 mm |  | largest aggregate size
38 mm
38 mm
38 mm
38 mm
38 mm
38 mm
38 mm
38 mm
38 mm
38 mm
38 mm
38 mm
38 mm
38 mm
[6 empty] | Pass/Fail: | Pass
Comments:
Initial By:

## Fine Agg. Splitter
Mechanical Splitter Calibration Worksheet
|  | Equipment ID: | Fine Agg. Splitter |  | Date: | 3/1/16
|  | Manufacturer: | Soiltest |  | Performed by: | Mary Fleck
|  | Model #: | CL - 283
|  | Serial #: | NA |  | Next Calibration Due: | March 2017
|  | UofA ID: | NA |  | Last Calibration: | March 2015
|  | Location: | AGG (233)
| Calibration Item: | Verify critical dimensions of the mechanical splitter
| Calibration Procedure: | In-House Procedure #00-11
| Calibration Equipment: | Ruler, readable to 1 mm.
|  | Number of Chutes: | 14 | Tolerance = even number
at least 8 for coarse, 12 for fine
|  | Chute Opening: | 19 mm | Tolerance = approx. 50% larger than
|  |  | 19 mm |  | largest aggregate size
18 mm
19 mm
19 mm
19 mm
19 mm
19 mm
19 mm
19 mm
19 mm
19 mm
19 mm
18 mm
[6 empty] | Pass/Fail: | Pass
Comments:
Initial By:

## Sieve Shaker - MaryAnn 1
Sieve Shaker Calibration Worksheet
|  | Equipment ID: | Mary Ann 1 |  | Date: | 3/3/16
|  | Manufacturer: | Rainhart |  | Performed by: | M. Fleck
|  | Model #: | 637
|  | Serial #: | NA |  | Next Calibration Due: | March 2017
|  | UofA ID: | NA |  | Last Calibration: | March 2015
|  | Location: | SAW (235)
| Calibration Item: | Verify sieving thoroughness of sieve shakers
| Calibration Procedure: | In-House Procedure #00-10
| Calibration Equipment: | Sieves, Scale CL1, pan
|  | Original Mass of Sample: | 1843.9
After Hand Sieving:
|  |  |  | Mass Passing | Percent Passing
|  | 3/4" | 0 |  | 0
|  | 1/2" | 2.8 |  | 0.15185
|  | #4 | 0.3 |  | 0.01627
|  | #8 | 0.9 |  | 0.04881
|  | #16 | 0.3 |  | 0.01627
|  | #50 | 0.2 |  | 0.010847
|  | #100 | 0.4 |  | 0.021693
| Tolerance: | No sieve shall have greater than 0.5% passing by total mass of sample.
[6 empty] | Pass/Fail: | Pass
| Comments: | Shake Time = 7 minutes
Initial By:
Maintenance (see mfg. instructions):
√ | Check for cleanliness?
√ | Paddles do not show excessive wear?
√ | Greased if needed?
| Comments: | No maintenance required.
Initial By:

## Sieve Shaker - MaryAnn 2
Sieve Shaker Calibration Worksheet
|  | Equipment ID: | Mary Ann 2 |  | Date: | 3/3/16
|  | Manufacturer: | Rainhart |  | Performed by: | M. Fleck
|  | Model #: | 637
|  | Serial #: | NA |  | Next Calibration Due: | March 2017
|  | UofA ID: | NA |  | Last Calibration: | March 2015
|  | Location: | SAW (235)
| Calibration Item: | Verify sieving thoroughness of sieve shakers
| Calibration Procedure: | In-House Procedure #00-10
| Calibration Equipment: | Sieves, Scale CL1, pan
|  | Original Mass of Sample: | 1844.6
After Hand Sieving:
|  |  |  | Mass Passing | Percent Passing
|  | 3/4" | 0 |  | 0
|  | 1/2" | 0 |  | 0
|  | #4 | 0.7 |  | 0.037949
|  | #8 | 0.8 |  | 0.04337
|  | #16 | 0.6 |  | 0.032527
|  | #50 | 0.2 |  | 0.010842
|  | #100 | 0.5 |  | 0.027106
| Tolerance: | No sieve shall have greater than 0.5% passing by total mass of sample.
[6 empty] | Pass/Fail: | Pass
| Comments: | Shake Time = 7 minutes
Initial By:
Maintenance (see mfg. instructions):
√ | Check for cleanliness?
√ | Paddles do not show excessive wear?
√ | Greased if needed?
| Comments: | No maintenance required
Initial By:

## Cone and Tamper Calibration
Cone and Tamper Calibration Worksheet
|  | Equipment ID: | Cone & Tamper |  |  | Date: | 3/1/16
|  | Manufacturer: | ELE |  |  | Performed by: | M. Fleck
|  | Model #: | G - 325
|  | Serial #: | NA |  |  | Next Calibration Due: | March 2017
|  | UofA ID: | NA |  |  | Last Calibration: | March 2015
|  | Location: | ST (218)
| Calibration Item: | Verify critical dimensions of the cone and tamper
| Calibration Procedure: | In-House Procedure #00-8
| Calibration Equipment: | Mititoyo Digimatic Calipers, Scale CL1, and ruler read. to 1 mm.
|  | Top Inside Diameter of Cone: | 41.26 mm | 41.34 mm | Tolerance = 40 +/- 3 mm
|  | Bottom Inside Diameter of Cone: | 90.15 mm | 90.23 mm | Tolerance = 90 +/- 3 mm
|  | Height of Cone: | 76.31 mm |  | Tolerance = 75 +/- 3 mm
|  | Thickness of Cone: | 1.18 mm | 1.29 mm | Tolerance = at least 0.8 mm
|  |  | 1.41 mm | 1.24 mm
|  | Diameter of tamper face: | 25.31 mm | 25.31 mm | Tolerance = 25 +/- 3 mm
|  | Mass of tamper: | 345.3 g |  | Tolerance = 340 +/- 15 g
[7 empty] | Pass/Fail: | Pass
Comments:
Initial By:

## Fine Agg. Angularity
Fine Aggregate Angularity Calibration Worksheet
|  | Equipment ID: | FAA # 2 |  | Date: | 3/1/16
|  | Manufacturer: | Gilson |  | Performed by: | M. Fleck
|  | Model #: | H - 1680
|  | Serial #: | NA |  | Next Calibration Due: | March 2017
|  | UofA ID: | NA |  | Last Calibration: | March 2015
|  | Location: | ST (218)
| Calibration Item: | Verify critical dimensions of the fine aggregate angularity device
| Calibration Procedure: | In-House Procedure #00-13
| Calibration Equipment: | Mititoyo Digimatic Calipers, balance read. to 0.1 g, and ruler read. to 1 mm.
|  | Inside Diameter of Cylinder: | 39.33 mm | Tolerance = approx. 39.02 mm
|  | Inside Height of Cylinder: | 82.77 mm | Tolerance = approx. 86 mm
|  | Thickness of Cylinder Bottom: | 6.30 mm | Tolerance = at least 6 mm
|  | Inside Diameter of Funnel: | 12.33 mm | Tolerance = 12.7 +/- 0.6 mm
|  | Height of Funnel: | 50.60 mm | Tolerance = at least 38 mm
|  | Distance from Funnel to Cylinder: | 115 mm | Tolerance = 115 +/- 2 mm
|  | Width of Glass Plate: | 59.80 mm | Tolerance = approx. 60 mm
|  | Length of Glass Plate: | 60.11 mm | Tolerance = approx. 60 mm
|  | Thickness of Glass Plate: | 4.75 mm | Tolerance = at least 4 mm
[6 empty] | Pass/Fail: | Pass
Comments:
Initial By: