## Maint. G1
Calibration Worksheet: |  | Pine G1 Gyratory Maintenance Log
Equipment ID: | Pine G1 Compactor |  |  | Maintenance Item: | Lubricate and adjustments
Manufacturer: | Pine Instruments |  |  | Maintenance Procedure: | See Pine G1 Inst. Manual
Model : | AFG 1A |  |  | Maintenance Equipment: | Grease, oil, cleaner
Serial # : | 1429 |  |  | Maintenance Interval: | 12 months
Location : | HM (225)
Date | Machine Hrs | Maintenance Performed |  |  |  | By
2012-04-13 | 54.3 | Fully serviced by Pine Instrument technician |  |  |  | M. D.
2013-02-05 | 54.6 | None Required |  |  |  | D.P.
2013-08-21 | 56.9 | None Required |  |  |  | M. Fleck
2014-01-27 | 61.4 | None Required |  |  |  | M. Fleck
2014-07-25 | 69.8 | DC power supply replaced |  |  |  | D. Peachee
2015-02-04 | 72.2 | All 25 hr checks - No maintenance required |  |  |  | D. Peachee
2015-07-01 | 72.2 | None Required |  |  |  | M. Fleck
2016-01-06 | 72.2 | None Required |  |  |  | M. Fleck
Initial :
Maintenance Items for G1 Compactory
Inspect Ram Key |  | 25 hr
Timing Belt Tension |  | 25 hr
Inspect Mold Clamps |  | 25 hr
Inspect Ram Foot for Wear |  | 25 hr
Lubricate Ball Screw Bearings |  | 100 hr | NLGI grade 2 Lithium Soap Grease
Lubricate Ball Screw |  | 100 hr | NLGI grade 2 Lithium Soap Grease
Lubricate Actuator Bearings |  | 100 hr | NLGI grade 2 Lithium Soap Grease
Lubricate Mold Clamp Pivot |  | 100 hr | Anti-Sieze Lubricant
Lubricate Top Clamps |  | 100 hr | Anti-Sieze Lubricant

## Maint. G2
Calibration Worksheet: |  | Pine G2 Gyratory Maintenance Log
Equipment ID: | Pine G2 Compactor |  |  | Maintenance Item: | Lubricate and adjustments
Manufacturer: | Pine Instruments |  |  | Maintenance Procedure: | See Pine G2 Inst. Manual
Model : | AFG 2AS |  |  | Maintenance Equipment: | Grease, oil, cleaner
Serial # : | 8001 |  |  | Maintenance Interval: | 12 months
Location : | HM (225)
Date | Machine Hrs | Maintenance Performed |  |  |  | By
2012-04-13 | 55.24 | Fully serviced by Pine Instrument technician |  |  |  | M. D.
2013-02-05 | 57.69 | None Required |  |  |  | D. P.
2013-08-21 | 70.94 | All Maintenance Items Listed on Next Page |  |  |  | D. P.
2014-01-27 | 84.46 | None Required |  |  |  | M. F.
2014-07-07 | 93.49 | None Required |  |  |  | M. F.
2015-02-04 | 115.8 | All 25 hr, 100 hr checks & 100 hr lubrications - OK |  |  |  | D. P.
2015-07-01 | 117 | None Required |  |  |  | M. F.
2016-01-06 | 119.8 | Right clamp adjusted. Cleaned. None other required. |  |  |  | M.F.
Initial :
Maintenance Items for G2 Compactory
Timing Belt Tension |  | 25 hr
Inspect Mold Clamps |  | 25 hr
Check Mold Base Clamps |  | 25 hr
Inspect Ram Key |  | 100 hr
Inspect Ram Foot for Wear |  | 100 hr
Lubricate Ball Screw Bearings |  | 100 hr | NLGI grade 2 Lithium Soap Grease
Lubricate Actuator Bearings |  | 100 hr | NLGI grade 2 Lithium Soap Grease
Lubricate Mold Clamp Pivot |  | 100 hr | Anti-Sieze Lubricant
Lubricate Top Clamps |  | 100 hr | Anti-Sieze Lubricant
Lubricate Ball Screw |  | 500 hr | NLGI grade 2 Lithium Soap Grease

## Pine G1 Compactor
| Calibration Worksheet: |  | Pine G1 Compactor Calibration Verification
| Equipment ID: | Pine G1 Compactor |  |  | Date: | 2016-01-06
| Manufacturer: | Pine Instruments |  |  | Performed by: | M. Fleck
| Model #: | AFG 1A
| Serial #: | 1429 |  |  | Next Calibration Due: | Jan. 2017
| UofA ID: | NA |  |  | Last Calibration: | Jan. 2015
| Location: | HM (225)
|  | Calibration Items: | Speed, Pressure, Height, Angle
|  | Calibration Procedure: | Follow Manufacturer's Instructions (summarized below)
|  | Calibration Equipment: | Calibration kit supplied by manufacturer
SPEED |  | Select CALIBRATION and press ENTER [ ]
| Select VERIFY, press ENTER. Install mold top. |  |  |  |  | Target Time: | 20 seconds
| Select VERIFY SPEED and press ENTER. Press START. |  |  |  |  | Actual Time: | 20.13
| Allow machine to rotate, then count and time 10 gyrations. |  |  |  |  | Tolerance: | +/- 0.33 sec
| Record stopwatch reading for 10 gyrations. Verify unloaded angle 1.21 ± 0.01° |  |  |  |  | Pass / Fail: | Pass
RAM PRESSURE |  |  |  | HEIGHT
| Select CALIBRATION and press ENTER. |  |  |  | Center 6" gage blocks under ram as directed.
| Select VERIFY and press ENTER. |  |  |  | Press START. Verify height measurements.
| Select VERIFY FORCE/HEIGHT and press ENTER. |  |  |  | Press START. Press (+) or (-) to accept. Remove
| Install mold top if not done. |  |  |  | gage blocks and spacer. Lube with anti-sieze.
Press START for insertion of ram foot spacer.
| Clean area & insert ram foot calibration spacer. |  |  |  | Target (mm) | Actual (mm)
| Press START. |  |  |  | 152.4 | 152.4
| Place proving ring into chamber. |  |  |  | Tolerance: | ± 0.05 mm
Center proving ring and 1/8" block in compactor.
| Press START. Compactor flexes proving ring. |  |  |  | Pass / Fail: | Pass
Zero dial. Press START. Follow screen prompts.
INTERNAL ANGLE
| Load (N) | Target Dial Rdg. | Actual Dial Rdg.
| 5000 | 113.8 | 112.8 |  | Place tube over RAM. Set to mode 3.
| 10500 | 238.1 | 237.8 |  | Rotate 1 revolution. Read angle.
|  | Tolerance : | ± 1 %
[5 empty] | Calib. Tube | 1) | 1.23
| Pass / Fail: | Pass |  |  | 1.238 | 2) | 1.23
[5 empty] | Pass / Fail | 3) | 1.23
[6 empty] | Avg | 1.23
| Press START to park ram. | [5 empty] | (± 0.01°)
| Press (+) or (-) for acceptance. |  |  |  | Insert RAM into mold.
| Remove proving ring and press ENTER to |  |  |  | Set gyratory # of gyrations to 15.
| continue with height verification. |  |  |  | Turn RAM off, then on. Lower and start test.
Read angle. Take 2 top and 2 bottom readings.
[5 empty] | Top | Bottom | Effective Interior
[5 empty] | 1.16 | 1.15 | Angle
[5 empty] | 1.15 | 1.16
[5 empty] | 1.155 | 1.155 | 1.155
[5 empty] | Average | Tolerance | 1.16± 0.02°
| Machine Hours : | 72.2 |  |  | Pass / Fail: | Pass
| Maint. Notes : | Gyratory operating normally during all verification checks. No maintenance required. |  |  |  | Initial By :
No routine maintenance required.

## Pine G2 Compactor
| Calibration Worksheet: |  | Pine G2 Compactor Calibration Verification
| Equipment ID: | Pine G2 Compactor |  |  | Date: | 2015-01-19
| Manufacturer: | Pine Instruments |  |  | Performed by: | M. Fleck
| Model #: | AFG2AS
| Serial #: | 8001 |  |  | Next Calibration Due: | Jan. 2017
| UofA ID: | 263529 |  |  | Last Calibration: | Jan. 2015
| Location: | HM (225)
|  | Calibration Items: | Speed, Pressure, Height, Angle
|  | Calibration Procedure: | Follow Manufacturer's Instructions (summarized below)
|  | Calibration Equipment: | Calibration kit supplied by manufacturer
SPEED |  | Select SETUP/CALIBRATE and press ENTER [ ]
| Select CALIBRATE, press ENTER. Select VERIFY and press ENTER. |  |  |  |  | Target Time: | 20 seconds
| Select VERIFY SPEED and press ENTER. Press START. |  |  |  |  | Actual Time: | 19.94
| Allow machine to rotate, then count and time 10 gyrations. |  |  |  |  | Tolerance: | +/- 0.33 sec
| Record stopwatch reading for 10 gyrations. Follow screen prompts. |  |  |  |  | Pass / Fail: | Pass
RAM PRESSURE |  |  |  | HEIGHT
| Select VERIFY FORCE and press ENTER. |  |  |  | Select VERIFY HEIGHTS and press ENTER.
| Center proving ring and 1/8" block in compactor. |  |  |  | Center 6" gage blocks under ram as directed.
| Press START. Compactor flexes proving ring. |  |  |  | Select 150mm and press ENTER. Press START.
| Zero dial. Press START. Follow screen prompts. |  |  |  | Record height. Follow screen prompts.
| Load (N) | Target Dial Rdg. | Actual Dial Rdg. |  | Target (mm) | Actual (mm)
| 5000 | 113.8 | 114.8 |  | 152.4 | 152.4
| 10500 | 238.1 | 238.9 |  | Tolerance: | ± 0.05 mm
|  | Tolerance : | ± 1 %
| Pass / Fail: | Pass |  |  | Pass / Fail: | Pass
Heights recalibrated and verified.
GYRATORY SHEAR |  |  |  | INTERNAL ANGLE
Clean and lubricate ram foot with "Safety Ease".
| Select VERIFY SHEAR and press ENTER. |  |  |  | Place tube over RAM. Set to mode 3.
| LOAD 32mm rings into RAM and lubricate edge |  |  |  | Rotate 1 revolution. Read angle.
vaseline. Follow screen prompts.
| Press START. Record reading from compactor. |  |  |  | Calib. Tube | 1) | 1.23
| Remove rings, and repeat loading. |  |  |  | 1.238 | 2) | 1.23
[5 empty] | Pass / Fail | 3) | 1.23
| Eccentricity : | (large rings) | 32 mm |  |  | Avg | 1.23
| Tilting Moment : | 678.5 N-m | 703.8 |  |  |  | (± 0.01°)
|  |  | (610 - 746) |  | Insert RAM into mold.
| Pass / Fail: | Pass |  |  | Set gyratory # of gyrations to 15.
Turn device off, then on. Lower and start test.
| Eccentricity : | (no rings) | 22 mm |  | Read angle. Take 2 top and 2 bottom readings.
| Tilting Moment : | 466.5 N-m | 476.8
|  |  | (419 - 513) |  | Top | Bottom | Effective Interior
| Pass / Fail: | Pass |  |  | 1.16 | 1.14 | Angle
[5 empty] | 1.16 | 1.14
| SHEAR |  |  |  | 1.16 | 1.14 | 1.15
| (requires both to pass) |  |  |  | Average | Tolerance | 1.16 ± 0.02°
| Pass / Fail: | Pass |  |  | Pass / Fail: | PASS
| Machine Hours : | 119.8
| Maint. Notes : | Gyratory operating normally during all verification checks. No maintenance required. |  |  |  | Initial By :
Right clamp adjusted and internal angle checked.

## Pine G2 - Full
| Calibration Worksheet: |  |  | Pine Gyratory Compactor
| Equipment ID: | Pine G2 Compactor |  |  |  | Date:
| Manufacturer: | Pine Instruments |  |  |  | Performed by: | Mary Fleck
| Model #: | AFG2AS
| Serial #: | 8001 |  |  |  | Next Calibration Due: | July 2011
| UofA ID: | 263529 |  |  |  | Last Calibration: | New
| Location: | HM (225)
|  | Calibration Items: |  | Speed, Pressure, Height, Angle
|  | Calibration Procedure: |  | Follow Manufacturer's Instructions (summarized below)
|  | Calibration Equipment: |  | Calibration kit supplied by manufacturer
SPEED
| Press ENTER and SELECT at the same time. | [5 empty] | Target Time: | 20 seconds
| Scroll to enter code (125) | [5 empty] | Actual Time:
| Select VERIFY from menu and press ENTER. Press START. | [5 empty] | Tolerance: | +/- 0.3 sec
| Record stopwatch reading for 10 gyrations | [5 empty] | Pass / Fail: | Pass
RAM PRESSURE |  |  |  |  | HEIGHT
| Place proving ring and 1/8" block in compactor. |  |  |  |  | Select VERIFY from ram position calibration menu.
| Select VERIFY from ram force calibration menu. |  |  |  |  | Place blocks under ram as directed.
| Follow screen prompts. |  |  |  |  | Follow screen prompts.
| Load (N) | Target Dial Rdg. |  | Actual Dial Rdg. |  | Target (mm) | Actual (mm)
| 1500 | 33.9 |  |  |  | 254
| 3500 | 79.9 |  |  |  | 228.6
| 5500 | 124.6 |  |  |  | 203.2
| 7500 | 169.9 |  |  |  | 177.8
| 9500 | 215.1 |  |  |  | 152.4
| 11500 | 261.8 |  |  |  | 127
| 13500 | 308.2 |  |  |  | 101.6
| 15500 | 354.1 |  |  |  | 76.2
| 17500 | 399.7
[6 empty] | Tolerance: | ± 0.05 mm
| Tolerance: | ±1% or ±3% |  |  |  | Pass / Fail: | Pass
| Pass / Fail: | Pass
INTERNAL ANGLE |  |  |  |  | SHEAR
1)
| Calib. Tube | 2) |  |  |  | Eccentricity : |  | 32 mm
| 1.234° | 3) |  |  |  | Tilting Moment :
|  | Avg | #DIV/0! |  |  |  |  | (678 N-m)
(± 0.01°)
| Top # 1 | [6 empty] | Pass / Fail
| Top # 2 |  |  | To calculate effective angle, average the top and bottom average angle measurements.
| Top Average |  |  |  |  | Eccentricity : |  | 22 mm
Tilting Moment :
| Bottom # 1 | [6 empty] | (466 N-m)
Bottom # 2
| Bottom Average | [6 empty] | Pass / Fail
| EFF. ANGLE |  |  | PASS |  | SHEAR |  | PASS
| Tolerance : | 1.16 ± 0.02° |  | Pass / Fail : |  | (requires both to pass) |  | Pass / Fail :
Maintenance performed according to mfg. instructions
| Initial | Date

## Pine Gyratory - Full
| Calibration Worksheet: |  | Pine Gyratory Compactor - Full Verification
| Equipment ID: | Pine Compactor |  |  | Date:
| Manufacturer: | Pine Instruments |  |  | Performed by:
| Model #: | AFGC125X
| Serial #: | 15 |  |  | Next Calibration Due:
| UofA ID: | 229874 |  |  | Last Calibration:
| Location: | HM (225)
|  | Calibration Items: | Speed, Pressure, Height, Angle
|  | Calibration Procedure: | Follow Manufacturer's Instructions (summarized below)
|  | Calibration Equipment: | Calibration kit supplied by manufacturer
SPEED
| Press ENTER and SELECT at the same time. |  |  |  |  | Target Time: | 20 seconds
| Scroll to enter code (125) |  |  |  |  | Actual Time:
| Select VERIFY from menu and press ENTER. Press START. |  |  |  |  | Tolerance: | +/- 0.3 sec
| Record stopwatch reading for 10 gyrations |  |  |  |  | Pass / Fail: | Pass
RAM PRESSURE |  |  |  | HEIGHT
| Place proving ring and 1/8" block in compactor. |  |  |  | Select VERIFY from ram position calibration menu.
| Select VERIFY from ram force calibration menu. |  |  |  | Place blocks under ram as directed.
| Follow screen prompts. |  |  |  | Follow screen prompts.
| Load (N) | Target Dial Rdg. | Actual Dial Rdg. |  | Target (mm) | Actual (mm)
| 1500 | 33.9 |  |  | 254
| 3500 | 79.9 |  |  | 228.6
| 5500 | 124.6 |  |  | 203.2
| 7500 | 169.9 |  |  | 177.8
| 9500 | 215.1 |  |  | 152.4
| 11500 | 261.8 |  |  | 127
| 13500 | 308.2 |  |  | 101.6
| 15500 | 354.1 |  |  | 76.2
| 17500 | 399.7 |  |  | Tolerance: | ± 0.05 mm
|  | Tolerance: | ±1%
| Pass / Fail: |  |  |  | Pass / Fail:
ANGLE
| 1) CHECK ROLLER CLEARANCE. |  |  | 3) VERIFY INTERNAL ANGLE.
| Set dials to zero. Lift gauge at each roller. |  |  |  |  | 1)
|  | Tolerance: | Actual: |  | Calib. Tube | 2)
| Right | 0.0015 - 0.002 |  |  | 1.234° | 3)
| Left | 0.002 - 0.004 |  |  | Pass / Fail | Avg | #DIV/0!
| Back | 0.002 - 0.004 |  |  |  |  | (± 0.01°)
| Pass / Fail : | Pass |  |  | Take 2 measurements on top
and 2 measurements on bottom.
2) CHECK ZERO POSITION.
| Set dials to 0.3500. Spin to 180˚ and read. |  |  |  | T1 = _1.14__ | B1 = _1.15__
| Tolerance: | 0.3500 ± 0.001 |  |  | T2 = _1.14__ | B2 = _1.15__
| Pass / Fail: | Pass |  |  | TAvg = _1.14_ | BAvg = _1.15_
[5 empty] | (Tavg + Bavg) / 2 | Tolerance: | Pass / Fail:
| Initial by: |  |  |  | EFF. INT. ANGLE = __1.15__ | 1.16 ± 0.02˚ | PASS
Maintenance performed according to mfg. instructions
| Initial | Date

## Pine Gyratory Compactor (1)
| Calibration Worksheet: |  | Pine Gyratory Compactor Calibration Verification
| Equipment ID: | Pine Compactor |  |  | Date: | Jan. 18, 2012
| Manufacturer: | Pine Instruments |  |  | Performed by: | Mary Fleck
| Model #: | AFGC125X
| Serial #: | 15 |  |  | Next Calibration Due: | Jan. 2013
| UofA ID: | 229874 |  |  | Last Calibration: | July 2011
| Location: | HM (225)
|  | Calibration Items: | Speed, Pressure, Height, Angle
|  | Calibration Procedure: | Follow Manufacturer's Instructions (summarized below)
|  | Calibration Equipment: | Calibration kit supplied by manufacturer
SPEED
| Select OTHER OPTIONS and press ENTER. |  |  |  |  | Target Time: | 20 seconds
| Select VERIFY CALIBRATION and press ENTER. |  |  |  |  | Actual Time:
| Press START. Allow machine to rotate, then count and time 10 gyrations. |  |  |  |  | Tolerance: | +/- 0.33 sec
| Record stopwatch reading for 10 gyrations. Press ROTATE to park carriage. |  |  |  |  | Pass / Fail: | Pass
RAM PRESSURE |  |  |  | HEIGHT
| Center proving ring and 1/8" block in compactor. |  |  |  | Center 6" gage blocks under ram as directed.
| Press START. Compactor flexes proving ring. |  |  |  | Press START.
| Zero dial. Press START. Follow screen prompts. |  |  |  | Follow screen prompts.
| Load (N) | Target Dial Rdg. | Actual Dial Rdg. |  | Target (mm) | Actual (mm)
| 3500 | 79.9 | 79.3 |  | 152.4 | 152.39
| 14500 | 331.1 | 328.3 |  | Tolerance: | ± 0.05 mm
|  | Tolerance : | ± 1 %
| Pass / Fail: | Pass |  |  | Pass / Fail: | Pass
ANGLE
| 1) CHECK ROLLER CLEARANCE. |  |  | 3) VERIFY INTERNAL ANGLE.
| Insert angle indicating apparatus into chamber. |  |  |  | Place tube over RAM. Set to mode 3.
| Set large dials to zero, small dials to 0.35. |  |  |  | Rotate 1 revolution. Read angle.
Lift device at each roller. Read dial nearest roller.
[5 empty] | Calib. Tube | 1) | 1.24
|  | Tolerance: | Actual: |  | 1.234° | 2) | 1.24
| Right | 0.0015 - 0.002 | 0.0038 |  | Pass / Fail | 3) | 1.24
| Left | 0.002 - 0.004 | 0.0038 |  |  | Avg | 1.24
| Back | 0.002 - 0.004 | 0.0035 |  |  |  | (± 0.01°)
Insert RAM into mold.
| Pass / Fail : | Fail |  |  | Set gyratory # of gyrations to 15.
Turn device off, then on. Lower and start test.
| 2) CHECK ZERO POSITION. |  |  |  | Read angle. Take 2 top and 2 bottom readings.
Set dials to 0.3500. Spin device 180˚.
| Read dials using mirror. |  |  |  | Top | Bottom | Effective Interior
[5 empty] | 1.06 | 1.14 | Angle
| Left Dial | Right Dial | Tolerance |  | 1.07 | 1.15
| 0.35005 | 0.35005 | 0.3500 ± 0.001 |  | 1.065 | 1.145 | 1.105
[5 empty] | Average | Tolerance | 1.16± 0.02°
| Pass / Fail: | Pass |  |  | Pass / Fail: | Fail
Maintenance to be performed by Pine Instruments in March.
| Initial | Date

## Gyratory Mold (Pine G1-G2)
Calibration Worksheet: | [6 empty] | Gyratory Mold
|  |  | Equipment ID: | Pine G1/G2 Molds | [7 empty] | Date: | 2015-09-01
|  |  | Manufacturer: | Pine Instruments | [7 empty] | Performed By: | Nice K.
|  |  | Location: | HM (225)
[12 empty] | Next Calibration Due: | Sept. 2016
|  |  | Calib. Interval: | 12 months | [7 empty] | Last Calibration : | Sept. 2014
[5 empty] | Calibration Items: |  |  | Height, Wall Thickness, Inside Diameter
[5 empty] | Calibration Procedure: |  |  | In-House Procedure # 00 - 4
[5 empty] | Calibration Equipment: |  |  | Mititoyo Calibrated Digimatic Calipers
Mititoyo Calibrated Digimatic 3-Prong Caliper
12" Metal Ruler
Mold # |  | 4
Total Height |  |  |  | Meas. # A |  | Meas. # B |  |  |  |  | Specification :
| (metal ruler) |  |  | 281 | mm | 282 |  | mm |  |  | > 250 mm |  | Pass / Fail
Wall Thickness |  |  |  | Meas. # A |  | Meas. # B |  |  | Meas. # C |  | Specification : |  | Pass
| (standard caliper) |  |  | 7.89 | mm | 7.93 |  | 8.05 | 8.05 | mm | > 7.5 mm
Inside Diameter (mm)
(3-prong caliper) |  |  |  | Meas. # A |  | Meas. # B |  |  | Meas. # C |  | Specification :
# 1 (2" from Top) |  |  |  | 149.93 | mm | 149.94 |  | mm | 149.93 | mm | ≤ 150.2 mm |  | Initial
# 2 (4" from Bottom) |  |  |  | 149.93 | mm | 149.93 |  | mm | 149.92 | mm | ≤ 150.2 mm |  |  | _______
# 3 ( 2" from Bottom) |  |  |  | 149.93 | mm | 149.93 |  | mm | 149.93 | mm | ≤ 150.2 mm
Comments:
Mold # |  | 5
Total Height |  |  |  | Meas. # A |  | Meas. # B |  |  |  |  | Specification :
| (metal ruler) |  |  | 280 | mm | 280 |  | mm |  |  | > 250 mm |  | Pass / Fail
Wall Thickness |  |  |  | Meas. # A |  | Meas. # B |  |  | Meas. # C |  | Specification : |  | Pass
| (standard caliper) |  |  | 7.82 | mm | 7.9 |  | mm | 7.88 | mm | > 7.5 mm
Inside Diameter (mm)
(3-prong caliper) |  |  |  | Meas. # A |  | Meas. # B |  |  | Meas. # C |  | Specification :
# 1 (2" from Top) |  |  |  | 149.94 | mm | 149.93 |  | mm | 149.92 | mm | ≤ 150.2 mm |  | Initial
# 2 (4" from Bottom) |  |  |  | 149.92 | mm | 149.92 |  | mm | 149.92 | mm | ≤ 150.2 mm |  |  | _______
# 3 ( 2" from Bottom) |  |  |  | 149.92 | mm | 149.92 |  | mm | 149.93 | mm | ≤ 150.2 mm
Comments:
Mold #
Total Height |  |  |  | Meas. # A |  | Meas. # B |  |  |  |  | Specification :
| (metal ruler) |  |  |  | mm |  |  | mm |  |  | > 250 mm |  | Pass / Fail
Wall Thickness |  |  |  | Meas. # A |  | Meas. # B |  |  | Meas. # C |  | Specification :
| (standard caliper) |  |  |  | mm |  |  | mm |  | mm | > 7.5 mm
Inside Diameter (mm)
(3-prong caliper) |  |  |  | Meas. # A |  | Meas. # B |  |  | Meas. # C |  | Specification :
# 1 (2" from Top) |  |  |  |  | mm |  |  | mm |  | mm | ≤ 150.2 mm |  | Initial
# 2 (4" from Bottom) |  |  |  |  | mm |  |  | mm |  | mm | ≤ 150.2 mm |  |  | _______
# 3 ( 2" from Bottom) |  |  |  |  | mm |  |  | mm |  | mm | ≤ 150.2 mm
Comments:

## Pine G Bottom Plates
Calibration Worksheet: |  | Pine G Series Bottom Plates
Equipment ID: | Pine G Bottom Plates |  | Date: | 2015-09-01
Manufacturer: | Pine Instruments |  | Performed by: | Nice K.
Location: | ST (218)
|  |  | Next Calibration Due: | September 2016
|  |  | Last Calibration: | September 2014
| Calibration Items: | Plate Diameter
| Calibration Procedure: | In-House Procedure
Calibration Equipment:
| Plate # 093812 |  | Plate # 093813
| Bottom Plate Diameter (mm) |  | Bottom Plate Diameter (mm)
| Meas. 1 | 149.51 | Meas. 1 | 149.67
| Meas. 2 | 149.64 | Meas. 2 | 149.69
| Meas. 3 | 149.66 | Meas. 3 | 149.63
| Meas. 4 | 149.56 | Meas. 4 | 149.65
| Avg: | 149.5925 | Avg: | 149.66
| Acc. Range: | 149.50-149.75 | Acc. Range: | 149.50-149.75
| Pass/Fail: | Pass | Pass/Fail: | Pass
| Bottom Plate Diameter (mm) |  | Bottom Plate Diameter (mm)
| Meas. 1 |  | Meas. 1
| Meas. 2 |  | Meas. 2
| Meas. 3 |  | Meas. 3
| Meas. 4 |  | Meas. 4
| Avg: |  | Avg:
| Acc. Range: |  | Acc. Range:
| Pass/Fail: |  | Pass/Fail:
Initial

## CoreLok 1
Calibration Worksheet: |  | CoreLok 1
Equipment ID: | Instrotek (SN 1028) |  |  | Date: | 2016-07-01
Manufacturer: | Instrotek |  |  | Performed by:
Location: | Corelok Room
|  |  |  | Next Calibration Due: | July 2017
|  |  |  | Last Calibration: | July 2015
| Calibration Items: | Density
| Calibration Procedure: | ASTM 6752
| Calibration Equipment: | 4.75mm specimen prepared to 4-8% air voids at 6% binder*
*Use existing calibration specimen: 4.02% air voids (Gmm = 2.362)
CoreLok | Rep 1 | Rep 2 | Rep 3
Initial wt of Core
Sealed core in air
Sealed core in water
Final wt of Core
Bag Density
Bulk Specific Gravity | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
SSD |  |  |  | Vacuum System Check : (3 month interval)
dry wt of core |  |  |  | Gauge : Instro Tek : SN 09070056
subm wt of core |  |  |  |  | Gauge Reading (torr) |  | Perf. By
SSD wt of core |  |  |  |  | Jan. 2016 | 3.5 | M. Fleck
Bulk Specific Gravity | #DIV/0! |  |  |  | Apr. 2016
Jul. 2016
Oct. 2016
Difference |  |  |  | Allowable tolerance = < 10 torr
SSD - CL | #DIV/0! | Pass
Allowable tolerance = +/- 0.010 |  | Pass / Fail |  |  |  | Initial by:

## CoreLok 2
Calibration Worksheet: |  | CoreLok 2
Equipment ID: | Gilson (SN 1004) |  |  | Date: | 2016-07-01
Manufacturer: | Instrotek |  |  | Performed by:
Location: | Corelok Room
|  |  |  | Next Calibration Due: | July 2017
|  |  |  | Last Calibration: | July 2015
| Calibration Items: | Density
| Calibration Procedure: | ASTM 6752
| Calibration Equipment: | 4.75mm specimen prepared to 4-8% air voids at 6% binder
*Use existing calibration specimen: 4.02% air voids (Gmm = 2.362)
CoreLok | Rep 1 | Rep 2 | Rep 3
Initial wt of Core
Sealed core in air
Sealed core in water
Final wt of Core
Bag Density
Bulk Specific Gravity | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
SSD |  |  |  | Vacuum System Check : (3 month interval)
Dry wt of core |  |  |  | Gauge : Instro Tek : SN 09070056
Subm wt of core |  |  |  |  | Gauge Reading (torr) |  | Perf. By
SSD wt of core |  |  |  |  | Jan. 2016 | 5.5 | M. Fleck
Bulk Specific Gravity | #DIV/0! |  |  |  | Apr. 2016
Jul. 2016
Oct. 2016
Difference |  |  |  | Allowable tolerance = < 10 torr
SSD - CL | #DIV/0! | Pass
Allowable tolerance = +/- 0.010 |  | Pass / Fail |  |  |  | Initial by:

## CoreLok
BULK SPECIFIC GRAVITY BY THE CORELOK METHOD
Project: | Bag Density |  | Date: |  | 2016-07-01
Source: | Verification Patty |  | Tested By:
Sample Number | 1 | 2 | 3 | 4 | 5 | 6
Wt. of Dry Core in Air before testing, g (A)
Wt. of Sealed Core in Air, g (B)
Wt of Sealed Core in Water, g (C )
Wt. of Dry Core in Air after testing, g (D)
Bag Weight E = (B - A) | 0 | 0 | 0 | 0 | 0 | 0
Bag Ratio F = A / E | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Large Bag Volume Correction (-0.00166*F+0.8596) (G) | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Total Volume H = (E + D) - C | 0 | 0 | 0 | 0 | 0 | 0
Bag Volume I = E / G | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Sample Volume J = H - I | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Bulk Specific Gravity (Gmb) K = A / J | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Check: % wt. change (must be -0.08% to +0.04%) | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Max. Sp. Gravity, Gmm (from AASHTO T-209) (J)
% Air Voids (AASHTO T-269) 100 * (1 - H/J) | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Computed By: |  |  | Checked By: