## Instructions
Scale Calibration Instructions
Ohaus EP12001 |  |  |  |  | AND HV60KGL
| → Main Menu |  |  |  |  | Remove back plate
| Enter |  |  |  |  | Press & hold Cal button
| ↕↔ Calibration |  |  |  |  | (black button on rim)
| Enter |  |  |  |  | Enter (Pan = 0)
| ↕ Span |  |  |  |  | Load 60 kg
| Enter |  |  |  |  | Enter
| Enter |  |  |  |  | Remove mass
| Place 10 kg on scales |  |  |  |  | Press Cal button
| Enter |  |  |  |  | Replace plate
Press any key to exit
AND FG 60KAL
I/O Ohaus GT 8000 | [5 empty] | Remove back plate
| I/O Off |  |  |  |  | Press & hold Cal button
| Open door (leave open) |  |  |  |  | Press mode (SPAN)
| Remove ceramic plate & put calibration |  |  |  |  | Load 60 kg
| on ceramic pegs |  |  |  |  | Press Mode
| Hold #4 and turn on I/O until beep |  |  |  |  | Remove mass
| Use paper clip to press ON button |  |  |  |  | Replace plate
(SPAN)
| Press ON |  |  |  | Metttler SG32000
| Press ON (load 8 kg) |  |  |  |  | Hold Cal button
| Press ON |  |  |  |  | (Int. Cal Set)
| Power OFF |  |  |  |  | Scale does it's thing
| Replace ceramic plate |  |  |  |  | (this stops cal flashing)
| . | . | . | .
| Print | Mode | Off | On/Tare
Adam Highland HCB302 |  |  |  |  | Mettler MS32000-L 01
| Press ON/OFF button |  |  |  |  | Press WEIGHT button
| Press [Smpl] & [Print] at same time |  |  |  |  | (Int. Cal Set)
| during self-check |  |  |  |  | Scale does it's thing
Press [Tare] button when stable
| Change calib. weight by pressing |  |  |  | Adam CPW Plus Scales (Concrete/Case)
| [Mode] if needed |  |  |  |  | Pres and hold [Tare] for 4 seconds
| Press [Tare] button to select weight |  |  |  |  | Press [Unit] button to change units
| Load weight |  |  |  |  | Press [Print] button
| Press [Tare] when stable |  |  |  |  | When calibration mass shows press
| If PASS shows, remove weight |  |  |  |  | [Unit] key to confirm
Load weight
Press [Unit] button
Remove weight

## Scale List
Current Scale List 11/13/15
Scale | Mfg | Model # | Serial # | U of A ID | Class | Capacity | Readability | Date In
[6 empty] | (g) | (g) | Service
G1 - 1 | ADAM | HCB 302 | AE76304522 | NA | G1 | 300 | 0.01 | 2015-11-13
G1 - 2 | OHAUS | EOD 120 | E2411119231431 | 250795 | G1 | 4100 | 0.01 | ?
AG - A | OHAUS | EP12001 | 1126012086 | NA | G2 | 12000 | 0.1 | ?
AG - B | OHAUS | EP12001 | 1125430302 | NA | G2 | 12000 | 0.1 | ?
AG - E | OHAUS | EP12001 | 1125330661 | NA | G2 | 12000 | 0.1 | ?
AG - F | OHAUS | EP12001 | 1126051479 | NA | G2 | 12000 | 0.1 | ?
CL - 1 | OHAUS | EP12001 | 1128061460 | NA | G2 | 12000 | 0.1 | ?
CL - 2 | OHAUS | EP12001 | 1128061463 | NA | G2 | 12000 | 0.1 | ?
G2 - X | OHAUS | EX24001 | B525089329 | 2692501 | G2 | 24000 | 0.1 | 2015-09-09
IGN. OVEN | OHAUS | GT 8000 | 945960260909 | 234502 | G2 | 8000 | 0.01 | ?
MX - 3 | METTLER | MS32000L / 01 | B508635397 | 2692698 | G5 | 32000 | 1 | 2015-03-23
MX - 2 | METTLER | SG32000 | 1120303652 | NA | G5 | 32000 | 1 | ?
C - 2 | AND | FG60KAM | Q1911021 | NA | G100 | 60000 | 20 | 2008-09-05
Calibrations should cover at least 90% of the scale's capacity. Calibration weights should be verified every 12 years.
Extra Scales in Storage:
| OHAUS | EP12001 | 1126051480 | NA | G2 | 12000 | 0.1
| OHAUS | EP12001 | 1125373087 | NA | G2 | 12000 | 0.1
| OHAUS | EP12001 | 1125373345 | NA | G2 | 12000 | 0.1
| OHAUS | EP12001 | 1126012088 | NA | G2 | 12000 | 0.1
| OHAUS | GT 400 | 5376 | NA | G1 | 400 | 0.01
| OHAUS | GT 400 | 5373 | NA | G1 | 400 | 0.01
| ADAM | CPW Plus 35 | AE39919806 | NA | G100 | 35000 | 10
| ADAM | CPW Plus 35 | AE39919807 | NA | G100 | 35000 | 10

## Scale (G1-1)
Scale Calibration Worksheet
| Equipment ID: | G1 - 1 |  |  |  |  | Date: | 2015-09-13
| Manufacturer: | Adam |  |  |  |  | Performed by: | M. Fleck
| Model #: | HCB 302
| Serial #: | AE76304522 |  |  |  |  | Next Calibration Due: | September 2016
| UofA ID: | NA |  |  |  |  | Last Calibration: | New
| Location: | STO (218)
| Calibration Item: | Verify mass readings for at least 5 points in the range of the balance
Check sensitivity (0.1%) for at least 5 points in the range of the balance
| Calibration Procedure: | AASHTO M-231
| Calibration Equipment: | Class F, Class I, and 10-kg Calibration Weights
Tolerances:
| Class |  | Readability and Sensitivity |  |  |  | Accuracy
| G1 |  |  | 0.01 g |  | 0.02 g or 0.1 percent
| G2 |  |  | 0.1 g |  | 0.2 or 0.1 percent
| G5 |  |  | 1 g |  | 2 g or 0.1 percent
| G20 |  |  | 5 g |  | 5 g or 0.1 percent
| G100 |  |  | 20 g |  | 20 g or 0.1 percent
| Weight Applied (g) |  | Scale Reading (g) |  | Tolerance (g) |  |  | Pass/Fail
| 1 |  | 1 |  | 0.02
| 1.01 |  | 1.01 |  | 0.02
| 5 |  | 5 |  | 0.02
| 10 |  | 10 |  | 0.02
| 10.01 |  | 10.01 |  | 0.02
| 20 |  | 20 |  | 0.02
| 20.01 |  | 20.01 |  | 0.02
| 50 |  | 50 |  | 0.05
| 50.01 |  | 50.01 |  | 0.05
| 100 |  | 100 |  | 0.1
| 100.01 |  | 100.01 |  | 0.1
| 200 |  | 200 |  | 0.2
| 300 |  | 300 |  | 0.3
[9 empty] | Pass/Fail: | Pass
Comments:
Initial By:

## Scale (G1-2)
Scale Calibration Worksheet
| Equipment ID: | G1 - 2 |  |  |  |  | Date: | 2015-09-08
| Manufacturer: | Ohaus |  |  |  |  | Performed by: | Danielle N.
| Model #: | EOD 120
| Serial #: | E2411119231431 |  |  |  |  | Next Calibration Due: | September 2016
| UofA ID: | [5 empty] | Last Calibration: | September 2014
| Location: | STO (218)
| Calibration Item: | Verify mass readings for at least 5 points in the range of the balance
Check sensitivity (0.1%) for at least 5 points in the range of the balance
| Calibration Procedure: | AASHTO M-231
| Calibration Equipment: | Class F, Class I, and 10-kg Calibration Weights
Tolerances:
| Class |  | Readability and Sensitivity |  |  |  | Accuracy
| G1 |  |  | 0.01 g |  | 0.02 g or 0.1 percent
| G2 |  |  | 0.1 g |  | 0.2 or 0.1 percent
| G5 |  |  | 1 g |  | 2 g or 0.1 percent
| G20 |  |  | 5 g |  | 5 g or 0.1 percent
| G100 |  |  | 20 g |  | 20 g or 0.1 percent
| Weight Applied (g) |  | Scale Reading (g) |  | Tolerance (g) |  |  | Pass/Fail
| 1 |  | 1 |  | 0.02
| 1.01 |  | 1.01 |  | 0.02
| 5 |  | 5 |  | 0.02
| 10 |  | 10 |  | 0.02
| 20 |  | 20 |  | 0.02
| 50 |  | 50 |  | 0.05
| 100 |  | 100.01 |  | 0.1
| 500 |  | 500 |  | 0.5
| 500.01 |  | 500.01 |  | 0.5
| 1000 |  | 999.98 |  | 1
| 1000.01 |  | 999.99 |  | 1
| 4000 |  | 4000.07 |  | 4
| 4000.01 |  | 4000.08 |  | 4
[9 empty] | Pass/Fail: | Pass
Comments:
Initial By:

## Scale (G1-3)
Scale Calibration Worksheet
| Equipment ID: | G1 - 3 |  |  |  |  | Date: | 2015-09-08
| Manufacturer: | Ohaus |  |  |  |  | Performed by: | Danielle N.
| Model #: | GT 400
| Serial #: | 5376 |  |  |  |  | Next Calibration Due: | September 2016
| UofA ID: | NA |  |  |  |  | Last Calibration: | September 2014
| Location: | STO (218)
| Calibration Item: | Verify mass readings for at least 5 points in the range of the balance
Check sensitivity (0.1%) for at least 5 points in the range of the balance
| Calibration Procedure: | AASHTO M-231
| Calibration Equipment: | Class F, Class I, and 10-kg Calibration Weights
Tolerances:
| Class |  | Readability and Sensitivity |  |  |  | Accuracy
| G1 |  |  | 0.01 g |  | 0.02 g or 0.1 percent
| G2 |  |  | 0.1 g |  | 0.2 or 0.1 percent
| G5 |  |  | 1 g |  | 2 g or 0.1 percent
| G20 |  |  | 5 g |  | 5 g or 0.1 percent
| G100 |  |  | 20 g |  | 20 g or 0.1 percent
| Weight Applied (g) |  | Scale Reading (g) |  | Tolerance (g) |  |  | Pass/Fail
| 0.01 |  | 0.01 |  | 0.02
| 1 |  | 1 |  | 0.02
| 5 |  | 5.01 |  | 0.02
| 10 |  | 10 |  | 0.02
| 10.01 |  | 10.01 |  | 0.02
| 50 |  | 50.02 |  | 0.05
| 50.01 |  | 50.03 |  | 0.05
| 100 |  | 100.05 |  | 0.1
| 100.01 |  | 100.06 |  | 0.1
| 200 |  | 200.08 |  | 0.2
| 200.01 |  | 200.09 |  | 0.2
| 300 |  | 300.12 |  | 0.3
| 400 |  | 400.16 |  | 0.4
[9 empty] | Pass/Fail: | Pass
Comments:
Initial By:

## Scale (G1-4)
Scale Calibration Worksheet
| Equipment ID: | G1 - 4 |  |  |  |  | Date: | 2015-09-08
| Manufacturer: | Ohaus |  |  |  |  | Performed by: | Danielle N.
| Model #: | GT 400
| Serial #: | 5373 |  |  |  |  | Next Calibration Due: | September 2016
| UofA ID: | NA |  |  |  |  | Last Calibration: | September 2014
| Location: | STO (218)
| Calibration Item: | Verify mass readings for at least 5 points in the range of the balance
Check sensitivity (0.1%) for at least 5 points in the range of the balance
| Calibration Procedure: | AASHTO M-231
| Calibration Equipment: | Class F, Class I, and 10-kg Calibration Weights
Tolerances:
| Class |  | Readability and Sensitivity |  |  |  | Accuracy
| G1 |  |  | 0.01 g |  | 0.02 g or 0.1 percent
| G2 |  |  | 0.1 g |  | 0.2 or 0.1 percent
| G5 |  |  | 1 g |  | 2 g or 0.1 percent
| G20 |  |  | 5 g |  | 5 g or 0.1 percent
| G100 |  |  | 20 g |  | 20 g or 0.1 percent
| Weight Applied (g) |  | Scale Reading (g) |  | Tolerance (g) |  |  | Pass/Fail
| 0.01 |  | 0.01 |  | 0.02
| 1 |  | 1 |  | 0.02
| 5 |  | 5 |  | 0.02
| 10 |  | 10 |  | 0.02
| 10.01 |  | 10.01 |  | 0.02
| 50 |  | 50.01 |  | 0.05
| 50.01 |  | 50.02 |  | 0.05
| 100 |  | 100.03 |  | 0.1
| 100.01 |  | 100.04 |  | 0.1
| 200 |  | 200.05 |  | 0.2
| 200.01 |  | 200.06 |  | 0.2
| 300 |  | 300.07 |  | 0.3
| 400 |  | 400.09 |  | 0.4
[9 empty] | Pass/Fail: | Pass
Comments:
Initial By:

## Scale (G2-X)
Scale Calibration Worksheet
| Equipment ID: | G2 - X |  |  |  |  | Date: | 2015-09-09
| Manufacturer: | Ohaus |  |  |  |  | Performed by: | M. Fleck
| Model #: | EX24001
| Serial #: | B525089329 |  |  |  |  | Next Calibration Due: | September 2016
| UofA ID: | NA |  |  |  |  | Last Calibration: | September 2014
| Location: | STO (218)
| Calibration Item: | Verify mass readings for at least 5 points in the range of the balance
Check sensitivity (0.1%) for at least 5 points in the range of the balance
| Calibration Procedure: | AASHTO M-231
| Calibration Equipment: | Class F, Class I, and 10-kg Calibration Weights
Tolerances:
| Class |  | Readability and Sensitivity |  |  |  | Accuracy
| G1 |  |  | 0.01 g |  | 0.02 g or 0.1 percent
| G2 |  |  | 0.1 g |  | 0.2 or 0.1 percent
| G5 |  |  | 1 g |  | 2 g or 0.1 percent
| G20 |  |  | 5 g |  | 5 g or 0.1 percent
| G100 |  |  | 20 g |  | 20 g or 0.1 percent
| Weight Applied (g) |  | Scale Reading (g) |  | Tolerance (g) |  |  | Pass/Fail
| 10 |  | 10 |  | 0.2
| 10.1 |  | 10.1 |  | 0.2
| 200 |  | 200 |  | 0.2
| 500 |  | 500 |  | 0.5
| 1000 |  | 1000 |  | 1
| 1000.1 |  | 1000.1 |  | 1
| 2000 |  | 2000 |  | 2
| 5000 |  | 5000 |  | 5
| 5000.1 |  | 5000.1 |  | 5
| 10000 |  | 10000 |  | 10
| 10000.1 |  | 10000.1 |  | 10
| 15000 |  | 15000 |  | 15
| 20000 |  | 20000.1 |  | 20
| 20000.1 |  | 20000.2 |  | 20
| 24000 |  | 24000.2 |  | 24
[9 empty] | Pass/Fail: | Pass
Comments:
Initial By:

## Scale (G2-A)
Scale Calibration Worksheet
| Equipment ID: | AG - A |  |  |  |  | Date: | 2015-09-01
| Manufacturer: | Ohaus |  |  |  |  | Performed by: | Danielle N.
| Model #: | EP12001
| Serial #: | 1126012086 |  |  |  |  | Next Calibration Due: | September 2016
| UofA ID: | NA |  |  |  |  | Last Calibration: | September 2014
| Location: | AG (233) Rear Table (West)
| Calibration Item: | Verify mass readings for at least 5 points in the range of the balance
Check sensitivity (0.1%) for at least 5 points in the range of the balance
| Calibration Procedure: | AASHTO M-231
| Calibration Equipment: | Class F, Class I, and 10-kg Calibration Weights
Tolerances:
| Class |  | Readability and Sensitivity |  |  |  | Accuracy
| G1 |  |  | 0.01 g |  | 0.02 g or 0.1 percent
| G2 |  |  | 0.1 g |  | 0.2 or 0.1 percent
| G5 |  |  | 1 g |  | 2 g or 0.1 percent
| G20 |  |  | 5 g |  | 5 g or 0.1 percent
| G100 |  |  | 20 g |  | 20 g or 0.1 percent
| Weight Applied (g) |  | Scale Reading (g) |  | Tolerance (g) |  |  | Pass/Fail
| 0.1 |  | 0.1 |  | 0.2
| 1 |  | 1 |  | 0.2
| 10 |  | 10.1 |  | 0.2
| 100 |  | 100 |  | 0.2
| 100.1 |  | 100.1 |  | 0.2
| 200 |  | 200 |  | 0.2
| 500 |  | 500.1 |  | 0.5
| 500.1 |  | 500.2 |  | 0.5
| 1000 |  | 1000 |  | 1
| 2000 |  | 2000.1 |  | 2
| 2000.1 |  | 2000.2 |  | 2
| 5000 |  | 5000 |  | 5
| 5000.1 |  | 5000.1 |  | 5
| 10000 |  | 10000.1 |  | 10
| 11000 |  | 11000.2 |  | 11
[9 empty] | Pass/Fail: | Pass
Comments:
Initial By:

## Scale (G2-B)
Scale Calibration Worksheet
| Equipment ID: | AG - B |  |  |  |  | Date: | 2015-09-01
| Manufacturer: | Ohaus |  |  |  |  | Performed by: | Danielle N.
| Model #: | EP12001
| Serial #: | 1125430302 |  |  |  |  | Next Calibration Due: | September 2016
| UofA ID: | NA |  |  |  |  | Last Calibration: | September 2014
| Location: | AG (233) Front Table (East)
| Calibration Item: | Verify mass readings for at least 5 points in the range of the balance
Check sensitivity (0.1%) for at least 5 points in the range of the balance
| Calibration Procedure: | AASHTO M-231
| Calibration Equipment: | Class F, Class I, and 10-kg Calibration Weights
Tolerances:
| Class |  | Readability and Sensitivity |  |  |  | Accuracy
| G1 |  |  | 0.01 g |  | 0.02 g or 0.1 percent
| G2 |  |  | 0.1 g |  | 0.2 or 0.1 percent
| G5 |  |  | 1 g |  | 2 g or 0.1 percent
| G20 |  |  | 5 g |  | 5 g or 0.1 percent
| G100 |  |  | 20 g |  | 20 g or 0.1 percent
| Weight Applied (g) |  | Scale Reading (g) |  | Tolerance (g) |  |  | Pass/Fail
| 0.1 |  | 0.1 |  | 0.2
| 1 |  | 1 |  | 0.2
| 10 |  | 10 |  | 0.2
| 100 |  | 100 |  | 0.2
| 100.1 |  | 100.1 |  | 0.2
| 200 |  | 200 |  | 0.2
| 500 |  | 500 |  | 0.5
| 500.1 |  | 500.1 |  | 0.5
| 1000 |  | 1000 |  | 1
| 2000 |  | 2000.1 |  | 2
| 2000.1 |  | 2000.2 |  | 2
| 5000 |  | 5000.3 |  | 5
| 5000.1 |  | 5000.4 |  | 5
| 10000 |  | 10000.6 |  | 10
| 11000 |  | 11000.7 |  | 11
[9 empty] | Pass/Fail: | Pass
Comments:
Initial By:

## Scale (G2-E)
Scale Calibration Worksheet
| Equipment ID: | AG - E |  |  |  |  | Date: | 2015-09-08
| Manufacturer: | Ohaus |  |  |  |  | Performed by: | Danielle N.
| Model #: | EP12001
| Serial #: | 1125330661 |  |  |  |  | Next Calibration Due: | September 2016
| UofA ID: | NA |  |  |  |  | Last Calibration: | September 2014
| Location: | AG (233) Front SpG (East)
| Calibration Item: | Verify mass readings for at least 5 points in the range of the balance
Check sensitivity (0.1%) for at least 5 points in the range of the balance
| Calibration Procedure: | AASHTO M-231
| Calibration Equipment: | Class F, Class I, and 10-kg Calibration Weights
Tolerances:
| Class |  | Readability and Sensitivity |  |  |  | Accuracy
| G1 |  |  | 0.01 g |  | 0.02 g or 0.1 percent
| G2 |  |  | 0.1 g |  | 0.2 or 0.1 percent
| G5 |  |  | 1 g |  | 2 g or 0.1 percent
| G20 |  |  | 5 g |  | 5 g or 0.1 percent
| G100 |  |  | 20 g |  | 20 g or 0.1 percent
| Weight Applied (g) |  | Scale Reading (g) |  | Tolerance (g) |  |  | Pass/Fail
| 0.1 |  | 0.1 |  | 0.2
| 1 |  | 1.1 |  | 0.2
| 10 |  | 10 |  | 0.2
| 100 |  | 100 |  | 0.2
| 100.1 |  | 100.1 |  | 0.2
| 200 |  | 200 |  | 0.2
| 500 |  | 500 |  | 0.5
| 500.1 |  | 500.1 |  | 0.5
| 1000 |  | 1000 |  | 1
| 2000 |  | 2000 |  | 2
| 2000.1 |  | 2000.1 |  | 2
| 5000 |  | 5000 |  | 5
| 5000.1 |  | 5000.1 |  | 5
| 10000 |  | 9999.9 |  | 10
| 11000 |  | 11000 |  | 11
[9 empty] | Pass/Fail: | Pass
Comments:
Initial By:

## Scale (G2-F)
Scale Calibration Worksheet
| Equipment ID: | AG - F |  |  |  |  | Date: | 2015-09-08
| Manufacturer: | Ohaus |  |  |  |  | Performed by: | Danielle N.
| Model #: | EP12001
| Serial #: | 1126051479 |  |  |  |  | Next Calibration Due: | September 2016
| UofA ID: | NA |  |  |  |  | Last Calibration: | September 2014
| Location: | AG (233) Rear SpG (West)
| Calibration Item: | Verify mass readings for at least 5 points in the range of the balance
Check sensitivity (0.1%) for at least 5 points in the range of the balance
| Calibration Procedure: | AASHTO M-231
| Calibration Equipment: | Class F, Class I, and 10-kg Calibration Weights
Tolerances:
| Class |  | Readability and Sensitivity |  |  |  | Accuracy
| G1 |  |  | 0.01 g |  | 0.02 g or 0.1 percent
| G2 |  |  | 0.1 g |  | 0.2 or 0.1 percent
| G5 |  |  | 1 g |  | 2 g or 0.1 percent
| G20 |  |  | 5 g |  | 5 g or 0.1 percent
| G100 |  |  | 20 g |  | 20 g or 0.1 percent
| Weight Applied (g) |  | Scale Reading (g) |  | Tolerance (g) |  |  | Pass/Fail
| 0.1 |  | 0.1 |  | 0.2
| 1 |  | 1 |  | 0.2
| 10 |  | 10 |  | 0.2
| 100 |  | 100 |  | 0.2
| 100.1 |  | 100.1 |  | 0.2
| 200 |  | 200 |  | 0.2
| 500 |  | 500 |  | 0.5
| 500.1 |  | 500.1 |  | 0.5
| 1000 |  | 1000 |  | 1
| 2000 |  | 2000.1 |  | 2
| 2000.1 |  | 2000.2 |  | 2
| 5000 |  | 5000 |  | 5
| 5000.1 |  | 5000.1 |  | 5
| 10000 |  | 10000.1 |  | 10
| 11000 |  | 11000.2 |  | 11
[9 empty] | Pass/Fail: | Pass
Comments:
Initial By:

## Scale (G2-CL1)
Scale Calibration Worksheet
| Equipment ID: | CL - 1 |  |  |  |  | Date: | 2015-09-01
| Manufacturer: | Ohaus |  |  |  |  | Performed by: | Danielle N.
| Model #: | EP12001
| Serial #: | 1128061460 |  |  |  |  | Next Calibration Due: | September 2016
| UofA ID: | NA |  |  |  |  | Last Calibration: | September 2014
| Location: | CL (220) Left SpG
| Calibration Item: | Verify mass readings for at least 5 points in the range of the balance
Check sensitivity (0.1%) for at least 5 points in the range of the balance
| Calibration Procedure: | AASHTO M-231
| Calibration Equipment: | Class F, Class I, and 10-kg Calibration Weights
Tolerances:
| Class |  | Readability and Sensitivity |  |  |  | Accuracy
| G1 |  |  | 0.01 g |  | 0.02 g or 0.1 percent
| G2 |  |  | 0.1 g |  | 0.2 or 0.1 percent
| G5 |  |  | 1 g |  | 2 g or 0.1 percent
| G20 |  |  | 5 g |  | 5 g or 0.1 percent
| G100 |  |  | 20 g |  | 20 g or 0.1 percent
| Weight Applied (g) |  | Scale Reading (g) |  | Tolerance (g) |  |  | Pass/Fail
| 0.1 |  | 0.2 |  | 0.2
| 1 |  | 1 |  | 0.2
| 10 |  | 10 |  | 0.2
| 100 |  | 100 |  | 0.2
| 100.1 |  | 100.1 |  | 0.2
| 200 |  | 200 |  | 0.2
| 500 |  | 500 |  | 0.5
| 500.1 |  | 500.1 |  | 0.5
| 1000 |  | 999.9 |  | 1
| 2000 |  | 1999.9 |  | 2
| 2000.1 |  | 2000 |  | 2
| 5000 |  | 4999.9 |  | 5
| 5000.1 |  | 5000 |  | 5
| 10000 |  | 9999.8 |  | 10
| 11000 |  | 10999.8 |  | 11
[9 empty] | Pass/Fail: | Pass
Comments:
Initial By:

## Scale (G2-CL2)
Scale Calibration Worksheet
| Equipment ID: | CL - 2 |  |  |  |  | Date: | 2015-09-01
| Manufacturer: | Ohaus |  |  |  |  | Performed by: | Danielle N.
| Model #: | EP12001
| Serial #: | 1128061463 |  |  |  |  | Next Calibration Due: | September 2016
| UofA ID: | NA |  |  |  |  | Last Calibration: | September 2014
| Location: | CL (220) Right SpG
| Calibration Item: | Verify mass readings for at least 5 points in the range of the balance
Check sensitivity (0.1%) for at least 5 points in the range of the balance
| Calibration Procedure: | AASHTO M-231
| Calibration Equipment: | Class F, Class I, and 10-kg Calibration Weights
Tolerances:
| Class |  | Readability and Sensitivity |  |  |  | Accuracy
| G1 |  |  | 0.01 g |  | 0.02 g or 0.1 percent
| G2 |  |  | 0.1 g |  | 0.2 or 0.1 percent
| G5 |  |  | 1 g |  | 2 g or 0.1 percent
| G20 |  |  | 5 g |  | 5 g or 0.1 percent
| G100 |  |  | 20 g |  | 20 g or 0.1 percent
| Weight Applied (g) |  | Scale Reading (g) |  | Tolerance (g) |  |  | Pass/Fail
| 0.1 |  | 0.1 |  | 0.2
| 1 |  | 1 |  | 0.2
| 10 |  | 10 |  | 0.2
| 100 |  | 100 |  | 0.2
| 100.1 |  | 100.1 |  | 0.2
| 200 |  | 200 |  | 0.2
| 500 |  | 500 |  | 0.5
| 500.1 |  | 500.1 |  | 0.5
| 1000 |  | 1000 |  | 1
| 2000 |  | 2000 |  | 2
| 2000.1 |  | 2000.1 |  | 2
| 5000 |  | 4999.8 |  | 5
| 5000.1 |  | 4999.9 |  | 5
| 10000 |  | 9999.6 |  | 10
| 11000 |  | 10999.6 |  | 11
[9 empty] | Pass/Fail: | Pass
Comments:
Initial By:

## Scale (G2-IO)
Scale Calibration Worksheet
| Equipment ID: | Ignition Oven |  |  |  |  | Date: | 2015-09-02
| Manufacturer: | Ohaus |  |  |  |  | Performed by: | M. Fleck
| Model #: | GT 8000
| Serial #: | 945960260909 |  |  |  |  | Next Calibration Due: | September 2016
| UofA ID: | NA |  |  |  |  | Last Calibration: | September 2014
| Location: | HM (225) Ignition Oven
| Calibration Item: | Verify mass readings for at least 5 points in the range of the balance
Check sensitivity (0.1%) for at least 5 points in the range of the balance
| Calibration Procedure: | AASHTO M-231
| Calibration Equipment: | Class F, Class I, and 10-kg Calibration Weights
Tolerances:
| Class |  | Readability and Sensitivity |  |  |  | Accuracy
| G1 |  |  | 0.01 g |  | 0.02 g or 0.1 percent
| G2 |  |  | 0.1 g |  | 0.2 or 0.1 percent
| G5 |  |  | 1 g |  | 2 g or 0.1 percent
| G20 |  |  | 5 g |  | 5 g or 0.1 percent
| G100 |  |  | 20 g |  | 20 g or 0.1 percent
| Weight Applied (g) |  | Scale Reading (g) |  | Tolerance (g) |  |  | Pass/Fail
| 0.1 |  | 0.1 |  | 0.2
| 1 |  | 1 |  | 0.2
| 10 |  | 10 |  | 0.2
| 100 |  | 100.1 |  | 0.2
| 100.1 |  | 100.2 |  | 0.2
| 200 |  | 200.1 |  | 0.2
| 500 |  | 500.1 |  | 0.5
| 500.1 |  | 500.2 |  | 0.5
| 1000 |  | 1000.2 |  | 1
| 2000 |  | 2000.3 |  | 2
| 2000.1 |  | 2000.4 |  | 2
| 5000 |  | 5000.4 |  | 5
| 5000.1 |  | 5000.5 |  | 5
| 8000 |  | 8000.2 |  | 8
[9 empty] | Pass/Fail: | Pass
Comments:
Initial By:

## Scale (G5-MX2)
Scale Calibration Worksheet
| Equipment ID: | MX - 2 |  |  |  |  | Date: | 2015-09-08
| Manufacturer: | Mettler |  |  |  |  | Performed by: | Danielle N.
| Model #: | SG 32000
| Serial #: | 1120303652 |  |  |  |  | Next Calibration Due: | September 2016
| UofA ID: | NA |  |  |  |  | Last Calibration: | September 2014
| Location: | MX (234) (Left)
| Calibration Item: | Verify mass readings for at least 5 points in the range of the balance
Check sensitivity (0.1%) for at least 5 points in the range of the balance
| Calibration Procedure: | AASHTO M-231
| Calibration Equipment: | Class F, Class I, and 10-kg Calibration Weights
Tolerances:
| Class |  | Readability and Sensitivity |  |  |  | Accuracy
| G1 |  |  | 0.01 g |  | 0.02 g or 0.1 percent
| G2 |  |  | 0.1 g |  | 0.2 or 0.1 percent
| G5 |  |  | 1 g |  | 2 g or 0.1 percent
| G20 |  |  | 5 g |  | 5 g or 0.1 percent
| G100 |  |  | 20 g |  | 20 g or 0.1 percent
| Weight Applied (g) |  | Scale Reading (g) |  | Tolerance (g) |  |  | Pass/Fail
| 1 |  | 1 |  | 2
| 5 |  | 5 |  | 2
| 10 |  | 10 |  | 2
| 50 |  | 50 |  | 2
| 100 |  | 100 |  | 2
| 500 |  | 500 |  | 2
| 501 |  | 501 |  | 2
| 1000 |  | 1000 |  | 2
| 5000 |  | 5000 |  | 5
| 5001 |  | 5001 |  | 5
| 10000 |  | 10000 |  | 10
| 10001 |  | 10001 |  | 10
| 20000 |  | 20001 |  | 20
| 30000 |  | 30002 |  | 30
| 30001 |  | 30003 |  | 30
[9 empty] | Pass/Fail: | Pass
Comments:
Initial By:

## Scale (G5-MX3)
Scale Calibration Worksheet
| Equipment ID: | MX - 3 |  |  |  |  | Date: | 2015-09-08
| Manufacturer: | Mettler |  |  |  |  | Performed by: | Danielle N.
| Model #: | MS 32000L / 01
| Serial #: | B508635397 |  |  |  |  | Next Calibration Due: | September 2016
| UofA ID: | NA |  |  |  |  | Last Calibration: | September 2014
| Location: | MX (234) (Left)
| Calibration Item: | Verify mass readings for at least 5 points in the range of the balance
Check sensitivity (0.1%) for at least 5 points in the range of the balance
| Calibration Procedure: | AASHTO M-231
| Calibration Equipment: | Class F, Class I, and 10-kg Calibration Weights
Tolerances:
| Class |  | Readability and Sensitivity |  |  |  | Accuracy
| G1 |  |  | 0.01 g |  | 0.02 g or 0.1 percent
| G2 |  |  | 0.1 g |  | 0.2 or 0.1 percent
| G5 |  |  | 1 g |  | 2 g or 0.1 percent
| G20 |  |  | 5 g |  | 5 g or 0.1 percent
| G100 |  |  | 20 g |  | 20 g or 0.1 percent
| Weight Applied (g) |  | Scale Reading (g) |  | Tolerance (g) |  |  | Pass/Fail
| 1 |  | 1 |  | 2
| 5 |  | 5 |  | 2
| 10 |  | 10 |  | 2
| 50 |  | 50 |  | 2
| 100 |  | 100 |  | 2
| 500 |  | 500 |  | 2
| 501 |  | 501 |  | 2
| 1000 |  | 1000 |  | 2
| 5000 |  | 5000 |  | 5
| 5001 |  | 5001 |  | 5
| 10000 |  | 10001 |  | 10
| 10001 |  | 10002 |  | 10
| 20000 |  | 20002 |  | 20
| 30000 |  | 30004 |  | 30
| 30001 |  | 30005 |  | 30
[9 empty] | Pass/Fail: | Pass
Comments:
Initial By:

## Scale (G100-C1)
Scale Calibration Worksheet
| Equipment ID: | C - 1 |  |  |  |  | Date: | 2015-09-02
| Manufacturer: | AND |  |  |  |  | Performed by: | M. Fleck
| Model #: | HV60KGL
| Serial #: | EM7306611 |  |  |  |  | Next Calibration Due: | September 2016
| UofA ID: | NA |  |  |  |  | Last Calibration: | September 2014
| Location: | CL (260)
| Calibration Item: | Verify mass readings for at least 5 points in the range of the balance
Check sensitivity (0.1%) for at least 5 points in the range of the balance
| Calibration Procedure: | AASHTO M-231
| Calibration Equipment: | Class F, Class I, and 10-kg Calibration Weights
Tolerances:
| Class |  | Readability and Sensitivity |  |  |  | Accuracy
| G1 |  |  | 0.01 g |  | 0.02 g or 0.1 percent
| G2 |  |  | 0.1 g |  | 0.2 or 0.1 percent
| G5 |  |  | 1 g |  | 2 g or 0.1 percent
| G20 |  |  | 5 g |  | 5 g or 0.1 percent
| G100 |  |  | 20 g |  | 20 g or 0.1 percent
| Weight Applied (g) |  | Scale Reading (g) |  | Tolerance (g) |  |  | Pass/Fail
| 20 |  | 20 |  | 20
| 500 |  | 500 |  | 20
| 520 |  | 520 |  | 20
| 1000 |  | 1000 |  | 20
| 5000 |  | 4995 |  | 20
| 10000 |  | 9995 |  | 20
| 10020 |  | 10015 |  | 20
| 15000 |  | 14995 |  | 20
| 30000 |  | 29995 |  | 30
| 30020 |  | 30015 |  | 30
| 45000 |  | 44990 |  | 45
| 45020 |  | 40010 |  | 45
| 60000 |  | 59995 |  | 60
[9 empty] | Pass/Fail: | Pass
| Comments: | Scales in spec but recalibrated to make better.
Initial By:

## Scale (G100-C2)
Scale Calibration Worksheet
| Equipment ID: | C - 2 |  |  |  |  | Date: | 2015-09-02
| Manufacturer: | AND |  |  |  |  | Performed by: | M. Fleck
| Model #: | FG60KAM
| Serial #: | Q1911021 |  |  |  |  | Next Calibration Due: | September 2016
| UofA ID: | NA |  |  |  |  | Last Calibration: | September 2014
| Location: | STO (216)
| Calibration Item: | Verify mass readings for at least 5 points in the range of the balance
Check sensitivity (0.1%) for at least 5 points in the range of the balance
| Calibration Procedure: | AASHTO M-231
| Calibration Equipment: | Class F, Class I, and 10-kg Calibration Weights
Tolerances:
| Class |  | Readability and Sensitivity |  |  |  | Accuracy
| G1 |  |  | 0.01 g |  | 0.02 g or 0.1 percent
| G2 |  |  | 0.1 g |  | 0.2 or 0.1 percent
| G5 |  |  | 1 g |  | 2 g or 0.1 percent
| G20 |  |  | 5 g |  | 5 g or 0.1 percent
| G100 |  |  | 20 g |  | 20 g or 0.1 percent
| Weight Applied (g) |  | Scale Reading (g) |  | Tolerance (g) |  |  | Pass/Fail
| 20 |  | 20 |  | 20
| 500 |  | 500 |  | 20
| 520 |  | 520 |  | 20
| 1000 |  | 1000 |  | 20
| 5000 |  | 5000 |  | 20
| 10000 |  | 10000 |  | 20
| 10020 |  | 10020 |  | 20
| 15000 |  | 15000 |  | 20
| 30000 |  | 30000 |  | 30
| 30020 |  | 30020 |  | 30
| 45000 |  | 45000 |  | 45
| 45020 |  | 45020 |  | 45
| 60000 |  | 60000 |  | 60
[9 empty] | Pass/Fail: | Pass
| Comments: | Scales in spec but recalibrated to be better.
Initial By: