## Capping Compound
Capping Compound Calibration Worksheet
|  | Equipment ID: | Ingot Capping Compound |  |  | Date Received: | 2007-04-15
|  | Manufacturer: | Gilson |  |  | Received By: | F. Griffith
|  | Model #: | Red Iron 9000 HM - 189 |  |  | Amount in Order: | 3 Bags
|  | Serial #: | NA
|  | UofA ID: | NA |  |  | Calibration Interval: | 3 months
|  | Location: | RM (225) Block Room |  |  |  | or before use
| Calibration Item: | Verify compressive strength of mortar
| Calibration Procedure: | ASTM C 617 - section 5
| Calibration Equipment: | cube mold, melting pot, capping compound, compression testing machine
Compressive Strength (psi)
Date | Cube 1 | Cube 2 | Cube 3 | Anticipated Cylinder Strength (psi) |  |  | Pass/Fail | Initial
2007-05-14 | 7553 | 7290 | 7410 |  | 5000 |  | Pass | sgw
2007-05-22 | 7238 | 7567 | 7029 |  | 5000 |  | Pass | sgw
2007-08-31 | 11313 | 11953 | 12368 |  | 5000 |  | Pass | sgw
2007-11-08 | 7848 | 8033 | 7928 |  | 5000 |  | Pass | sgw
2008-02-05 | 8479 | 7925 | 8107 |  | 5000 |  | Pass | ftg
2008-05-14 | 7547 | 7263 | 7039 |  | 5000 |  | Pass | ftg
2008-08-06 | 9011 | 8897 | 9129 |  | 5000 |  | Pass | ftg
2008-12-03 | 7235 | 7356 | 7702 |  | 5000 |  | Pass | ftg
2009-02-18 | 9143 | 9325 | 9079 |  | 5000 |  | Pass | ftg
2009-05-13 | 7970 | 8615 | 8047 |  | 5000 |  | Pass | ftg
2009-08-19 | 12140 | 12325 | 11845 |  | 5000 |  | Pass | ftg
2009-11-30 | 8963 | 9133 | 9172 |  | 5000 |  | Pass | ftg
Comments:

## 2" Cubes
2" Cube Test Set
Equipment ID: | Water Storage Tank | 2" cube test set |  |  | Date: | 2009-09-03
Manufacturer: |  | Gilson |  |  | Performed by: | F. Griffith
Model #: |  | MCA-6
Serial #: |  |  |  |  | Calibration Interval: | Yearly
UofA ID: | [5 empty] | New
Location: | Concrete Lab Moist Room | ST (218)
Maintenance Item | Calibration Item: | Head and faces of set
| Calibration Procedure: | Lubricate and check for gouges
| Calibration Equipment: | Steel bar and feeler gauges and oil
Date |  | Plane to .002 in 6"? | Lubricate head? | Performed by:
2009-09-03 |  | Yes | Yes | F. Griffith

## Capping Plates
Capping Plates Calibration Worksheet
|  | Equipment ID: | 6" Blue Capping Plate |  | Date: | May
|  | Manufacturer: | Gilson |  | Performed by: | Various
|  | Model #: | HM - 166
|  | Serial #: | NA
|  | UofA ID: | NA |  | Calibration Interval: | 12 months
|  | Location: | ST (262)
| Calibration Item: | Verify plate dimensions
| Calibration Procedure: | ASTM C 617 - 4.1
| Calibration Equipment: | digital calipers
Date | Thickness (≥ 1/4") | Recessed Depth (≤ 1/2") | Plane within 0.002 in.? | Comments: | Performed by: |  | Pass Fail
2007-05-14 | 0.6090" | 0.4375" | Yes |  | S. G. Williams |  | Pass
2008-05-12 | 0.6110" | 0.4275" | Yes |  | F. Griffith |  | Pass
2009-05-12 | 0.6140" | 0.432 | Yes |  | F. Griffith |  | Pass
[6 empty] | Initial by: | sgw

## Alignment Device
Alignment Device Calibration Worksheet
|  | Equipment ID: | Alignment Devices |  | Date: | 2009-05-12
|  | Manufacturer: | Test Mark |  | Performed by: | F. Griffith
|  | Model #: | CA - 0110
|  | Serial #: | NA |  | Next Calibration Due: | 2010-05-01
|  | UofA ID: | NA |  | Last Calibration: | 2008-05-01
|  | Location: | ST (218)
| Calibration Item: | Verify the device dimensions
| Calibration Procedure: | ASTM C 617 - 4.1
| Calibration Equipment: | digital calipers, straightedge, feeler gauges
Verify the device dimensions:
|  | Thickness of plate: | 0.6300" | Tolerance = at least 1/4"
|  | Depth of recessed area: | 0.4595" | Tolerance = no more than 1/2"
|  | Diameter of plate: | 8.026" | Tolerance = at least 1" greater than
specimen diameter
|  | Is surface plane within 0.002 inches? | Yes
Are alignment guides perpendicular
|  | within 0.5 degrees (1/8" in 12")? | Yes
[6 empty] | Pass/Fail: | Pass
Comments:
[7 empty] | Initial By: | ftg

## Retainer Rings
Calibration Worksheet for Retainer Rings
|  | Equipment ID: | 6" Retainer Rings |  |  | Date: | May
|  | Manufacturer: | Test Mark |  |  | Performed by: | Various
Model #:
|  | Serial #: | NA
|  | UofA ID: | NA |  |  | Calibration Interval: | 12 mos.
|  | Location: | ST (218)
| Calibration Item: | Verify retainer ring dimensions
| Calibration Procedure: | ASTM C 1231 - 5.3
| Calibration Equipment: | digital calipers
| Depth of recession (in) ≥ 2* pad thickness |  | Inside Diameter of Ring 102 - 107% of cylinder dia.
Date | Ring A | Ring B | Ring A | Ring B | Any gouges > 0.01" or > 0.05 in.2? |  | Pass/Fail | Initial
2007-05-14 | 1 | 1 | 6.074 | 6.106 | no |  | Pass | sgw
2008-05-12 | 1 | 1 | 6.125 | 6.123 | no |  | Pass | ftg
2009-05-12 | 1.0645 | 1.0527 | 6.1215 | 6.1245 | no |  | Pass | ftg
| Comments: | [5 empty] | Initial by: | sgw

## Neoprene Pads
Calibration Worksheet for Elastomeric Pads
|  | Equipment ID: | Elastomeric Pad - # 70 |  |  |  | Date: | 2009-05-12
|  | Manufacturer: | Engineered Plastics |  |  |  | Performed by: | F. Griffith
|  | Model #: | CA - 0190 - 01
|  | Serial #: | NA
|  | UofA ID: | NA |  |  |  | Calibration Interval: | prior to first use
|  | Location: | ST (218)
| Calibration Item: | Verify pad dimensions
| Calibration Procedure: | ASTM C 1231 - 5.2
| Calibration Equipment: | digital calipers
| Pad Thickness (in) |  | Pad Diameter (in)
Date | Pad 01 | Pad 02 | Pad 01 | Pad 02 |  | Any cracks > 3/8"? |  | Pass/Fail | Initial
2007-05-15 | 0.4845" | 0.4980" | 6.085" | 6.101" | No |  |  | Pass | ftg
2008-05-12 | 0.4930" | 0.5155" | 6.109" | 6.122" | No |  |  | Pass | ftg
2009-05-12 | 0.4950" | 0.4900" | 6.0905" | 6.0985" | No |  |  | Pass | ftg
| Comments: | [6 empty] | Initial By: | sgw

## Pad Usage Log
Pad Usage Log Sheet
|  | Pad ID: | 01 - 09 |  |  | Durometer: | 70 |  |  |  | Date Entered into Service: | 2009-02-17
|  | Initial | Date |  |  | Initial | Date |  | Initial | Date |  | Initial | Date
X | 1 | FG | 2009-02-01 | X | 26 | FG |  | 51 |  |  | 76
X | 2 | FG |  | X | 27 | FG | 2009-09-01 | 52 |  |  | 77
X | 3 | FG |  | X | 28 | FG |  | 53 |  |  | 78
X | 4 | FG |  |  | 29 |  |  | 54 |  |  | 79
X | 5 | FG |  |  | 30 |  |  | 55 |  |  | 80
X | 6 | FG |  |  | 31 |  |  | 56 |  |  | 81
X | 7 | FG |  |  | 32 |  |  | 57 |  |  | 82
X | 8 | FG |  |  | 33 |  |  | 58 |  |  | 83
X | 9 | FG | 2009-08-01 |  | 34 |  |  | 59 |  |  | 84
X | 10 | FG |  |  | 35 |  |  | 60 |  |  | 85
X | 11 | FG |  |  | 36 |  |  | 61 |  |  | 86
X | 12 | FG |  |  | 37 |  |  | 62 |  |  | 87
X | 13 | FG |  |  | 38 |  |  | 63 |  |  | 88
X | 14 | FG |  |  | 39 |  |  | 64 |  |  | 89
X | 15 | FG |  |  | 40 |  |  | 65 |  |  | 90
X | 16 | FG |  |  | 41 |  |  | 66 |  |  | 91
X | 17 | FG |  |  | 42 |  |  | 67 |  |  | 92
X | 18 | FG |  |  | 43 |  |  | 68 |  |  | 93
X | 19 | FG |  |  | 44 |  |  | 69 |  |  | 94
X | 20 | FG |  |  | 45 |  |  | 70 |  |  | 95
X | 21 | FG |  |  | 46 |  |  | 71 |  |  | 96
X | 22 | FG |  |  | 47 |  |  | 72 |  |  | 97
X | 23 | FG |  |  | 48 |  |  | 73 |  |  | 98
X | 24 | FG |  |  | 49 |  |  | 74 |  |  | 99
X | 25 | FG |  |  | 50 |  |  | 75 |  |  | 100
| Notes: | [9 empty] | Last Date of Service:

## Pad Usage Log (2)
Pad Usage Log Sheet
|  | Pad ID: | 02 - 09 |  |  | Durometer: | 70 |  |  |  | Date Entered into Service: | 2009-02-17
|  | Initial | Date |  |  | Initial | Date |  | Initial | Date |  | Initial | Date
X | 1 | FG | 2009-02-01 | X | 26 | FG |  | 51 |  |  | 76
X | 2 | FG |  | X | 27 | FG | 2009-09-01 | 52 |  |  | 77
X | 3 | FG |  | X | 28 | FG |  | 53 |  |  | 78
X | 4 | FG |  |  | 29 |  |  | 54 |  |  | 79
X | 5 | FG |  |  | 30 |  |  | 55 |  |  | 80
X | 6 | FG |  |  | 31 |  |  | 56 |  |  | 81
X | 7 | FG |  |  | 32 |  |  | 57 |  |  | 82
X | 8 | FG |  |  | 33 |  |  | 58 |  |  | 83
X | 9 | FG | 2009-08-01 |  | 34 |  |  | 59 |  |  | 84
X | 10 | FG |  |  | 35 |  |  | 60 |  |  | 85
X | 11 | FG |  |  | 36 |  |  | 61 |  |  | 86
X | 12 | FG |  |  | 37 |  |  | 62 |  |  | 87
X | 13 | FG |  |  | 38 |  |  | 63 |  |  | 88
X | 14 | FG |  |  | 39 |  |  | 64 |  |  | 89
X | 15 | FG |  |  | 40 |  |  | 65 |  |  | 90
X | 16 | FG |  |  | 41 |  |  | 66 |  |  | 91
X | 17 | FG |  |  | 42 |  |  | 67 |  |  | 92
X | 18 | FG |  |  | 43 |  |  | 68 |  |  | 93
X | 19 | FG |  |  | 44 |  |  | 69 |  |  | 94
X | 20 | FG |  |  | 45 |  |  | 70 |  |  | 95
X | 21 | FG |  |  | 46 |  |  | 71 |  |  | 96
X | 22 | FG |  |  | 47 |  |  | 72 |  |  | 97
X | 23 | FG |  |  | 48 |  |  | 73 |  |  | 98
X | 24 | FG |  |  | 49 |  |  | 74 |  |  | 99
X | 25 | FG |  |  | 50 |  |  | 75 |  |  | 100
| Notes: | [9 empty] | Last Date of Service:

## Water Tank Maintenance
Water Storage Tank Maintenance Worksheet
Equipment ID: |  | Water Storage Tank |  |  | Date:
Manufacturer: |  |  |  |  | Performed by: | Various
Model #:
Serial #: |  |  |  |  | Next Calibration Due:
UofA ID: |  |  |  |  | Last Calibration:
Location: |  | Concrete Moist Room
Maintenance Item: | Calibration Item: | Stir water in tank to replace calcium ions
Maintenance Interval: | Calibration Procedure: | Once per month
Maintenance Item: | Calibration Equipment: | Change water and saturate with sodium hydroxide
Maintenance Interval: |  | Every 24 months
Date: |  | Water Stirred: | Water Changed: | Comments: | Performed by:
2007 - 2008 |  |  |  | OOS | F. Griffith
3/2009 |  | Yes |  | In Service | F. Griffith
4/2009 |  | Yes |  |  | F. Griffith
5/2009 |  | Yes | Yes |  | F. Griffith
6/2009 |  | Yes |  |  | F. Griffith
7/2009 |  | Yes |  |  | F. Griffith
8/2009 |  | Yes |  |  | F. Griffith
9/2009 |  | Yes |  |  | F. Griffith
10/2009 |  | Yes |  |  | F. Griffith
11/2009 |  | Yes |  |  | F. Griffith
12/2009 |  |  |  | OOS | F. Griffith
[6 empty] | Initial by: | sgw

## Water Tank Temperature
Water Storage Tank Temperature Worksheet
Equipment ID: |  | Water Storage Tank |  | Date: |  | Weekly
Manufacturer: |  | County Line |  | Performed by: |  | Various
Model #: |  | 100 gal (2' x 4' x 2')
Serial #: |  | NA
UofA ID: |  | NA
Location: |  | Concrete Moist Room
| Calibration Item: | Maintenance of 23.0 ± 2 °C (73.5 ± 3.5° F)
Maintenance Item: | Calibration Procedure: | Verify that correct temperature is being maintained in tank
Interval: | Calibration Equipment: | Weekly when in use
(Allowable Range 70 - 77 °F)
Date: |  | Wheel Temperature: | Comments: | Verified by: |  |  | Pass / Fail
2009-03-23 |  | 69 °F | NIST temp 73°F - adj. recorder | F. Griffith |  |  | Pass
2009-03-30 |  | 73 °F |  | F. Griffith |  |  | Pass
2009-04-06 |  | 74 °F |  | F. Griffith |  |  | Pass
2009-04-13 |  | 73 °F |  | F. Griffith |  |  | Pass
2009-04-20 |  | 73 °F |  | F. Griffith |  |  | Pass
2009-04-27 |  | 74 °F |  | F. Griffith |  |  | Pass
2009-05-04 |  | 73 °F |  | F. Griffith |  |  | Pass
2009-05-12 |  | 72 °F | Water changed | F. Griffith |  |  | Pass
2009-05-18 |  | 72 °F |  | F. Griffith |  |  | Pass
2009-05-26 |  | 72 °F |  | F. Griffith |  |  | Pass
2009-06-01 |  | 72 °F |  | F. Griffith |  |  | Pass
2009-06-08 |  | 73 °F |  | R. Conley |  |  | Pass
2009-06-15 |  | 73 °F |  | F. Griffith |  |  | Pass
2009-06-22 |  | 73 °F |  | R. Conley |  |  | Pass
2009-06-29 |  | 73 °F |  | F. Griffith |  |  | Pass
2009-07-09 |  | 73 °F |  | F. Griffith |  |  | Pass
2009-07-13 |  | 74 °F |  | F. Griffith |  |  | Pass
2009-07-20 |  | 73 °F |  | F. Griffith |  |  | Pass
2009-07-27 |  | 73 °F |  | F. Griffith |  |  | Pass
2009-08-03 |  | 74 °F |  | F. Griffith |  |  | Pass
2009-08-10 |  | 73 °F |  | F. Griffith |  |  | Pass
2009-08-17 |  | 73 °F |  | F. Griffith |  |  | Pass
2009-08-26 |  | 74 °F |  | F. Griffith |  |  | Pass
2009-09-01 |  | 75 °F | Water added | F. Griffith |  |  | Pass
2009-09-03 |  | 72 °F |  | F. Griffith |  |  | Pass
2009-09-07 |  | 72 °F |  | F. Griffith |  |  | Pass
Wheel saved and temperature recorded verified by calibrated thermometer.
[5 empty] | Initial by: |  | sgw

## Temperature Recorder
Water Tank Temperature Recorder Calibration Worksheet
Equipment ID: | Water Storage Tank | Temp. Recorder |  |  | Date: | 2009
Manufacturer: |  | Palmer Instruments |  |  | Performed by: | Various
Model #: |  | X4-16133 |  |  |  | 6 Month Interval
Serial #: |  | NA |  |  | Next Calibration Due: | May / Nov.
UofA ID: |  | NA |  |  | Last Calibration: | May / Nov.
Location: | Concrete Moist Room
Calibration: |  | Automatic control of 23.0 + 2 °C
C. Procedure: | Calibration Procedure: | Verify that temperature recorder is correct (± 1° C = ± 1.8° F)
Equipment: |  | Nist Traceable Thermometer (T-Cal 2)
Date: |  | Recorder: | NIST Thermometer: | Verified by: | Pass / Fail:
2007 |  | OOS |  | F. Griffith
2008 |  | OOS |  | F. Griffith
2009-03-23 |  | 69° F | 73.0° F | F. Griffith | Adjustment made.
2009-05-05 |  | 73° F | 72.6° F | F. Griffith | Pass
2009-09-03 |  | 72° F | 72.0° F | F. Griffith | Pass
2009-09-15 |  | 72° F | 73.0° F | F. Griffith | Pass
[6 empty] | Initial by: | sgw