## Slump Cone
Slump Cone
|  | Equipment ID: | Slump Cone |  | Date: | 2010-05-18
|  | Manufacturer: | Gilson |  | Performed by: | DAP
|  | Model #: | HM - 45
|  | Serial #: | NA |  | Next Calibration Due: | May 2011
|  | UofA ID: | NA |  | Last Calibration: | May 2009
|  | Location: | ST (218)
| Calibration Item: | Verify critical dimensions of the cone
| Calibration Procedure: | In-House Procedure #07-1
| Calibration Equipment: | Mititoyo Digimatic Calipers, and ruler read. to 1 mm.
|  | Thickness of cone: | 0.061" | Tolerance = Not less than 0.060 inches
|  | Bottom Inside Diameter of Cone: | 7.963" | Tolerance = 8 +/- 1/8 inches
|  | Top Inside Diameter of Cone: | 4.009" | Tolerance = 4 +/- 1/8 inches
|  | Height of Cone: | 12.175" | Tolerance = 12 +/- 1/8 inches
|  | Cone Interior Smooth: | Yes | Pass/Fail
|  | Base and top opening parallel: | Yes | Pass/Fail
[6 empty] | Pass/Fail: | Pass
Comments:
[7 empty] | Initial By: | DAP

## Cylinder Molds
Single-Use Cylinder Mold
|  | Equipment ID: | Cylinder Molds |  |  |  | Date: | 2010-05-27
|  | Manufacturer: | Test Mark |  |  |  | Performed by: | F. Griffith
|  | Model #: | MA Industries
Serial #:
|  | UofA ID: |  |  |  |  | Next Calibration Due: | May 2011
|  | Location: |  |  |  |  | Last Calibration: | Sept. 2009
| Calibration Item: | Verify critical dimensions of the cylinder
| Calibration Procedure: | ASTM C 31 - C 470
Resistance to damage under use and water leakage
| Calibration Equipment: | Mititoyo Digimatic Calipers, and ruler read. to 1 mm.
Date | Length | Diameter | L/D ratio (approx. 2) |  | Comments |  |  |  | Pass/Fail | Initial
2010-05-27 | 12.05" | 6.10" | 2 | Brand: |  | MA Industries
|  |  |  | Shipment: |  | 2008-07-01
|  |  |  | Water Leakage: |  |  |  |  | Pass | ftg
|  |  |  | Resistance to Damage: |  |  |  |  | Pass | ftg
2010-05-27 | 12.10" | 6.05" | 2 | Brand: |  | MA Industries
|  |  |  | Shipment: |  | 2008-07-01
|  |  |  | Water Leakage: |  |  |  |  | Pass | ftg
|  |  |  | Resistance to Damage: |  |  |  |  | Pass | ftg
2010-05-27 | 12.08" | 6.07" | 2 | Brand: |  | MA Industries
|  |  |  | Shipment: |  | 2008-07-01
|  |  |  | Water Leakage: |  |  |  |  | Pass | ftg
|  |  |  | Resistance to Damage: |  |  |  |  | Pass | ftg
Brand:
Shipment:
Water Leakage:
Resistance to Damage:
Brand:
Shipment:
Water Leakage:
Resistance to Damage:
| Comments: | [7 empty] | Initial by: | sgw