## Capping Compound
Capping Compound Calibration Worksheet
|  | Equipment ID: | Ingot Capping Compound |  |  | Date Received: | 2007-04-15
|  | Manufacturer: | Gilson |  |  | Received By: | F. Griffith
|  | Model #: | Red Iron 9000 HM - 189 |  |  | Amount in Order: | 3 Bags
|  | Serial #: | NA
|  | UofA ID: | NA |  |  | Calibration Interval: | 3 months
|  | Location: | RM (225) Block Room |  |  |  | or before use
| Calibration Item: | Verify compressive strength of mortar
| Calibration Procedure: | ASTM C 617 - section 5
| Calibration Equipment: | cube mold, melting pot, capping compound, compression testing machine
Compressive Strength (psi)
Date | Cube 1 | Cube 2 | Cube 3 | Anticipated Cylinder Strength (psi) |  |  | Pass/Fail | Initial
2007-05-14 | 7553 | 7290 | 7410 |  | 5000 |  | Pass | sgw
2007-05-22 | 7238 | 7567 | 7029 |  | 5000 |  | Pass | sgw
2007-08-31 | 11313 | 11953 | 12368 |  | 5000 |  | Pass | sgw
2007-11-08 | 7848 | 8033 | 7928 |  | 5000 |  | Pass | sgw
2008-02-05 | 8479 | 7925 | 8107 |  | 5000 |  | Pass | ftg
2008-05-14 | 7547 | 7263 | 7039 |  | 5000 |  | Pass | ftg
2008-08-06 | 9011 | 8897 | 9129 |  | 5000 |  | Pass | ftg
2008-12-03 | 7235 | 7356 | 7702 |  | 5000 |  | Pass | ftg
2009-02-18 | 9143 | 9325 | 9079 |  | 5000 |  | Pass | ftg
2009-05-13 | 7970 | 8615 | 8047 |  | 5000 |  | Pass | ftg
2009-08-19 | 12140 | 12325 | 11845 |  | 5000 |  | Pass | ftg
2009-11-30 | 8963 | 9133 | 9172 |  | 5000 |  | Pass | ftg
2010-02-23 | 8274 | 7853 | 8401 |  | 5000 |  | Pass | ftg
2010-04-27 | 7738 | 8031 | 7982 |  | 5000 |  | Pass | ftg
2010-07-30 | 10272 | 9364 | 9558 |  | 5000 |  | Pass | ftg
2010-10-19 | 6651 | 6753 | 6709 |  | 5000 |  | Pass | ftg
2011-02-24 | 6317 | 6385 | 6403 |  | 5000 |  | Pass | ftg
2011-05-16 | 6387 | 6241 | 6348 |  | 5000 |  | Pass | ftg
2011-08-22 | Not in use this quarter |  |  |  | Next Class ; Nov., 2011 |  |  | mmf
2011-11-15 | 6339 | 6407 | 6451 |  | 5000 |  | Pass | ftg
Comments:

## 2" Cubes
2" Cube Test Set
| Equipment ID: | 2" cube test set |  |  | Date: | 2011-05-16
| Manufacturer: | Gilson |  |  | Performed by: | F. Griffith
| Model #: | MCA-6
| Serial #: |  |  | Calibration Interval: |  | Yearly
| UofA ID: |  |  | Next Calibration Due: |  | 2012-05-01
| Location: | ST (218) |  | Last Calibration: |  | 2010-05-01
| Maintenance Item: | Check tightness and lubricate
| Calibration Procedure: | Verify planeness and critical dimensions
| Calibration Equipment: | Caliper, steel bar, feeler gauges and oil
Cube | Mold parts | Sides Plane? | Distance b/t sides | Height of Sides |  | Angle 90 ± 0.5°
| fit tightly? | < 0.002" | 2" ± 0.02" | 2" (+0.01 to -0.015) |  | between faces?
1 | Yes | Yes | 2.01 |  |  | Yes
Yes
|  | Yes | 2.09
Yes
2 | Yes | Yes | 2.075 |  |  | Yes
Yes
|  | Yes | 2.01
Yes
3 | Yes | Yes | 2.095 |  |  | Yes
Yes
Yes
Yes
Initial:
Notes : | 0.5° is equivalent to 1/8" over a distance of 12" = 0.02" in 2"

## Capping Plates
Capping Plates Calibration Worksheet
| Equipment ID: | 6" Blue Capping Plate |  |  | Date: | May
| Manufacturer: | Gilson |  |  | Performed by: | Various
| Model #: | HM - 166
| Serial #: | NA
| UofA ID: | NA |  |  | Calibration Interval: | 12 months
| Location: | ST (262)
| Calibration Item: | Verify plate dimensions
| Calibration Procedure: | ASTM C 617 - 4.1
| Calibration Equipment: | digital calipers
Date | Thickness (≥ 1/4") | Recessed Depth (≤ 1/2") | Plane within 0.002 in.? | Comments: | Performed by: |  | Pass Fail
2007-05-14 | 0.6090" | 0.4375" | Yes |  | S. G. Williams |  | Pass
2008-05-12 | 0.6110" | 0.4275" | Yes |  | F. Griffith |  | Pass
2009-05-12 | 0.6140" | 0.432 | Yes |  | F. Griffith |  | Pass
2010-05-18 | 0.6109" | 0.4327 | Yes |  | F. Griffith |  | Pass
2011-05-16 | 0.6180" | 0.375 | Yes |  | F. Griffith |  | Pass
[6 empty] | Initial by: | sgw

## Alignment Device
Alignment Device Calibration Worksheet
|  | Equipment ID: | Alignment Devices |  | Date: | 2011-05-01
|  | Manufacturer: | Test Mark |  | Performed by: | F. Griffith
|  | Model #: | CA - 0110
|  | Serial #: | NA |  | Next Calibration Due: | 2012-05-01
|  | UofA ID: | NA |  | Last Calibration: | 2010-05-01
|  | Location: | ST (218)
| Calibration Item: | Verify the device dimensions
| Calibration Procedure: | ASTM C 617 - 4.1
| Calibration Equipment: | digital calipers, straightedge, feeler gauges
Verify the device dimensions:
|  | Thickness of plate: | 0.618" | Tolerance = at least 1/4"
|  | Depth of recessed area: | 0.375" | Tolerance = no more than 1/2"
|  | Diameter of plate: | 8.02" | Tolerance = at least 1" greater than
specimen diameter
|  | Is surface plane within 0.002 inches? | Yes
Are alignment guides perpendicular
|  | within 0.5 degrees (1/8" in 12")? | Yes
[6 empty] | Pass/Fail: | Pass
Comments:
[7 empty] | Initial By: | ftg

## Retainer Rings
Calibration Worksheet for Retainer Rings
|  | Equipment ID: | 6" Retainer Rings |  |  | Date: | May
|  | Manufacturer: | Test Mark |  |  | Performed by: | Various
Model #:
|  | Serial #: | NA
|  | UofA ID: | NA |  |  | Calibration Interval: | 12 mos.
|  | Location: | ST (218)
| Calibration Item: | Verify retainer ring dimensions
| Calibration Procedure: | ASTM C 1231 - 5.3
| Calibration Equipment: | digital calipers
| Depth of recession (in) ≥ 2* pad thickness |  | Inside Diameter of Ring 102 - 107% of cylinder dia.
Date | Ring A | Ring B | Ring A | Ring B | Any gouges > 0.01" or > 0.05 in.2? |  | Pass/Fail | Initial
2007-05-14 | 1 | 1 | 6.074 | 6.106 | no |  | Pass | sgw
2008-05-12 | 1 | 1 | 6.125 | 6.123 | no |  | Pass | ftg
2009-05-12 | 1.0645 | 1.0527 | 6.1215 | 6.1245 | no |  | Pass | ftg
2010-05-18 | 1.0538 | 1.0529 | 6.122 | 6.124 | no |  | Pass | ftg
2011-05-16 | 0.938 | 0.938 | 6.12 | 6.12 | no |  | Pass | ftg
| Comments: | [5 empty] | Initial by: | sgw

## Neoprene Pads
Calibration Worksheet for Elastomeric Pads
|  | Equipment ID: | Elastomeric Pad - # 70 |  |  |  | Date: | 2011-05-14
|  | Manufacturer: | Engineered Plastics |  |  |  | Performed by: | F. Griffith
|  | Model #: | CA - 0190 - 01
|  | Serial #: | NA
|  | UofA ID: | NA |  |  |  | Calibration Interval: | prior to first use
|  | Location: | ST (218)
| Calibration Item: | Verify pad dimensions
| Calibration Procedure: | ASTM C 1231 - 5.2
| Calibration Equipment: | digital calipers
| Pad Thickness (in) |  | Pad Diameter (in)
Date | Pad 01 | Pad 02 | Pad 01 | Pad 02 |  | Any cracks > 3/8"? |  | Pass/Fail | Initial
2007-05-15 | 0.4845" | 0.4980" | 6.085" | 6.101" | No |  |  | Pass | ftg
2008-05-12 | 0.4930" | 0.5155" | 6.109" | 6.122" | No |  |  | Pass | ftg
2009-05-12 | 0.4950" | 0.4900" | 6.0905" | 6.0985" | No |  |  | Pass | ftg
2010-05-18 | 0.5105" | 0.5090" | 6.109" | 6.087" | No |  |  | Pass | ftg
2011-05-14 | 0.4945" | 0.4910" | 6.050" | 6.060" | No |  |  | Pass | ftg
| Comments: | [6 empty] | Initial By: | sgw

## Pad Usage Log
Pad Usage Log Sheet
| Pad ID: |  |  | Durometer: | 70 |  |  |  | Date Entered into Service:
| Initial | Date |  | Initial | Date |  | Initial | Date |  | Initial | Date
1 |  |  | 26 |  |  | 51 |  |  | 76
2 |  |  | 27 |  |  | 52 |  |  | 77
3 |  |  | 28 |  |  | 53 |  |  | 78
4 |  |  | 29 |  |  | 54 |  |  | 79
5 |  |  | 30 |  |  | 55 |  |  | 80
6 |  |  | 31 |  |  | 56 |  |  | 81
7 |  |  | 32 |  |  | 57 |  |  | 82
8 |  |  | 33 |  |  | 58 |  |  | 83
9 |  |  | 34 |  |  | 59 |  |  | 84
10 |  |  | 35 |  |  | 60 |  |  | 85
11 |  |  | 36 |  |  | 61 |  |  | 86
12 |  |  | 37 |  |  | 62 |  |  | 87
13 |  |  | 38 |  |  | 63 |  |  | 88
14 |  |  | 39 |  |  | 64 |  |  | 89
15 |  |  | 40 |  |  | 65 |  |  | 90
16 |  |  | 41 |  |  | 66 |  |  | 91
17 |  |  | 42 |  |  | 67 |  |  | 92
18 |  |  | 43 |  |  | 68 |  |  | 93
19 |  |  | 44 |  |  | 69 |  |  | 94
20 |  |  | 45 |  |  | 70 |  |  | 95
21 |  |  | 46 |  |  | 71 |  |  | 96
22 |  |  | 47 |  |  | 72 |  |  | 97
23 |  |  | 48 |  |  | 73 |  |  | 98
24 |  |  | 49 |  |  | 74 |  |  | 99
25 |  |  | 50 |  |  | 75 |  |  | 100
Last Date of Service:

## Pad Usage Log (2)
Pad Usage Log Sheet
| Pad ID: |  |  | Durometer: | 70 |  |  |  | Date Entered into Service:
| Initial | Date |  | Initial | Date |  | Initial | Date |  | Initial | Date
1 |  |  | 26 |  |  | 51 |  |  | 76
2 |  |  | 27 |  |  | 52 |  |  | 77
3 |  |  | 28 |  |  | 53 |  |  | 78
4 |  |  | 29 |  |  | 54 |  |  | 79
5 |  |  | 30 |  |  | 55 |  |  | 80
6 |  |  | 31 |  |  | 56 |  |  | 81
7 |  |  | 32 |  |  | 57 |  |  | 82
8 |  |  | 33 |  |  | 58 |  |  | 83
9 |  |  | 34 |  |  | 59 |  |  | 84
10 |  |  | 35 |  |  | 60 |  |  | 85
11 |  |  | 36 |  |  | 61 |  |  | 86
12 |  |  | 37 |  |  | 62 |  |  | 87
13 |  |  | 38 |  |  | 63 |  |  | 88
14 |  |  | 39 |  |  | 64 |  |  | 89
15 |  |  | 40 |  |  | 65 |  |  | 90
16 |  |  | 41 |  |  | 66 |  |  | 91
17 |  |  | 42 |  |  | 67 |  |  | 92
18 |  |  | 43 |  |  | 68 |  |  | 93
19 |  |  | 44 |  |  | 69 |  |  | 94
20 |  |  | 45 |  |  | 70 |  |  | 95
21 |  |  | 46 |  |  | 71 |  |  | 96
22 |  |  | 47 |  |  | 72 |  |  | 97
23 |  |  | 48 |  |  | 73 |  |  | 98
24 |  |  | 49 |  |  | 74 |  |  | 99
25 |  |  | 50 |  |  | 75 |  |  | 100
Last Date of Service:

## Water Tank Maintenance
Water Storage Tank Maintenance Worksheet
Equipment ID: |  | Water Storage Tank |  |  | Date: | 2011
Manufacturer: |  |  |  |  | Performed by: | Various
Model #:
Serial #: |  |  |  |  | Next Calibration Due:
UofA ID: |  |  |  |  | Last Calibration:
Location: |  | Concrete Moist Room
Maintenance Item: | Calibration Item: | Stir water in tank to replace calcium ions
Maintenance Interval: | Calibration Procedure: | Once per month
Maintenance Item: | Calibration Equipment: | Change water and saturate with sodium hydroxide
Maintenance Interval: |  | Every 24 months
Date: |  | Water Stirred: | Water Changed: | Comments: | Performed by:
2007 - 2008 |  |  |  | OOS | F. Griffith
3/2009 |  | Yes |  | In Service | F. Griffith
4/2009 |  | Yes |  |  | F. Griffith
5/2009 |  | Yes | Yes |  | F. Griffith
6/2009 |  | Yes |  |  | F. Griffith
7/2009 |  | Yes |  |  | F. Griffith
8/2009 |  | Yes |  |  | F. Griffith
9/2009 |  | Yes |  |  | F. Griffith
10/2009 |  | Yes |  |  | F. Griffith
11/2009 |  | Yes |  |  | F. Griffith
12/2009 |  |  |  | OOS | F. Griffith
5/2010 |  | Yes | Yes | In Service | F. Griffith
6/2010 |  |  |  | OOS | F. Griffith
11/2010 |  | Yes |  | In Service | F. Griffith
12/2010 |  |  |  | OOS | F. Griffith
5/2011 |  | Yes |  | In Service | F. Griffith
6/2011 |  |  |  | OOS | F. Griffith
11/2011 |  | Yes |  | In Service | F. Griffith
12/2011 |  |  |  | OOS | F. Griffith
[6 empty] | Initial by: | sgw

## Temperature Recorder
Water Tank Temperature Recorder Calibration Worksheet
Equipment ID: | Water Storage Tank | Temp. Recorder |  |  | Date: | 2011
Manufacturer: |  | Palmer Instruments |  |  | Performed by: | Various
Model #: |  | X4-16133 |  |  |  | 6 Month Interval
Serial #: |  | NA |  |  | Next Calibration Due: | May / Nov.
UofA ID: |  | NA |  |  | Last Calibration: | May / Nov.
Location: | Concrete Moist Room
Calibration: |  | Automatic control of 23.0 + 2 °C
C. Procedure: | Calibration Procedure: | Verify that temperature recorder is correct (± 1° C = ± 1.8° F)
Equipment: |  | Nist Traceable Thermometer (T-Cal 2)
Date: |  | Recorder: | NIST Thermometer: | Verified by: | Pass / Fail:
2009-03-23 |  | 69° F | 73.0° F | F. Griffith | Adjustment made.
2009-05-05 |  | 73° F | 72.6° F | F. Griffith | Pass
2009-09-03 |  | 72° F | 72.0° F | F. Griffith | Pass
2009-09-15 |  | 72° F | 73.0° F | F. Griffith | Pass
2010-05-19 |  | 73° F | 73.0° F | F. Griffith | Pass
2010-12-13 |  | 73°F | 73.0° F | F. Griffith | Pass
2011-05-09 |  | 73° F | 73.0° F | F. Griffith | Pass
2011-05-16 |  | 73° F | 73.0° F | F. Griffith | Pass
2011-05-19 |  | 73° F | 73.0° F | F. Griffith | Pass
2011-12-12 |  | 73° F | 73.4° F | F. Griffith | Pass
[6 empty] | Initial by: | sgw

## Water Tank Temperature
Water Storage Tank Temperature Worksheet
Equipment ID: |  | Water Storage Tank |  | Date: |  | Weekly
Manufacturer: |  | County Line |  | Performed by: |  | Various
Model #: |  | 100 gal (2' x 4' x 2')
Serial #: |  | NA
UofA ID: |  | NA
Location: |  | Concrete Moist Room
| Calibration Item: | Maintenance of 23.0 ± 2 °C (73.5 ± 3.5° F)
Maintenance Item: | Calibration Procedure: | Verify that correct temperature is being maintained in tank
Interval: | Calibration Equipment: | Weekly when in use
(Allowable Range 70 - 77 °F)
Date: |  | Wheel Temperature: | Comments: | Verified by: |  |  | Pass / Fail
2009-03-23 |  | 69 °F | NIST temp 73°F - adj. recorder | F. Griffith |  |  | Pass
2009-03-30 |  | 73 °F |  | F. Griffith |  |  | Pass
2009-04-06 |  | 74 °F |  | F. Griffith |  |  | Pass
2009-04-13 |  | 73 °F |  | F. Griffith |  |  | Pass
2009-04-20 |  | 73 °F |  | F. Griffith |  |  | Pass
2009-04-27 |  | 74 °F |  | F. Griffith |  |  | Pass
2009-05-04 |  | 73 °F |  | F. Griffith |  |  | Pass
2009-05-12 |  | 72 °F | Water changed | F. Griffith |  |  | Pass
2009-05-18 |  | 72 °F |  | F. Griffith |  |  | Pass
2009-05-26 |  | 72 °F |  | F. Griffith |  |  | Pass
2009-06-01 |  | 72 °F |  | F. Griffith |  |  | Pass
2009-06-08 |  | 73 °F |  | R. Conley |  |  | Pass
2009-06-15 |  | 73 °F |  | F. Griffith |  |  | Pass
2009-06-22 |  | 73 °F |  | R. Conley |  |  | Pass
2009-06-29 |  | 73 °F |  | F. Griffith |  |  | Pass
2009-07-09 |  | 73 °F |  | F. Griffith |  |  | Pass
2009-07-13 |  | 74 °F |  | F. Griffith |  |  | Pass
2009-07-20 |  | 73 °F |  | F. Griffith |  |  | Pass
2009-07-27 |  | 73 °F |  | F. Griffith |  |  | Pass
2009-08-03 |  | 74 °F |  | F. Griffith |  |  | Pass
2009-08-10 |  | 73 °F |  | F. Griffith |  |  | Pass
2009-08-17 |  | 73 °F |  | F. Griffith |  |  | Pass
2009-08-26 |  | 74 °F |  | F. Griffith |  |  | Pass
2009-09-01 |  | 75 °F | Water added | F. Griffith |  |  | Pass
2009-09-03 |  | 72 °F |  | F. Griffith |  |  | Pass
2009-09-07 |  | 72 °F |  | F. Griffith |  |  | Pass
2010-05-12 |  | 72 °F | Water Added and Stirred | F. Griffith |  |  | Pass
2010-05-19 |  | 72 °F |  | F. Griffith |  |  | Pass
2010-05-26 |  | 73 °F |  | F. Griffith |  |  | Pass
2010-12-06 |  | 73 °F |  | F. Griffith |  |  | Pass
2010-12-13 |  | 73 °F |  | F. Griffith |  |  | Pass
Wheel saved and temperature recorded verified by calibrated thermometer.
Initial by:

## Water Tank Temperature (2)
Water Storage Tank Temperature Worksheet
Equipment ID: |  | Water Storage Tank |  | Date: |  | Weekly
Manufacturer: |  | County Line |  | Performed by: |  | Various
Model #: |  | 100 gal (2' x 4' x 2')
Serial #: |  | NA
UofA ID: |  | NA
Location: |  | Concrete Moist Room
| Calibration Item: | Maintenance of 23.0 ± 2 °C (73.5 ± 3.5° F)
Maintenance Item: | Calibration Procedure: | Verify that correct temperature is being maintained in tank
Interval: | Calibration Equipment: | Weekly when in use
(Allowable Range 70 - 77 °F)
Date: |  | Wheel Temperature: | Comments: | Verified by: |  |  | Pass / Fail
2011-05-08 |  | 73 °F |  | F. Griffith |  |  | Pass
2011-05-16 |  | 73 °F |  | F. Griffith |  |  | Pass
2011-05-19 |  | 73 °F |  | F. Griffith |  |  | Pass
2011-12-05 |  | 74 °F |  | F. Griffith |  |  | Pass
2011-12-12 |  | 73 °F |  | F. Griffith |  |  | Pass
2012-05-10 |  | 72 °F | Water changed | M. Fleck |  |  | Pass
2012-05-17 |  | 73 °F |  | M. Fleck |  |  | Pass
Wheel saved and temperature recorded verified by calibrated thermometer.
Initial by: