## Air Pressure Meter (1)
Air Pressure Meter Calibration Worksheet
|  | Equipment ID: | Air Pressure Meter |  |  |  | Date: | 2012-02-23
|  | Manufacturer: | Gilson |  |  |  | Performed by: | F. Griffith
|  | Model #: | HM - 30
|  | Serial #: | NA |  |  |  | Next Calibration Due: | May 2012
|  | UofA ID: | NA |  |  |  | Last Calibration: | November 2011
|  | Location: | ST (218)
| Calibration Item: | Verify calibration vessel, measuring bowl, and air readings
| Calibration Procedure: | In-House Procedure #07-3
| Calibration Equipment: | balance read. to 0.01 lb, calibration tubing
Verify the calibration vessel:
|  | Wt. of vessel: |  | lbs | Temperature of Water °F |  |  |  | 79.4
|  | Wt. of vessel and water: |  | lbs | Density of Water @ T |  |  |  | 62.221 | lbs/ft³
|  | Wt. of Water in vessel (w): |  | lbs | 70° F |  | 62.301 |  | lbs/ft°
|  | Wt. of pressure bowl + plate: |  | lbs | 75° F |  | 62.261 |  | lbs/ft°
|  | Wt. of water + pressure bowl + plate: |  | lbs | 80° F |  | 62.216 |  | lbs/ft°
|  | Wt. of Water in pressure bowl (W): |  | lbs | 85° F |  | 62.166 |  | lbs/ft°
|  | Effective Volume (R = w/W): |  | % | Tolerance = Approximately 5.0%
Air pressure bowl diameter should be 0.75 - 1.25 x the height, and capacity of at least 0.20 ft3
|  | Inside Diameter of bowl: |  | in
|  | Inside Height of bowl: |  | in
|  | Diameter/Ht. Ratio: |  |  |  | Tolerance = 0.75 to 1.25
|  | Volume of Bowl: |  | ft³ |  | At least 0.20 ft3
Check of Allowance for Expansion Factor (D):
|  | Initial Pressure: |  | %
|  | Final Pressure: |  | % |  | Final Pressure should be 0%
Check air content gradations on the Pressure Gage:
|  | Initial Pressure: | 3 | %
|  | % Water Removed: | 5 | %
|  | Gage Reading: | 5 | % |  | Tolerance = ± 0.1 %
[8 empty] | Pass/Fail: | Pass
Comments:
[9 empty] | Initial By: | ftg

## Air Pressure Meter (2)
Air Pressure Meter Calibration Worksheet
|  | Equipment ID: | Air Pressure Meter |  |  |  | Date: | 2012-05-30
|  | Manufacturer: | Gilson |  |  |  | Performed by: | M. Fleck
|  | Model #: | HM - 30
|  | Serial #: | NA |  |  |  | Next Calibration Due: | August 2012
|  | UofA ID: | NA |  |  |  | Last Calibration: | New
|  | Location: | ST (218)
| Calibration Item: | Verify calibration vessel, measuring bowl, and air readings
| Calibration Procedure: | In-House Procedure #07-3
| Calibration Equipment: | balance read. to 0.01 lb, calibration tubing
Verify the calibration vessel:
|  | Wt. of vessel: | 0.3313 | lbs | Temperature of Water °F |  |  |  | 79.4
|  | Wt. of vessel and water: | 1.1093 | lbs | Density of Water @ T |  |  |  | 62.221 | lbs/ft³
|  | Wt. of Water in vessel (w): | 0.778 | lbs | 70° F |  | 62.301 |  | lbs/ft°
|  | Wt. of pressure bowl + plate: | 12.2716 | lbs | 75° F |  | 62.261 |  | lbs/ft°
|  | Wt. of water + pressure bowl + plate: | 27.81 | lbs | 80° F |  | 62.216 |  | lbs/ft°
|  | Wt. of Water in pressure bowl (W): | 15.5384 | lbs | 85° F |  | 62.166 |  | lbs/ft°
|  | Effective Volume (R = w/W): | 5.007 | % | Tolerance = Approximately 5.0%
Air pressure bowl diameter should be 0.75 - 1.25 x the height, and capacity of at least 0.20 ft3
|  | Inside Diameter of bowl: | 8.1 | in
|  | Inside Height of bowl: | 8.62 | in
|  | Diameter/Ht. Ratio: | 0.93968 |  |  | Tolerance = 0.75 to 1.25
|  | Volume of Bowl: | 0.25 | ft³ |  | At least 0.20 ft3
Check of Allowance for Expansion Factor (D):
|  | Initial Pressure: | 3 | %
|  | Final Pressure: | 0 | % |  | Final Pressure should be 0%
Check air content gradations on the Pressure Gage:
|  | Initial Pressure: | 3 | %
|  | % Water Removed: | 5 | %
|  | Gage Reading: | 5 | % |  | Tolerance = ± 0.1 %
[8 empty] | Pass/Fail: | Pass
Comments:
[9 empty] | Initial By: | sgw

## Air Pressure Meter (3)
Air Pressure Meter Calibration Worksheet
|  | Equipment ID: | Air Pressure Meter |  |  |  | Date: | 2012-08-02
|  | Manufacturer: | Gilson |  |  |  | Performed by: | M. Fleck
|  | Model #: | HM - 30
|  | Serial #: | NA |  |  |  | Next Calibration Due: | November 2012
|  | UofA ID: | NA |  |  |  | Last Calibration: | May 2012
|  | Location: | ST (218)
| Calibration Item: | Verify calibration vessel, measuring bowl, and air readings
| Calibration Procedure: | In-House Procedure #07-3
| Calibration Equipment: | balance read. to 0.01 lb, calibration tubing
Verify the calibration vessel:
|  | Wt. of vessel: |  | lbs | Temperature of Water °F
|  | Wt. of vessel and water: |  | lbs | Density of Water @ T |  |  |  |  | lbs/ft³
|  | Wt. of Water in vessel (w): |  | lbs | 70° F |  | 62.301 |  | lbs/ft°
|  | Wt. of pressure bowl + plate: |  | lbs | 75° F |  | 62.261 |  | lbs/ft°
|  | Wt. of water + pressure bowl + plate: |  | lbs | 80° F |  | 62.216 |  | lbs/ft°
|  | Wt. of Water in pressure bowl (W): |  | lbs | 85° F |  | 62.166 |  | lbs/ft°
|  | Effective Volume (R = w/W): |  | % | Tolerance = Approximately 5.0%
Air pressure bowl diameter should be 0.75 - 1.25 x the height, and capacity of at least 0.20 ft3
|  | Inside Diameter of bowl: |  | in
|  | Inside Height of bowl: |  | in
|  | Diameter/Ht. Ratio: |  |  |  | Tolerance = 0.75 to 1.25
|  | Volume of Bowl: |  | ft³ |  | At least 0.20 ft3
Check of Allowance for Expansion Factor (D):
|  | Initial Pressure: |  | %
|  | Final Pressure: |  | % |  | Final Pressure should be 0%
Check air content gradations on the Pressure Gage:
|  | Initial Pressure: | 3 | %
|  | % Water Removed: | 5 | %
|  | Gage Reading: | 5 | % |  | Tolerance = ± 0.1 %
[8 empty] | Pass/Fail: | Pass
Comments:
Initial By:

## Air Pressure Meter (4)
Air Pressure Meter Calibration Worksheet
|  | Equipment ID: | Air Pressure Meter |  |  |  | Date: | 2012-11-12
|  | Manufacturer: | Gilson |  |  |  | Performed by: | M. Fleck
|  | Model #: | HM - 30
|  | Serial #: | NA |  |  |  | Next Calibration Due: | February 2013
|  | UofA ID: | NA |  |  |  | Last Calibration: | August 2012
|  | Location: | ST (218)
| Calibration Item: | Verify calibration vessel, measuring bowl, and air readings
| Calibration Procedure: | In-House Procedure #07-3
| Calibration Equipment: | balance read. to 0.01 lb, calibration tubing
Verify the calibration vessel:
|  | Wt. of vessel: |  | lbs | Temperature of Water °F
|  | Wt. of vessel and water: |  | lbs | Density of Water @ T |  |  |  |  | lbs/ft³
|  | Wt. of Water in vessel (w): |  | lbs | 70° F |  | 62.301 |  | lbs/ft°
|  | Wt. of pressure bowl + plate: |  | lbs | 75° F |  | 62.261 |  | lbs/ft°
|  | Wt. of water + pressure bowl + plate: |  | lbs | 80° F |  | 62.216 |  | lbs/ft°
|  | Wt. of Water in pressure bowl (W): |  | lbs | 85° F |  | 62.166 |  | lbs/ft°
|  | Effective Volume (R = w/W): |  | % | Tolerance = Approximately 5.0%
Air pressure bowl diameter should be 0.75 - 1.25 x the height, and capacity of at least 0.20 ft3
|  | Inside Diameter of bowl: |  | in
|  | Inside Height of bowl: |  | in
|  | Diameter/Ht. Ratio: |  |  |  | Tolerance = 0.75 to 1.25
|  | Volume of Bowl: |  | ft³ |  | At least 0.20 ft3
Check of Allowance for Expansion Factor (D):
|  | Initial Pressure: |  | %
|  | Final Pressure: |  | % |  | Final Pressure should be 0%
Check air content gradations on the Pressure Gage:
|  | Initial Pressure: | 3 | %
|  | % Water Removed: | 5 | %
|  | Gage Reading: | 5 | % |  | Tolerance = ± 0.1 %
[8 empty] | Pass/Fail: | Pass
Comments:
Initial By:

## APM Strike Off Plate
Air Pressure Meter Strike Off Plate Calibration Worksheet
|  | Equipment ID: | Strike Off Plate |  | Date: | 2012-05-30
|  | Manufacturer: | Gilson |  | Performed by: | M. Fleck
|  | Model #: | HM SO Plate
|  | Serial #: | NA
|  | UofA ID: | NA |  | Calibration Interval: | 12 months
|  | Location: | ST (218)
| Calibration Item: | Length, width, and thickness of bar
| Calibration Procedure: | ASTM C 231, Section 4.11
| Calibration Equipment: | Mitutoyo Digital calipers
Date | Length (2"≥ bowl dia.) | Width (2"≥ bowl dia.) | Thickness (≥ 13 mm) | Comments |  |  | Pass/Fail | Initial
2012-05-30 | 14" | 14" | 17.8955 | New Plate |  |  | Pass
Comments:

## APM Strike Off Bar
Air Pressure Meter Strike Off Bar Calibration Worksheet
|  | Equipment ID: | Stike Off Bar |  | Date: | 2012-05-30
|  | Manufacturer: | Gilson |  | Performed by: | M. Fleck
|  | Model #: | RPHM-30KEY34R
|  | Serial #: | NA
|  | UofA ID: | NA |  | Calibration Interval: | 12 months
|  | Location: | ST (218)
| Calibration Item: | Length, width, and thickness of bar
| Calibration Procedure: | ASTM C 231, Section 4.11
| Calibration Equipment: | Mitutoyo Digital calipers
Date | Length (≥ 300 mm) | Width (≥ 20 mm) | Thickness (≥ 3 mm) | Comments |  |  | Pass/Fail | Initial
2012-05-30 | 305 mm | 25.54 mm | 4.95 mm | New Bar |  |  | Pass | sgw
Comments: