## Slump Cone
Slump Cone
|  | Equipment ID: | Slump Cone |  |  | Date: |  | 2013-05-06
|  | Manufacturer: | Gilson |  |  | Performed by: |  | M. Fleck
|  | Model #: | HM - 45
|  | Serial #: | NA |  |  | Next Calibration Due: |  | May 2014
|  | UofA ID: | NA |  |  | Last Calibration: |  | May 2012
|  | Location: | ST (218)
| Calibration Item: | Verify critical dimensions of the cone
| Calibration Procedure: | In-House Procedure #07-1
| Calibration Equipment: | Mititoyo Digimatic Calipers, and ruler read. to 1 mm.
|  | Thickness of cone: | 0.0672" | 0.0715" | Tolerance = Not less than 0.060 inches
|  | Bottom Inside Diameter of Cone: | 8" | 8" | Tolerance = 8 +/- 1/8 inches
|  | Top Inside Diameter of Cone: | 4" | 4" | Tolerance = 4 +/- 1/8 inches
|  | Height of Cone: | 12" | 12" | Tolerance = 12 +/- 1/8 inches
|  | Cone Interior Smooth: | Yes | Yes | Pass / Fail
|  | Base and top opening parallel: | Yes | Yes | Pass / Fail
[7 empty] | Pass/Fail: | Pass
Comments:
Initial By:

## Tamping Rod 24
24" Tamping Rod
|  | Equipment ID: | 24" Tamping Rod |  |  | Date: | May
|  | Manufacturer: | Test Mark |  |  | Performed by: | Various
|  | Model #: | CA - 0600 - T Rod
|  | Serial #: | NA
|  | UofA ID: | NA |  |  | Calibration Interval: | 12 months
|  | Location: | ST (218)
| Calibration Item: | Length, diameter, and shape of rod
| Calibration Procedure: | In-House #07-2
| Calibration Equipment: | Mitutoyo Digital calipers, tape measure
Date | Length (approx. 24") | Diameter (approx. 5/8") | Hemispherical Tip? |  | Comments |  |  | Pass/Fail | Initial
2012-05-21 | 24" | 0.625" | Yes | New Rod |  |  |  | Pass | MF
2013-05-06 | 24" | 0.625" | Yes |  |  |  |  | Pass | MF
| Comments: | Diameter: 5/8" = 0.625"
Initial By:

## Rubber Mallet
Rubber Mallet Calibration Worksheet
|  | Equipment ID: | Rubber Mallet |  | Date: | May
|  | Manufacturer: | Test Mark |  | Performed by: | Various
|  | Model #: | CA - 0470
|  | Serial #: | NA
|  | UofA ID: | NA |  | Calibration Interval: | 12 months
|  | Location: | ST (218)
| Calibration Item: | Length, width, and thickness of plate
| Calibration Procedure: | ASTM C 231 - 4.10, ASTM C 138 - 4.6, ASTM C 173 - 4.11
| Calibration Equipment: | Balance
Date | Weight (0.75 - 1.75 lb) |  | General Condition | Comments |  |  | Pass/Fail | Initial
2012-05-21 | 1.4 | lbs | New |  |  |  | Pass | MF
2013-05-06 | 1.4 | lbs |  |  |  |  | Pass | MF
lbs
lbs
lbs
lbs
lbs
lbs
lbs
lbs
lbs
lbs
lbs
lbs
lbs
lbs
lbs
lbs
lbs
| Comments: | Mallet Weight = 600 ± 200 g = 1.25 ±0.50 lbs
Initial By:

## Unit Weight Bucket
Unit Weight Bucket Calibration Worksheet
|  | Equipment ID: | Unit Weight Bucket |  |  | Date: | 2013-05-06
|  | Manufacturer: | Gilson |  |  | Performed by: | M. Fleck
|  | Model #: | HM - 30
|  | Serial #: | NA |  |  | Next Calibration Due: | May 2014
|  | UofA ID: | NA |  |  | Last Calibration: | May 2012
|  | Location: | ST (218)
| Calibration Item: | Verify the unit weight bucket volume
| Calibration Procedure: | In-House Procedure #07-3
| Calibration Equipment: | balance read. to 0.01 lb, glass plate, thermometer
Verify the unit weight bucket:
|  | Wt. of bucket: | 8.5 | lbs
|  | Wt. of glass: | 2.93 | lbs
|  | Wt. of bucket + glass + water: | 26.99 | lbs
|  | Wt. of Water in vessel (w): | 15.56 | lbs
|  | Temperature of water (T): | 73° F
|  | Density of Water at T (C 29): | 62.277 | lbs/ft³
|  | Volume of bucket (mass/density): | 0.24985 | ft³ | Tolerance = Approximately 0.25 cf
[7 empty] | Pass/Fail: | Pass
Comments:
Initial By: