## Kessler (Dial)(SP 7)
Thermometer Calibration Worksheet
Equipment ID: | Thermometer (Sulfur Pot) 7 |  |  | Date: | 4/23/15
Manufacturer: | Kessler |  |  | Performed by: | M. Fleck
Model #: | Gilson MA - 122
Serial #: | NA |  |  | Next Calibration Due: | April 2016
UofA ID: | 30 - 300° F (2°) |  |  | Last Calibration: | April 2014
Location: | STO (218)
|  |  |  | Typical Set Temperature: | 265 °F | deg.
(Sulfur)
| Calibration Item: |  | Verify temperature readings
| Calibration Procedure: |  | In-House Procedure #00-9
| Calibration Equipment: |  | NIST Traceable Calibration Thermometers, ASTM 63F, 65F, 67F
Temperature Measurement (deg.)
| Temp. Setting |  | Reading | Correction
| 40.6 |  | 40.5 | 0.1
| 209 |  | 209 | 0
| 283 |  | 284 | -1
.
| Pass/Fail: |  | Pass
| Comments: |  | No correction needed.
Initial By:

## Kessler (Dial)(CA 2)
Thermometer Calibration Worksheet
Equipment ID: | Concrete Thermometer 2 |  |  | Date: | 4/23/15
Manufacturer: | Gilson (Kessler) |  |  | Performed by: | M. Fleck
Model #: | MA 121F
Serial #: | NA |  |  | Next Calibration Due: | April 2016
UofA ID: | 25 - 125° F (1 °) |  |  | Last Calibration: | April 2014
Location: | Stor (218)
|  |  |  | Typical Set Temperature: | 70 ° F | deg.
| Calibration Item: |  | Verify temperature readings
| Calibration Procedure: |  | In-House Procedure #00-9
| Calibration Equipment: |  | NIST Traceable Calibration Thermometers, ASTM 63F, cat# 1322
Temperature Measurement (deg.)
| Temp. Setting |  | Reading | Correction
| 41.5 |  | 41.5 | 0
| 66.6 |  | 67 | -0.4
| 111.6 |  | 112 | -0.4
.
| Pass/Fail: |  | Pass
| Comments: |  | No correction needed.
Initial By: