## Compression Machine (1)
Compression Machine Maintenance Worksheet
Equipment ID: |  | Compression Tester |  |  |  |  | Date: | Aug. 2010
Manufacturer: |  | TMI |  |  |  |  | Performed by: | Various
Model #: |  | CM - 5000 - GB2
Serial #: |  | 100317 |  |  |  |  | Next Calibration Due:
UofA ID: | [6 empty] | Last Calibration: |  | New
Location: |  | ST (262)
Maintenance Item: | Calibration Item: | Clean and lubricate head
Maintenance Interval: | Calibration Procedure: | Annually or as needed
|  | Head |  | Fluid Level
Date |  | Cleaned | Lubricated | Checked | Adjusted | Performed By
8/12/10 |  | Yes | Yes | Yes | No | FG
8/23/10 |  | Yes | Yes | Yes | No | FG
12/14/11 |  | Yes | Yes | Yes | No | FG
7/26/12 |  | Yes | Yes | Yes | No | FG
11/20/12 |  | Yes | Yes | Yes | No | FG
5/14/13 |  | Yes | Yes | Yes | No | FG
12/17/13 |  | Yes | Yes | Yes | Yes | Calser Corp
5/16/14 |  | Yes | Yes | Yes | No | FG
12/22/14 |  | Yes | Yes | Yes | Yes | Calser Corp
5/19/15 |  | Yes | Yes | Yes | No | FG
Initial by:

## Compression Machine (2)
Compression Machine Maintenance Worksheet
Equipment ID: |  | Compression Tester |  |  |  |  | Date: | Feb. 2006
Manufacturer: |  | Forney |  |  |  |  | Performed by: | Various
Model #: |  | F - 400 F - LC1
Serial #: |  | 95037 |  |  |  |  | Next Calibration Due: | Yearly
UofA ID: | [6 empty] | Last Calibration: |  | New
Location: |  | ENRC 260
Maintenance Item: | Calibration Item: | Clean and lubricate head
Maintenance Interval: | Calibration Procedure: | Annually or as needed
|  | Head |  | Fluid Level
Date |  | Cleaned | Lubricated | Checked | Adjusted | Performed By
2/14/06 |  | Yes | Yes | Yes | No | FG
2/13/07 |  | Yes | Yes | Yes | No | FG
5/15/07 |  | Yes | Yes | Yes | No | FG
2/5/08 |  | Yes | Yes | Yes | No | FG
3/3/09 |  | Yes | Yes | Yes | No | FG
9/15/09 |  | Yes | Yes | Yes | No | FG
3/30/10 |  | Yes | Yes | Yes | No | FG
8/12/10 |  | Yes | Yes | Yes | No | FG
8/23/11 |  | Yes | Yes | Yes | No | FG
12/14/11 |  | Yes | Yes | Yes | No | FG
11/20/12 |  | Yes | Yes | Yes | No | FG
4/23/13 |  | Yes | Yes | Yes | No | FG
5/28/13 |  | No | No | Yes | Yes | FG
12/17/13 |  | Yes | Yes | Yes | No | Calser Corp
5/6/14 |  | Yes | Yes | Yes | No | FG
12/22/14 |  | Yes | Yes | Yes | Yes | Calser Corp
4/23/15 |  | Yes | Yes | Yes | No | FG
Initial by:

## Capping Compound
Capping Compound Calibration Worksheet
|  | Equipment ID: | Ingot Capping Compound |  |  | Date Received: | 2009-03-15
|  | Manufacturer: | Gilson |  |  | Received By: | F. Griffith
|  | Model #: | Red Iron 9000 HM - 189 |  |  | Amount in Order: | 3 Bags
|  | Serial #: | NA
|  | UofA ID: | NA |  |  | Calibration Interval: | 3 months
|  | Location: | RM (264) Chemical Room |  |  |  | or before use
| Calibration Item: | Verify compressive strength of mortar
| Calibration Procedure: | ASTM C 617 - section 5
| Calibration Equipment: | cube mold, melting pot, capping compound, compression testing machine
Compressive Strength (psi)
Date | Cube 1 | Cube 2 | Cube 3 | Anticipated Cylinder Strength (psi) |  |  | Pass/Fail | Initial
2012-03-05 | 5856 | 5736 | 5785 |  | 5000 |  | Pass | ftg
2012-06-12 | 6532 | 6376 | 6646 |  | 5000 |  | Pass | ftg
2012-11-08 | 6672 | 6534 | 6721 |  | 5000 |  | Pass | ftg
2013-04-23 | 8127 | 7899 | 8015 |  | 5000 |  | Pass | ftg
2013-06-27 | 7074 | 7155 | 7046 |  | 5000 |  | Pass | ftg
2013-08-20 | 6154 | 6258 | 6068 |  | 5000 |  | Pass | ftg
2013-10-15 | 7063 | 6408 | 6870 |  | 5000 |  | Pass | ftg
2014-02-25 | 7419 | 7236 | 6974 |  | 5000 |  | Pass | ftg
2014-05-19 | 5834 | 6153 | 6320 |  | 5000 |  | Pass | ftg
2014-08-17 | 6860 | 5655 | 6243 |  | 5000 |  | Pass | ftg
2014-11-11 | 7225 | 7595 | 7795 |  | 5000 |  | Pass | ftg
2015-01-22 | 7738 | 7630 | 7555 |  | 5000 |  | Pass | ftg
2015-04-23 | 8040 | 8215 | 8293 |  | 5000 |  | Pass | ftg
2015-07-16 | 8248 | 8032 | 8083 |  | 5000 |  | Pass | ftg
2015-10-12 | 7876 | 7966 | 8029 |  | 5000 |  | Pass | ftg
| Comments: | 200 - 400 psi/s |  |  |  |  | Initial by:

## Sulfur Cap Log
Sulfur Cap Log
| Equipment ID: |  | Sulfur Caps | [5 empty] | Date : | 2012
| Manufacturer: |  | NA | [5 empty] | Performed By : | Various
| Model #: |  | NA
| Serial #: |  | NA
| UofA ID: |  | NA | [5 empty] | Calibration Interval: | Daily when capping
| Location: |  | NA
| Calibration Item: |  | Verify sulfur cap thicknesses, planeness, and check for hollow areas
Strengths ≤ 7000 psi : Max. ind. thickness 8 mm; Max. avg. thickness 6 mm
Strengths > 7000 psi : Max. ind. thickness 5 mm; Max. avg. thickness 3 mm
| Calibration Procedure: |  | ASTM C 617 - section 6.2.5
| Calibration Equipment: |  | Mitoyo Calipers, coin, 0.002" feeler gauge, straight edge
Date | Plane ? (0.002" across 3 diameters) | Any hollow areas > 1/4" in diameter? |  | Cap Thickness (nearest 0.2 mm) (measure 6 individual pieces) | [5 empty] | Average Cap Thickness | Performed by | Pass / Fail
2012-07-26 | Yes | No |  | 2.8 | 2.6 | 2.8 | 2.6 | 2.6 | 2.6 | 2.6667 | M. Fleck | Pass
2013-04-23 | Yes | No |  | 4.6 | 5 | 4.8 | 4.6 | 4.8 | 4.6 | 4.7333 | F. Griffith | Pass
2013-08-21 | Yes | No |  | 4.5 | 4.3 | 4.8 | 3.7 | 4.5 | 4.4 | 4.3667 | Matthew W. | Pass
2014-02-25 | Yes | No |  | 4.1 | 4.2 | 4.7 | 4.9 | 3.9 | 4.1 | 4.3167 | F. Griffith | Pass
2014-05-27 | Yes | No |  | 3.7 | 3.6 | 3.7 | 3.5 | 3.8 | 3.9 | 3.7 | F. Griffith | Pass
2014-08-27 | Yes | No |  | 4.5 | 4.4 | 5.7 | 4.7 | 4.4 | 6.5 | 5.0333 | F. Griffith | Pass
2014-11-10 | Yes | No |  | 6.5 | 6.1 | 6.1 | 6 | 6.3 | 5.9 | 6.15 | F. Griffith | Pass
2015-01-22 | No capped cylinders broken in this class. | [9 empty] | F. Griffith
2015-04-23 | Yes | No |  | 4.1 | 3.7 | 3.7 | 4.2 | 3.5 | 3.9 | 3.85 | F. Griffith | Pass
2015-07-16 | No capped cylinders broken in this class. | [9 empty] | F. Griffith
2015-10-12 | No capped cylinders broken in this class. | [9 empty] | F. Griffith
| Comments: |  | One cylinder capped during certification | [6 empty] | Initial by:
classes.

## 2" Cubes
2" Cube Test Set
| Equipment ID: | 2" cube test set |  |  | Date: | 2015-05-19
| Manufacturer: | Gilson |  |  | Performed by: | FG
| Model #: | MCA-6
| Serial #: |  |  | Calibration Interval: |  | Yearly
| UofA ID: |  |  | Next Calibration Due: |  | 2016-05-01
| Location: | ST (218) |  | Last Calibration: |  | 2014-05-05
| Maintenance Item: | Check tightness and lubricate |  | X |  | Completed
| Calibration Procedure: | Verify planeness and critical dimensions
| Calibration Equipment: | Caliper, steel bar, feeler gauges and oil
Cube | Mold parts | Sides Plane? | Distance b/t Sides | Height of Sides |  | Angle 90 ± 0.5°
| fit tightly? | < 0.002" | 2" ± 0.02" | 2" (+0.01 to -0.015) |  | between faces?
1 | Yes | Yes | 2.003 | 2.004 |  | Yes
|  | Yes |  | 2.004
|  | Yes | 2.008 | 2.003
|  | Yes |  | 2.004
2 | Yes | Yes | 2.009 | 2.005 |  | Yes
|  | Yes |  | 2.009
|  | Yes | 2.008 | 2.003
|  | Yes |  | 2.008
3 | Yes | Yes | 2.006 | 2.008 |  | Yes
|  | Yes |  | 2.006
|  | Yes | 1.999 | 2.007
|  | Yes |  | 2.009
Notes : | 0.5° is equivalent to 1/8" over a distance of 12" = 0.02" in 2"
Initial By:

## 6" Alignment Device
Capping Plate & Alignment Device Calibration Worksheet
|  | Equipment ID: | 6" Blue Alignment Device |  | Date: | 2015-05-19
|  | Manufacturer: | Gilson (Cyl. Capper) |  | Performed by: | FG
|  | Model #: | HM - 166
|  | Serial #: | NA |  | Next Calibration Due: | May 2016
|  | UofA ID: | NA |  | Last Calibration: | May 2014
|  | Location: | ST (262)
| Calibration Item: | Verify the device dimensions and condition
| Calibration Procedure: | ASTM C 617 - 4.1
| Calibration Equipment: | digital calipers, straightedge, feeler gauges
Verify the device dimensions:
|  | Thickness of plate: | 0.991" | Tolerance = at least 0.45"
|  | Depth of recessed area: | 7/16" | Tolerance = no more than 1/2"
|  | Diameter of plate: | 7.999" | Tolerance = at least 1" greater than
specimen diameter
Check condition of plate:
|  | Any gouges deeper than 0.01"? | No
Any gouges > 0.05 in² surface area? |  |  | No
Check for planeness and perpendicularity:
|  | Is surface plane within 0.002 inches? | Yes
Are alignment guides perpendicular
|  | within 0.5 degrees (1/8" in 12")? | Yes
[6 empty] | Pass/Fail: | Pass
Comments:
Initial By:

## 4" Alignment Device
Capping Plate & Alignment Device Calibration Worksheet
|  | Equipment ID: | 4" Alignment Device |  | Date: | 2015-05-19
|  | Manufacturer: | Gilson (Cyl. Capper) |  | Performed by: | FG
|  | Model #: | HM - 166
|  | Serial #: | NA |  | Next Calibration Due: | May 2016
|  | UofA ID: | NA |  | Last Calibration: | May 2014
|  | Location: | ST (262)
| Calibration Item: | Verify the device dimensions and condition
| Calibration Procedure: | ASTM C 617 - 4.1
| Calibration Equipment: | digital calipers, straightedge, feeler gauges
Verify the device dimensions:
|  | Thickness of plate: | 0.614" | Tolerance = at least 0.45"
|  | Depth of recessed area: | 3/8" | Tolerance = no more than 1/2"
|  | Diameter of plate: | 5.99" | Tolerance = at least 1" greater than
specimen diameter
Check condition of plate:
|  | Any gouges deeper than 0.01"? | No
Any gouges > 0.05 in² surface area? |  |  | No
Check for planeness and perpendicularity:
|  | Is surface plane within 0.002 inches? | Yes
Are alignment guides perpendicular
|  | within 0.5 degrees (1/8" in 12")? | Yes
[6 empty] | Pass/Fail: | Pass
Comments:
Initial By:

## 6" Retainer Ring (A)
Calibration Worksheet for Retainer Rings
| Equipment ID: | 6" Retainer Ring "A" |  |  | Date: | May
| Manufacturer: | Test Mark |  |  | Performed by: | Various
Model #:
| Serial #: | NA
| UofA ID: | NA |  |  | Calibration Interval: | 12 mos.
| Location: | ST (218)
| Calibration Item: | Verify retainer ring dimensions
| Calibration Procedure: | ASTM C 1231 - 5.3
| Calibration Equipment: | digital calipers
Tolerance | (1.0 ± 0.1") | (102 -107%) (6.12" - 6.42") | Plane to 0.002 in? | Any > 0.010 in? | Any > 0.05 in²?
Date | Inside Height of Ring (in) | Inside Dia. of Ring (in) | Base Planeness | Gouge Depth | Gouge Area | Pass/Fail | Initial
2007-05-14 | 1 | 6.074 |  |  |  | Pass | sgw
2008-05-12 | 1 | 6.125 |  |  |  | Pass | ftg
2009-05-12 | 1.0645 | 6.1215 |  |  |  | Pass | ftg
2010-05-18 | 1.0538 | 6.122 |  |  |  | Pass | ftg
2011-05-16 | 0.938 | 6.12 |  |  |  | Pass | ftg
2012-05-30 | 1.0625 | 6.13 |  |  |  | Pass | ftg
2013-05-14 | 1.021 | 6.1285 | Yes | No | No | Pass | AMM
2014-05-14 | 1.005 | 6.1214 | Yes | No | No | Pass | ftg
2015-05-18 | 1.01 | 6.121 | Yes | No | No | Pass | ftg
| Comments: |  |  |  |  | Initial by:

## 6" Retainer Ring (B)
Calibration Worksheet for Retainer Rings
| Equipment ID: | 6" Retainer Ring "B" |  |  | Date: | May
| Manufacturer: | Test Mark |  |  | Performed by: | Various
Model #:
| Serial #: | NA
| UofA ID: | NA |  |  | Calibration Interval: | 12 mos.
| Location: | ST (218)
| Calibration Item: | Verify retainer ring dimensions
| Calibration Procedure: | ASTM C 1231 - 5.3
| Calibration Equipment: | digital calipers
Tolerance | (1.0 ± 0.1") | (102 -107%) (6.12" - 6.42") | Plane to 0.002 in? | Any > 0.010 in? | Any > 0.05 in²?
Date | Inside Height of Ring (in) | Inside Dia. of Ring (in) | Base Planeness | Gouge Depth | Gouge Area | Pass/Fail | Initial
2007-05-14 | 1 | 6.106 |  |  |  | Pass | sgw
2008-05-12 | 1 | 6.123 |  |  |  | Pass | ftg
2009-05-12 | 1.0527 | 6.1245 |  |  |  | Pass | ftg
2010-05-18 | 1.0529 | 6.124 |  |  |  | Pass | ftg
2011-05-16 | 0.938 | 6.12 |  |  |  | Pass | ftg
2012-05-30 | 1 | 6.12 |  |  |  | Pass | ftg
2013-05-14 | 1.008 | 6.1295 | Yes | No | No | Pass | AMM
2014-05-14 | 1.001 | 6.126 | Yes | No | No | Pass | ftg
2015-05-19 | 1 | 6.121 | Yes | No | No | Pass | ftg
| Comments: |  |  |  |  | Initial by:

## 4" Retainer Ring (A)
Calibration Worksheet for Retainer Rings
| Equipment ID: | 4" Retainer Ring "A" |  |  | Date: | May
| Manufacturer: | Test Mark |  |  | Performed by: | Various
Model #:
| Serial #: | NA
| UofA ID: | NA |  |  | Calibration Interval: | 12 mos.
| Location: | ST (218)
| Calibration Item: | Verify retainer ring dimensions
| Calibration Procedure: | ASTM C 1231 - 5.3
| Calibration Equipment: | digital calipers
Tolerance | (1.0 ± 0.1") | (102 -107%) (4.08" - 4.28") | Plane to 0.002 in? | Any > 0.010 in? | Any > 0.05 in²?
Date | Inside Height of Ring (in) | Inside Dia. of Ring (in) | Base Planeness | Gouge Depth | Gouge Area | Pass/Fail | Initial
2015-05-19 | 1.004 | 4.113 | Yes | No | No | Pass | ftg
| Comments: |  |  |  |  | Initial by:

## 4" Retainer Ring (B)
Calibration Worksheet for Retainer Rings
| Equipment ID: | 4" Retainer Ring "B" |  |  | Date: | May
| Manufacturer: | Test Mark |  |  | Performed by: | Various
Model #:
| Serial #: | NA
| UofA ID: | NA |  |  | Calibration Interval: | 12 mos.
| Location: | ST (218)
| Calibration Item: | Verify retainer ring dimensions
| Calibration Procedure: | ASTM C 1231 - 5.3
| Calibration Equipment: | digital calipers
Tolerance | (1.0 ± 0.1") | (102 -107%) (4.08" - 4.28") | Plane to 0.002 in? | Any > 0.010 in? | Any > 0.05 in²?
Date | Inside Height of Ring (in) | Inside Dia. of Ring (in) | Base Planeness | Gouge Depth | Gouge Area | Pass/Fail | Initial
2015-05-19 | 1.001 | 4.122 | Yes | No | No | Pass | ftg
| Comments: |  |  |  |  | Initial by:

## 6" Neoprene Pads
Calibration Worksheet for Elastomeric Pads
|  | Equipment ID: | 6" Elastomeric Pad - # 70 |  |  |  | Date: | 2012-05-23
|  | Manufacturer: | Gilson |  |  |  | Performed by: | F. Griffith
|  | Model #: | HM - 177
|  | Serial #: | NA
|  | UofA ID: | NA |  |  |  | Calibration Interval: | prior to first use
|  | Location: | ST (218)
| Calibration Item: | Verify pad dimensions
| Calibration Procedure: | ASTM C 1231 - 5.2
| Calibration Equipment: | digital calipers
Tolerance | (4.9375" - 5.0625") |  | Not < 1/16" of Ret. Ring (6.0575" - 6.4200")
| Pad Thickness (in) |  | Pad Diameter (in)
Date | Pad 01 | Pad 02 | Pad 01 | Pad 02 |  | Any cracks > 3/8"? |  | Pass/Fail | Initial
2012-05-23 | 0.4940" | 0.5100" | 6.1100" | 6.1200" | No |  |  | Pass | ftg
2013-05-14 | 0.4850" | 0.5005" | 6.0635" | 6.0550" | No |  |  | Pass | AMM
2014-05-14 | 0.4940" | 0.5020" | 6.1101" | 6.0975" | No |  |  | Pass | ftg
2014-08-27 | Pads #01 and #02 taken out of service. | [7 empty] | ftg
| Comments: | [6 empty] | Initial By:

## 6" Neoprene Pads (2)
Calibration Worksheet for Elastomeric Pads
|  | Equipment ID: | 6" Elastomeric Pad - # 70 |  |  | Date: | 2014-08-27
|  | Manufacturer: | Gilson |  |  | Performed by: | F. Griffith
|  | Model #: | HM - 177
|  | Serial #: | NA
|  | UofA ID: | NA |  |  | Calibration Interval: | prior to first use
|  | Location: | ST (218)
| Calibration Item: | Verify pad dimensions
| Calibration Procedure: | ASTM C 1231 - 5.2
| Calibration Equipment: | digital calipers
| (4.9375" - 5.0625") |  | Not < 1/16" of Ret. Ring (6.0575" - 6.4200") |  | Exceded Table 1 Reuse Limits?
| Pad Thickness (in) |  | Pad Diameter (in)
Date | Pad 03 | Pad 04 | Pad 03 | Pad 04 |  |  | Pass/Fail | Initial
2014-08-27 | 0.505" | 0.504" | 6.121" | 6.054" | No |  | Pass | ftg
2015-05-18 | 0.505" | 0.502" | 6.079" | 6.048" | No |  | Pass | ftg
| Comments: | [5 empty] | Initial By:

## 4" Neoprene Pads
Calibration Worksheet for Elastomeric Pads
|  | Equipment ID: | 4" Elastomeric Pad - # 60 |  |  | Date: | 2015-05-18
|  | Manufacturer: | Gilson |  |  | Performed by: | F. Griffith
|  | Model #: | HM - 361
|  | Serial #: | NA
|  | UofA ID: | NA |  |  | Calibration Interval: | prior to first use
|  | Location: | ST (218)
| Calibration Item: | Verify pad dimensions
| Calibration Procedure: | ASTM C 1231 - 5.2
| Calibration Equipment: | digital calipers
| (4.9375" - 5.0625") |  | Not < 1/16" of Ret. Ring (4.0175" - 4.2175") |  | Exceded Table 1 Reuse Limits?
| Pad Thickness (in) |  | Pad Diameter (in)
Date | Pad 01 | Pad 02 | Pad 01 | Pad 02 |  |  | Pass/Fail | Initial
2015-05-18 | 0.490" | 0.501" | 4.061" | 4.073" | No |  | Pass | ftg
| Comments: | [5 empty] | Initial By:

## 6" Pad Usage Log
6" Pad Usage Log Sheet
|  | Pad ID: | 01 |  |  | Durometer: | 70 | [5 empty] | Date Entered into Service: | 2012-05-17
|  | Initial | Date |  |  | Initial | Date |  |  | Initial | Date |  |  | Initial | Date
X | 1 | FG | 5/17/12 | X | 26 | FG | 2013-04-24 | X | 51 | FG | 2014-06-28 | X | 76 | FG | 2014-02-05
X | 2 | FG | 5/17/12 | X | 27 | FG | 2013-04-24 | X | 52 | FG | 2014-06-28 | X | 77 | FG | 2014-02-05
X | 3 | FG | 5/17/12 | X | 28 | FG | 2013-04-24 | X | 53 | FG | 2014-08-21 | X | 78 | FG | 2014-02-05
X | 4 | FG | 5/17/12 | X | 29 | FG | 2013-04-24 | X | 54 | FG | 2014-08-21 | X | 79 | FG | 2014-05-21
X | 5 | FG | 2012-06-12 | X | 30 | FG | 2013-04-24 | X | 55 | FG | 2014-08-21 | X | 80 | OOS | 2014-08-27
X | 6 | FG | 2012-06-13 | X | 31 | FG | 2013-04-24 | X | 56 | FG | 2014-08-21 |  | 81
X | 7 | FG | 2012-06-13 | X | 32 | FG | 2013-04-24 | X | 57 | FG | 2014-08-21 |  | 82
X | 8 | FG | 2012-06-13 | X | 33 | FG | 2013-04-24 | X | 58 | FG | 2014-08-21 |  | 83
X | 9 | FG | 2012-06-13 | X | 34 | FG | 2013-04-24 | X | 59 | FG | 2014-08-21 |  | 84
X | 10 | FG | 2012-07-25 | X | 35 | FG | 2013-05-09 | X | 60 | FG | 2014-08-21 |  | 85
X | 11 | FG | 2012-11-08 | X | 36 | FG | 2013-05-09 | X | 61 | FG | 2014-08-21 |  | 86
X | 12 | FG | 2012-11-09 | X | 37 | FG | 2013-05-09 | X | 62 | FG | 2014-08-21 |  | 87
X | 13 | FG | 2012-11-09 | X | 38 | FG | 2013-05-09 | X | 63 | FG | 2014-08-21 |  | 88
X | 14 | FG | 2012-11-09 | X | 39 | FG | 2013-05-09 | X | 64 | FG | 2014-08-21 |  | 89
X | 15 | FG | 2012-11-09 | X | 40 | FG | 2013-05-09 | X | 65 | FG | 2014-08-21 |  | 90
X | 16 | FG | 2012-11-09 | X | 41 | FG | 2014-06-28 | X | 66 | FG | 2014-08-21 |  | 91
X | 17 | FG | 2012-11-09 | X | 42 | FG | 2014-06-28 | X | 67 | FG | 2014-02-05 |  | 92
X | 18 | FG | 2012-11-09 | X | 43 | FG | 2014-06-28 | X | 68 | FG | 2014-02-05 |  | 93
X | 19 | FG | 2012-11-09 | X | 44 | FG | 2014-06-28 | X | 69 | FG | 2014-02-05 |  | 94
X | 20 | FG | 2012-11-20 | X | 45 | FG | 2014-06-28 | X | 70 | FG | 2014-02-05 |  | 95
X | 21 | FG | 2012-11-20 | X | 46 | FG | 2014-06-28 | X | 71 | FG | 2014-02-05 |  | 96
X | 22 | FG | 2012-11-20 | X | 47 | FG | 2014-06-28 | X | 72 | FG | 2014-02-05 |  | 97
X | 23 | FG | 2012-11-20 | X | 48 | FG | 2014-06-28 | X | 73 | FG | 2014-02-05 |  | 98
| 24 |  |  | X | 49 | FG | 2014-06-28 | X | 74 | FG | 2014-02-05 |  | 99
| 25 |  |  | X | 50 | FG | 2014-06-28 | X | 75 | FG | 2014-02-05 |  | 100
[13 empty] | Last Date of Service: | 2014-08-27

## 6" Pad Usage Log (2)
6" Pad Usage Log Sheet
|  | Pad ID: | 02 |  |  | Durometer: | 70 | [5 empty] | Date Entered into Service: | 2012-05-17
|  | Initial | Date |  |  | Initial | Date |  |  | Initial | Date |  |  | Initial | Date
X | 1 | FG | 5/17/12 | X | 26 | FG | 2013-04-24 | X | 51 | FG | 2014-06-28 | X | 76 | FG | 2014-02-05
X | 2 | FG | 5/17/12 | X | 27 | FG | 2013-04-24 | X | 52 | FG | 2014-06-28 | X | 77 | FG | 2014-02-05
X | 3 | FG | 5/17/12 | X | 28 | FG | 2013-04-24 | X | 53 | FG | 2014-08-21 | X | 78 | FG | 2014-02-05
X | 4 | FG | 5/17/12 | X | 29 | FG | 2013-04-24 | X | 54 | FG | 2014-08-21 | X | 79 | FG | 2014-05-21
X | 5 | FG | 2012-06-12 | X | 30 | FG | 2013-04-24 | X | 55 | FG | 2014-08-21 |  | 80 | OOS | 2014-08-27
X | 6 | FG | 2012-06-13 | X | 31 | FG | 2013-04-24 | X | 56 | FG | 2014-08-21 |  | 81
X | 7 | FG | 2012-06-13 | X | 32 | FG | 2013-04-24 | X | 57 | FG | 2014-08-21 |  | 82
X | 8 | FG | 2012-06-13 | X | 33 | FG | 2013-04-24 | X | 58 | FG | 2014-08-21 |  | 83
X | 9 | FG | 2012-06-13 | X | 34 | FG | 2013-04-24 | X | 59 | FG | 2014-08-21 |  | 84
X | 10 | FG | 2012-07-25 | X | 35 | FG | 2013-05-09 | X | 60 | FG | 2014-08-21 |  | 85
X | 11 | FG | 2012-11-08 | X | 36 | FG | 2013-05-09 | X | 61 | FG | 2014-08-21 |  | 86
X | 12 | FG | 2012-11-09 | X | 37 | FG | 2013-05-09 | X | 62 | FG | 2014-08-21 |  | 87
X | 13 | FG | 2012-11-09 | X | 38 | FG | 2013-05-09 | X | 63 | FG | 2014-08-21 |  | 88
X | 14 | FG | 2012-11-09 | X | 39 | FG | 2013-05-09 | X | 64 | FG | 2014-08-21 |  | 89
X | 15 | FG | 2012-11-09 | X | 40 | FG | 2013-05-09 | X | 65 | FG | 2014-08-21 |  | 90
X | 16 | FG | 2012-11-09 | X | 41 | FG | 2014-06-28 | X | 66 | FG | 2014-08-21 |  | 91
X | 17 | FG | 2012-11-09 | X | 42 | FG | 2014-06-28 | X | 67 | FG | 2014-02-05 |  | 92
X | 18 | FG | 2012-11-09 | X | 43 | FG | 2014-06-28 | X | 68 | FG | 2014-02-05 |  | 93
X | 19 | FG | 2012-11-09 | X | 44 | FG | 2014-06-28 | X | 69 | FG | 2014-02-05 |  | 94
X | 20 | FG | 2012-11-20 | X | 45 | FG | 2014-06-28 | X | 70 | FG | 2014-02-05 |  | 95
X | 21 | FG | 2012-11-20 | X | 46 | FG | 2014-06-28 | X | 71 | FG | 2014-02-05 |  | 96
X | 22 | FG | 2012-11-20 | X | 47 | FG | 2014-06-28 | X | 72 | FG | 2014-02-05 |  | 97
X | 23 | FG | 2012-11-20 | X | 48 | FG | 2014-06-28 | X | 73 | FG | 2014-02-05 |  | 98
| 24 |  |  |  | 49 | FG | 2014-06-28 | X | 74 | FG | 2014-02-05 |  | 99
| 25 |  |  |  | 50 | FG | 2014-06-28 | X | 75 | FG | 2014-02-05 |  | 100
[13 empty] | Last Date of Service: | 2014-08-27

## 6" Pad Usage Log (3)
6" Pad Usage Log Sheet
|  | Pad ID: | 03 |  | Durometer: | 70 |  |  |  | Date Entered into Service: | 2014-08-27
|  | Initial | Date |  | Initial | Date |  | Initial | Date |  | Initial | Date
X | 1 | FG | 2014-08-27 | 26 |  |  | 51 |  |  | 76
X | 2 | FG | 2014-08-27 | 27 |  |  | 52 |  |  | 77
X | 3 | FG | 2014-08-27 | 28 |  |  | 53 |  |  | 78
X | 4 | FG | 2014-08-27 | 29 |  |  | 54 |  |  | 79
X | 5 | FG | 2014-08-27 | 30 |  |  | 55 |  |  | 80
X | 6 | FG | 2014-08-27 | 31 |  |  | 56 |  |  | 81
X | 7 | FG | 2014-11-11 | 32 |  |  | 57 |  |  | 82
X | 8 | FG | 2014-11-12 | 33 |  |  | 58 |  |  | 83
X | 9 | FG | 2014-11-12 | 34 |  |  | 59 |  |  | 84
X | 10 | FG | 2014-11-12 | 35 |  |  | 60 |  |  | 85
X | 11 | FG | 2014-11-12 | 36 |  |  | 61 |  |  | 86
X | 12 | FG | 2014-11-12 | 37 |  |  | 62 |  |  | 87
X | 13 | FG | 2014-11-26 | 38 |  |  | 63 |  |  | 88
X | 14 | FG | 2014-11-26 | 39 |  |  | 64 |  |  | 89
X | 15 | FG | 2014-11-26 | 40 |  |  | 65 |  |  | 90
X | 16 | FG | 2014-11-26 | 41 |  |  | 66 |  |  | 91
| 17 |  |  | 42 |  |  | 67 |  |  | 92
| 18 |  |  | 43 |  |  | 68 |  |  | 93
| 19 |  |  | 44 |  |  | 69 |  |  | 94
| 20 |  |  | 45 |  |  | 70 |  |  | 95
| 21 |  |  | 46 |  |  | 71 |  |  | 96
| 22 |  |  | 47 |  |  | 72 |  |  | 97
| 23 |  |  | 48 |  |  | 73 |  |  | 98
| 24 |  |  | 49 |  |  | 74 |  |  | 99
| 25 |  |  | 50 |  |  | 75 |  |  | 100
Last Date of Service:

## 6" Pad Usage Log (4)
6" Pad Usage Log Sheet
|  | Pad ID: | 04 |  |  | Durometer: | 70 |  |  |  | Date Entered into Service: | 2014-08-27
|  | Initial | Date |  |  | Initial | Date |  | Initial | Date |  | Initial | Date
X | 1 | FG | 2014-08-27 | X | 26 |  |  | 51 |  |  | 76
X | 2 | FG | 2014-08-27 | X | 27 |  |  | 52 |  |  | 77
X | 3 | FG | 2014-08-27 | X | 28 |  |  | 53 |  |  | 78
X | 4 | FG | 2014-08-27 | X | 29 |  |  | 54 |  |  | 79
X | 5 | FG | 2014-08-27 | X | 30 |  |  | 55 |  |  | 80
X | 6 | FG | 2014-08-27 |  | 31 |  |  | 56 |  |  | 81
X | 7 | FG | 2014-11-11 |  | 32 |  |  | 57 |  |  | 82
X | 8 | FG | 2014-11-12 |  | 33 |  |  | 58 |  |  | 83
X | 9 | FG | 2014-11-12 |  | 34 |  |  | 59 |  |  | 84
X | 10 | FG | 2014-11-12 |  | 35 |  |  | 60 |  |  | 85
X | 11 | FG | 2014-11-12 |  | 36 |  |  | 61 |  |  | 86
X | 12 | FG | 2014-11-12 |  | 37 |  |  | 62 |  |  | 87
X | 13 | FG | 2014-11-26 |  | 38 |  |  | 63 |  |  | 88
X | 14 | FG | 2014-11-26 |  | 39 |  |  | 64 |  |  | 89
X | 15 | FG | 2014-11-26 |  | 40 |  |  | 65 |  |  | 90
X | 16 | FG | 2014-11-26 |  | 41 |  |  | 66 |  |  | 91
| 17 |  |  |  | 42 |  |  | 67 |  |  | 92
| 18 |  |  |  | 43 |  |  | 68 |  |  | 93
| 19 |  |  |  | 44 |  |  | 69 |  |  | 94
| 20 |  |  |  | 45 |  |  | 70 |  |  | 95
| 21 |  |  |  | 46 |  |  | 71 |  |  | 96
| 22 |  |  |  | 47 |  |  | 72 |  |  | 97
| 23 |  |  |  | 48 |  |  | 73 |  |  | 98
| 24 |  |  |  | 49 |  |  | 74 |  |  | 99
| 25 |  |  |  | 50 |  |  | 75 |  |  | 100
Last Date of Service:

## 4" Pad Usage Log (1)
4" Pad Usage Log Sheet
|  | Pad ID: | 01 |  | Durometer: | 60 |  |  |  | Date Entered into Service: | 2015-01-21
|  | Initial | Date |  | Initial | Date |  | Initial | Date |  | Initial | Date
X | 1 | FG | 2015-01-22 | 26 |  |  | 51 |  |  | 76
X | 2 | FG | 2015-01-23 | 27 |  |  | 52 |  |  | 77
X | 3 | FG | 2015-04-23 | 28 |  |  | 53 |  |  | 78
X | 4 | FG | 2015-04-23 | 29 |  |  | 54 |  |  | 79
X | 5 | FG | 2015-04-23 | 30 |  |  | 55 |  |  | 80
X | 6 | FG | 2015-04-23 | 31 |  |  | 56 |  |  | 81
X | 7 | FG | 2015-04-23 | 32 |  |  | 57 |  |  | 82
X | 8 | FG | 2015-04-23 | 33 |  |  | 58 |  |  | 83
X | 9 | FG | 2015-04-24 | 34 |  |  | 59 |  |  | 84
X | 10 | FG | 2015-04-24 | 35 |  |  | 60 |  |  | 85
X | 11 | FG | 2015-04-24 | 36 |  |  | 61 |  |  | 86
X | 12 | FG | 2015-04-24 | 37 |  |  | 62 |  |  | 87
X | 13 | FG | 2015-04-24 | 38 |  |  | 63 |  |  | 88
X | 14 | FG | 2015-04-24 | 39 |  |  | 64 |  |  | 89
X | 15 | FG | 2015-07-17 | 40 |  |  | 65 |  |  | 90
X | 16 | FG | 2015-07-17 | 41 |  |  | 66 |  |  | 91
X | 17 | FG | 2015-07-17 | 42 |  |  | 67 |  |  | 92
X | 18 | FG | 2015-07-17 | 43 |  |  | 68 |  |  | 93
X | 19 | FG | 2015-07-17 | 44 |  |  | 69 |  |  | 94
X | 20 | FG | 2015-07-17 | 45 |  |  | 70 |  |  | 95
X | 21 | FG | 2015-07-17 | 46 |  |  | 71 |  |  | 96
| 22 |  |  | 47 |  |  | 72 |  |  | 97
| 23 |  |  | 48 |  |  | 73 |  |  | 98
| 24 |  |  | 49 |  |  | 74 |  |  | 99
| 25 |  |  | 50 |  |  | 75 |  |  | 100
Last Date of Service:

## 4" Pad Usage Log (2)
4" Pad Usage Log Sheet
|  | Pad ID: | 02 |  | Durometer: | 60 |  |  |  | Date Entered into Service: | 2015-01-21
|  | Initial | Date |  | Initial | Date |  | Initial | Date |  | Initial | Date
X | 1 | FG | 2015-01-22 | 26 |  |  | 51 |  |  | 76
X | 2 | FG | 2015-01-23 | 27 |  |  | 52 |  |  | 77
X | 3 | FG | 2015-04-23 | 28 |  |  | 53 |  |  | 78
X | 4 | FG | 2015-04-23 | 29 |  |  | 54 |  |  | 79
X | 5 | FG | 2015-04-23 | 30 |  |  | 55 |  |  | 80
X | 6 | FG | 2015-04-23 | 31 |  |  | 56 |  |  | 81
X | 7 | FG | 2015-04-23 | 32 |  |  | 57 |  |  | 82
X | 8 | FG | 2015-04-23 | 33 |  |  | 58 |  |  | 83
X | 9 | FG | 2015-04-24 | 34 |  |  | 59 |  |  | 84
X | 10 | FG | 2015-04-24 | 35 |  |  | 60 |  |  | 85
X | 11 | FG | 2015-04-24 | 36 |  |  | 61 |  |  | 86
X | 12 | FG | 2015-04-24 | 37 |  |  | 62 |  |  | 87
X | 13 | FG | 2015-04-24 | 38 |  |  | 63 |  |  | 88
X | 14 | FG | 2015-04-24 | 39 |  |  | 64 |  |  | 89
X | 15 | FG | 2015-07-17 | 40 |  |  | 65 |  |  | 90
X | 16 | FG | 2015-07-17 | 41 |  |  | 66 |  |  | 91
X | 17 | FG | 2015-07-17 | 42 |  |  | 67 |  |  | 92
X | 18 | FG | 2015-07-17 | 43 |  |  | 68 |  |  | 93
X | 19 | FG | 2015-07-17 | 44 |  |  | 69 |  |  | 94
X | 20 | FG | 2015-07-17 | 45 |  |  | 70 |  |  | 95
X | 21 | FG | 2015-07-17 | 46 |  |  | 71 |  |  | 96
| 22 |  |  | 47 |  |  | 72 |  |  | 97
| 23 |  |  | 48 |  |  | 73 |  |  | 98
| 24 |  |  | 49 |  |  | 74 |  |  | 99
| 25 |  |  | 50 |  |  | 75 |  |  | 100
Last Date of Service:

## Water Tank Maintenance
Water Storage Tank Maintenance Worksheet
Equipment ID: |  | Water Storage Tank |  |  | Date: | 2012
Manufacturer: |  |  |  |  | Performed by: | Various
Model #:
Serial #: |  |  |  |  | Next Calibration Due:
UofA ID: |  |  |  |  | Last Calibration:
Location: |  | Concrete Moist Room
Maintenance Item: | Calibration Item: | Stir water in tank to replace calcium ions
Maintenance Interval: | Calibration Procedure: | Once per month
Maintenance Item: | Calibration Equipment: | Change water and saturate with sodium hydroxide
Maintenance Interval: |  | Every 24 months
Date: |  | Water Stirred: | Water Changed: | Comments: | Performed by:
2007 - 2008 |  |  |  | OOS | F. Griffith
3/2009 |  | Yes |  | In Service | F. Griffith
4/2009 |  | Yes |  |  | F. Griffith
5/2009 |  | Yes | Yes |  | F. Griffith
6/2009 |  | Yes |  |  | F. Griffith
7/2009 |  | Yes |  |  | F. Griffith
8/2009 |  | Yes |  |  | F. Griffith
9/2009 |  | Yes |  |  | F. Griffith
10/2009 |  | Yes |  |  | F. Griffith
11/2009 |  | Yes |  |  | F. Griffith
12/2009 |  |  |  | OOS | F. Griffith
5/2010 |  | Yes | Yes | In Service | F. Griffith
6/2010 |  |  |  | OOS | F. Griffith
11/2010 |  | Yes |  | In Service | F. Griffith
12/2010 |  |  |  | OOS | F. Griffith
5/2011 |  | Yes |  | In Service | F. Griffith
6/2011 |  |  |  | OOS | F. Griffith
11/2011 |  | Yes |  | In Service | F. Griffith
12/2011 |  |  |  | OOS | F. Griffith
5/2012 |  | Yes | Yes | In Service | M. Fleck
6/2012 |  |  |  | OOS | M. Fleck
7/20/12 |  | Yes |  | In Service | M. Fleck
7/27/12 |  |  |  | OOS | M. Fleck
11/12/12 |  | Yes | No | In Service | M. Fleck
11/26/12 |  |  |  | OOS | M. Fleck
[6 empty] | Initial by: | sgw

## Water Tank Maintenance (2)
Water Storage Tank Maintenance Worksheet
Equipment ID: |  | Water Storage Tank |  |  | Date: | 2013
Manufacturer: |  |  |  |  | Performed by: | Various
Model #:
Serial #: |  |  |  |  | Next Calibration Due:
UofA ID: |  |  |  |  | Last Calibration:
Location: |  | Concrete Moist Room
Maintenance Item: | Calibration Item: | Stir water in tank to replace calcium ions
Maintenance Interval: | Calibration Procedure: | Once per month
Maintenance Item: | Calibration Equipment: | Change water and saturate with sodium hydroxide
Maintenance Interval: |  | Every 24 months
Date: |  | Water Stirred: | Water Changed: | Comments: | Performed by:
5/1/13 |  | Yes | No | In Service | M. Fleck
5/10/13 |  |  |  | OOS | M.Fleck
5/2/14 |  | Yes | Yes | In Service | Matt Watters
5/14/14 |  |  |  | OOS | M. Fleck
8/18/14 |  | Yes | Yes | OOS | M. Fleck
9/8/14 |  | No | No | In Service | M. Fleck
9/10/14 |  | Yes | No | OOS | M. Fleck
11/19/14 |  | Yes | No | In Service | M. Fleck
11/27/14 |  | No | No | OOS | M. Fleck
4/14/15 |  | Yes | Yes | OOS | M. Fleck
5/4/15 |  | Yes | No | In Service | M. Fleck
5/11/15 |  | No | No | OOS | M. Fleck
11/10/15 |  | Yes | No | In Service | M. Fleck
11/29/15 |  | No | No | OOS | M. Fleck
Initial by:

## Water Tank Temperature
Water Storage Tank Temperature Worksheet
Equipment ID: |  | Water Storage Tank |  | Date: |  | Weekly
Manufacturer: |  | County Line |  | Performed by: |  | Various
Model #: |  | 100 gal (2' x 4' x 2')
Serial #: |  | NA
UofA ID: |  | NA
Location: |  | Concrete Moist Room
| Calibration Item: | Maintenance of 23.0 ± 2 °C (73.5 ± 3.5° F)
Maintenance Item: | Calibration Procedure: | Verify that correct temperature is being maintained in tank
Interval: | Calibration Equipment: | Weekly when in use
(Allowable Range 70 - 77 °F)
Date: |  | Wheel Temperature: | Comments: | Verified by: |  |  | Pass / Fail
2009-03-23 |  | 69 °F | NIST temp 73°F - adj. recorder | F. Griffith |  |  | Pass
2009-03-30 |  | 73 °F |  | F. Griffith |  |  | Pass
2009-04-06 |  | 74 °F |  | F. Griffith |  |  | Pass
2009-04-13 |  | 73 °F |  | F. Griffith |  |  | Pass
2009-04-20 |  | 73 °F |  | F. Griffith |  |  | Pass
2009-04-27 |  | 74 °F |  | F. Griffith |  |  | Pass
2009-05-04 |  | 73 °F |  | F. Griffith |  |  | Pass
2009-05-12 |  | 72 °F | Water changed | F. Griffith |  |  | Pass
2009-05-18 |  | 72 °F |  | F. Griffith |  |  | Pass
2009-05-26 |  | 72 °F |  | F. Griffith |  |  | Pass
2009-06-01 |  | 72 °F |  | F. Griffith |  |  | Pass
2009-06-08 |  | 73 °F |  | R. Conley |  |  | Pass
2009-06-15 |  | 73 °F |  | F. Griffith |  |  | Pass
2009-06-22 |  | 73 °F |  | R. Conley |  |  | Pass
2009-06-29 |  | 73 °F |  | F. Griffith |  |  | Pass
2009-07-09 |  | 73 °F |  | F. Griffith |  |  | Pass
2009-07-13 |  | 74 °F |  | F. Griffith |  |  | Pass
2009-07-20 |  | 73 °F |  | F. Griffith |  |  | Pass
2009-07-27 |  | 73 °F |  | F. Griffith |  |  | Pass
2009-08-03 |  | 74 °F |  | F. Griffith |  |  | Pass
2009-08-10 |  | 73 °F |  | F. Griffith |  |  | Pass
2009-08-17 |  | 73 °F |  | F. Griffith |  |  | Pass
2009-08-26 |  | 74 °F |  | F. Griffith |  |  | Pass
2009-09-01 |  | 75 °F | Water added | F. Griffith |  |  | Pass
2009-09-03 |  | 72 °F |  | F. Griffith |  |  | Pass
2009-09-07 |  | 72 °F |  | F. Griffith |  |  | Pass
2010-05-12 |  | 72 °F | Water Added and Stirred | F. Griffith |  |  | Pass
2010-05-19 |  | 72 °F |  | F. Griffith |  |  | Pass
2010-05-26 |  | 73 °F |  | F. Griffith |  |  | Pass
2010-12-06 |  | 73 °F |  | F. Griffith |  |  | Pass
2010-12-13 |  | 73 °F |  | F. Griffith |  |  | Pass
Wheel saved and temperature recorded verified by calibrated thermometer.
[5 empty] | Initial by: |  | sgw

## Water Tank Temperature (2)
Water Storage Tank Temperature Worksheet
Equipment ID: |  | Water Storage Tank |  | Date: |  | Weekly
Manufacturer: |  | County Line |  | Performed by: |  | Various
Model #: |  | 100 gal (2' x 4' x 2')
Serial #: |  | NA
UofA ID: |  | NA
Location: |  | Concrete Moist Room
| Calibration Item: | Maintenance of 23.0 ± 2 °C (73.5 ± 3.5° F)
Maintenance Item: | Calibration Procedure: | Verify that correct temperature is being maintained in tank
Interval: | Calibration Equipment: | Weekly when in use
(Allowable Range 70 - 77 °F)
Date: |  | Wheel Temperature: | Comments: | Verified by: |  |  | Pass / Fail
2011-05-08 |  | 73 °F |  | F. Griffith |  |  | Pass
2011-05-16 |  | 73 °F |  | F. Griffith |  |  | Pass
2011-05-19 |  | 73 °F |  | F. Griffith |  |  | Pass
2011-12-05 |  | 74 °F |  | F. Griffith |  |  | Pass
2011-12-12 |  | 73 °F |  | F. Griffith |  |  | Pass
2012-05-10 |  | 72 °F | Water changed | M. Fleck |  |  | Pass
2012-05-17 |  | 73 °F |  | M. Fleck |  |  | Pass
2012-07-20 |  | 73° F |  | M. Fleck |  |  | Pass
2012-07-27 |  | 73° F |  | M. Fleck |  |  | Pass
2012-11-12 |  | 74° F |  | M. Fleck |  |  | Pass
2012-11-19 |  | 74° F |  | M. Fleck |  |  | Pass
2013-05-01 |  | 74° F |  | M. Fleck |  |  | Pass
2013-05-08 |  | 75° F |  | M. Fleck |  |  | Pass
2014-05-02 |  | 74° F | Water Changed | M. Fleck |  |  | Pass
2014-05-08 |  | 74° F |  | M. Fleck |  |  | Pass
2014-05-14 |  | 74° F |  | M. Fleck |  |  | Pass
2014-08-18 |  | 74° F | Water Changed | M. Fleck |  |  | Pass
2014-09-08 |  | 75° F |  | F. Griffith |  |  | Pass
2014-11-19 |  | 74° F |  | M. Fleck |  |  | Pass
2015-11-26 |  | 74° F |  | F. Griffith |  |  | Pass
2015-04-14 |  | 72° F | Water Changed | M. Fleck |  |  | Pass
2015-05-04 |  | 73° F |  | M. Fleck |  |  | Pass
2015-05-11 |  | 73° F |  | M. Fleck |  |  | Pass
2015-11-10 |  | 73° F |  | M. Fleck |  |  | Pass
2015-11-17 |  | 73 °F |  | M. Fleck |  |  | Pass
Wheel saved and temperature recorded verified by calibrated thermometer.
Initial by:

## Temperature Recorder
Water Tank Temperature Recorder Calibration Worksheet
Equipment ID: | Water Storage Tank | Temp. Recorder |  |  | Date: | 2012
Manufacturer: |  | Palmer Instruments |  |  | Performed by: | Various
Model #: |  | X4-16133 |  |  |  | 6 Mo. Interval
Serial #: |  | NA |  |  | Next Calibration Due: | May / Nov.
UofA ID: |  | NA |  |  | Last Calibration: | May / Nov.
Location: | Concrete Moist Room
Calibration: |  | Automatic control of 23.0 + 2 °C
C. Procedure: | Calibration Procedure: | Verify that temperature recorder is correct (± 1° C = ± 1.8° F)
Equipment: |  | Nist Traceable Thermometer (T-Cal 2)
Date: |  | Recorder: | NIST Thermometer: | Verified by: | Pass / Fail:
2009-03-23 |  | 69° F | 73.0° F | F. Griffith | Adjustment made.
2009-05-05 |  | 73° F | 72.6° F | F. Griffith | Pass
2009-09-03 |  | 72° F | 72.0° F | F. Griffith | Pass
2009-09-15 |  | 72° F | 73.0° F | F. Griffith | Pass
2010-05-19 |  | 73° F | 73.0° F | F. Griffith | Pass
2010-12-13 |  | 73°F | 73.0° F | F. Griffith | Pass
2011-05-09 |  | 73° F | 73.0° F | F. Griffith | Pass
2011-05-16 |  | 73° F | 73.0° F | F. Griffith | Pass
2011-05-19 |  | 73° F | 73.0° F | F. Griffith | Pass
2011-12-12 |  | 73° F | 73.4° F | F. Griffith | Pass
2012-05-07 |  | 72° F | 72.8° F | M. Fleck | Pass
2012-05-23 |  | 73° F | 73.6° F | M. Fleck | Pass
2012-07-20 |  | 73° F | 73.4° F | M. Fleck | Pass
2012-07-27 |  | 73° F | 73.8° F | M. Fleck | Pass
2012-11-12 |  | 74° F | 73.8° F | M. Fleck | Pass
2012-11-19 |  | 74° F | 73.8° F | M. Fleck | Pass
2013-05-01 |  | 74° F | 74.6° F | M. Fleck | Pass
2014-05-06 |  | 74° F | 73.4° F | M. Fleck | Pass
2014-08-18 |  | 74° F | 74.5° F | M. Fleck | Pass
2014-09-08 |  | 75° F | 75.6° F | F. Griffith | Pass
2014-11-19 |  | 74° F | 72.5° F | M. Fleck | Pass
2015-05-04 |  | 73° F | 73.3° F | M. Fleck | Pass
2015-11-17 |  | 73° F | 72.5° F | M. Fleck | Pass
Initial by: